- Purpose:

To stipulate the required production and tests proceures for Yankauer suction set to ensure good quality products before releasing to the market.

- Scope:

This procedure is applicable for sterile Yankauer suction set producing by Saudi Mais company.

- Definition:

Yankauer Suction Set

The Yankauer suction set is oral suctioning medical device used in medical procedures. It is typically a firm suction tip with a large opening surrounded by a bulbous head and is designed to allow effective suction of blood or fluids without damaging surrounding tissue during surgical procedures.

- Responsibilities:

  - Quality Controllers are responsible for sampling & testing.

  - Quality control supervisors are responsible for checking the obtained results.

  - Machine operators and assembly peoples are responsible to produced good quantity and good quality products

- Procedure:

- Ensure the receiving slip from supplier is correct material as mentioned in purchase order.

- Inspect the material bag condition, must be no damaged at all bags with material.

- Take samples & inspect in order to AQL procedure, product description, specification with apply the AQL sampling plans for normal inspection.

- Record into OBM inspection report.

  - Inspection and test report for Raw material:

| Description | Requirement | Status |
| --- | --- | --- |
| Condition of Bag / Appearance | Good without damage bag | Visual |
| Material Grade and description | As per P. O | Visual |
| Appearance of Granules | Good / Transparent / colored | Visual / Physical |
| Granules Formation | No lobes formation | Visual |
| Lot no printed status | Clear printed | Visual |
| Arrangement of material bag | Should be specified and clear label | Visual |

  - Injection Machine Procedure

5.2.1  Injection Molding Components process.

Production controller is responsible to issue required raw material and production order for

the following components.

- Yankauer handle. (CTV Type) / (PTV Type)

- Connector for suction connecting tube.

    - Mold maintenance technician is responsible to prepare required mold for Yankauer handle and connector for suction connecting tube.

    - Machine maintenance technicians are responsible to mount the mold on specified machine and set approved parameters.

    - Machine operator is responsible to make sure that material is same as per production order before feeding the material into mutant feeding machine or feeding through by manual.

    - Machine operator is responsible to give initial sample to quality controller for inspection, quality controller is responsible to evaluate the products & make sure that the product is meeting with the criteria of specification.

    - Quality controller is responsible to take samples as per AQL general sampling method for visual inspection and AQL Special inspection method for performance testing of components.

    - Yankauer Hanlde

    - Qualtiy controller is responsible to make sure that Yankauer handle should be clear and good appearance, smooth surface, smooth end tip, no sharp, no flashing, no bubbles, no black, spots, crown tip vented or plan tip vented type, good ribs, vacuum vent holes proper open without blockage by flashing, good fitting of suction tube connector no loose and free from black spots, dust particle tip flashing.

    - Connector

- Qualtiy controller is responsible to make sure that connector should be soft and smooth surface accurate in dimeison in ID & OD as per specification, and should be free from flashing, black spots, crack marks, burn marks,  and irregularities defects.

- Quality controller is responsible to check the products randomly every 2 hours & record the results in process control form.

- Machine operator is responsible to arrange good products in clean polyethylene bag.

Production controller are responsible to weight the products and make inspection tag of products then the QC must be stamp on inspection tag to be confirm acceptable before transfer to store.

- Handle Bending process.

- Proudictotn supeviosrs are responsible to transfer required quantity of handle to Yankauer bending machine.

- Machine operators are responsible to make sure that standard parameters are set, and make sure that each units of machines and heaters are working.

- Place Yankauer handle on conveyer belt of mahcine, during placing yanker handle make sure that the air vent holes of yankauer handle should be right side.

- The given parameters will be used for bending Yankauer suction handles are.

- Parameters for bending Yankauer Handle:

| Yankauer Handle Bending M/C #01 | Temperature Setting | Holding Time | Pressure PSI |
| --- | --- | --- | --- |
|  | 180 - 240oC | 3 – 5 Sec | 90 –120 PSI |
| Yankauer Handle Bending M/C #02 | 180 – 240oC | 3 – 5 Sec |  |

    - Quality controller is responsible to check the products randomly every 2 hours & record the results in process control form.

    - Machine operator is responsible to arrange good products in clean polyethylene bag.

    - Production controller are responsible to weight the products and make inspection tag of products then the QC must be stamp on inspection tag to be confirm acceptable before transfer to store.

- In process control for Yankauer Handle.

| Sr | Description | Requirement | Status |
| --- | --- | --- | --- |
|  | Clarity / Appearance | Clean Good appearance / Clear | Visual |
|  | Handle Status | Smooth, No flashing, No bubbles, No black spot, etc. | Visual |
|  | Type of Handle | Vented Unvented | Visual |
|  | Tip quality and type | Crown tip Plain tip | Visual |
|  | Ribs quality | Good ribs / No offset | Visual /Physical |
|  | Windows & Vent status | Vented open, window open, free for flashing | Visual |
|  | Handle ID x OD | As per specification | Measurement |
|  | Bended quality | Must be perfect bended (Angle 20⁰ - 25⁰) | Physical / Measurement |
|  | Air Vent direction | Right position / Upper side direction | Visual |
|  | Fitting for suction tube | Good fiiting / Middle tight | Physical |

- Extrusion Machine Procedure:

  - Suction Connecting Tube process

- Production controller is responsible to issue raw material for Suciton connecting tube, and

issue production order after received raw material.

- Maintenance technicians are responsible to prepare required die for suction connecting tube and mount mold on specified machine by using standard parameters.

- Machine operators are responsible to feed the material by manual feeding.

- Machine operators are responsible to submit initial sample to quality controller for inspection.

    - Suction Connecting Tube Process Inspection

- Quality controller is responsible to make sure that tube should be accurate in dimension in ID / OD and length as per specification, smooth surface and tube should be free from black spots, sink marks, fish eyes, color faded, dust particles and should meet the standard criteria and specification as per in process control form.

- Quality controller is responsible to check the products randomly every 2 hours & record the results in process control form.

- Machine operator is responsible to arrange good products in clean polyethylene bag.

- Production controller are responsible to weight the products and make inspection tag of products then the QC must be stamp on inspection tag to be confirm acceptable before transfer to store.

- Yankauer Suction Set Specification.

- The Yankauer suction set complete dimension of components (tube, handle, connector) and material specification refer to TF-22-PS (product and raw material specification for Yankauer suction set).

- Assembly process

- Production supervisors are responsible to issue required components of Yankauer handle, tube and connector and issue production order for components assembly.

  - Connector and tube Assemlby process.

- Production controller are responsible to arrange components in specified tray and componnets should be with identification tag.

- Dip both end of tube in glue then removes excess glue by place tube on sponge and insert one end into universal connector and repeat process to assemble connector at opposite end of tube and make sure tube should be reached the internal walls of both connectors.

- Quality controller is responsible to check the products randomly every 2 hours & record the results in process control form.

- Machine operator is responsible to arrange good products in clean polyethylene bag.

- Production controller are responsible to weight the products and make inspection tag of products then the QC must be stamp on inspection tag to be confirm acceptable before transfer  for next process.

    - Q.C Inprocess control for suction connecting Tube.

| Description | Requirement | Status |
| --- | --- | --- |
| Material used speification | PVC material | Visual |
| Clarity / Appearance | Good Appearnce | Visual |
| Condition of Tubing | Flexible / Strong without any defect | Visual / Physical |
| Dimension of Tubing | As per specifications | Physical |
| Status of Gluing | Strong Gluing and not pull out | Physical |
| Fitting status for connector | Must be good fitting without leaking | Visual / physical |
| Condition of winding tube | Good winding without bend | Visual / Physical |

- Visual Inspection.

- There must be no visible contamination on the product.

- The tube surface must be smooth, no bumps, no bubbles and no flashing at all.

- Connector must be proper assemble with tube by strong gluing.

- Handle should clear no flashing and block spots.

- Handle tip should be crown tip or straight and no flashing.

- Tip vent holes should be open, no blockage.

- The tube and handle OD & ID should be as per our required specification , size and length.

- Over all good quality products no irregularities at all.

- Packaging and Labeling:

  - Production controller is responsible to issue packaging film and paper as per approved packaging film and paper specification and issue production order for packaging.

  - Machine operator is responsible to set the required packaging machine die and set approved standard parameters.

  - Machine operator is responsible to set the approved clichés and make sure that clichés information are correct and same as approved label.

  - Machine operator is responsible to give initial sample to quality controller for inspection and get approval from quality controller before packaging of the products.

  - Quality controller is responsible to the check, Product name and code no., reference no., size, lot no., type, temperature logo, sterilization logos, tender information, production date, expiry date, package forming, printing, sealing and cutting etc and quality must be stamp on pouch as approved.

  - After packaging then it will pack in specified small boxes, and arrange in big carton.

  - The big carton must be with label and detail information.

- Sterilization:

- The finished products are sterilized as per work instruction (WI-7-5-1-1).

- Final Inspection and Release:

- The Quality controller must be check, the unloaded products after sterilization.

- Quality controller is responsible to take sample as per AQL general inspection method for visual inspection and AQL special inspection method for performance testing of product and inspect the products as requirements mention in final inspection and release.

- The quality assurance manager or quality control manager have to review the completed cycle parameters and inspection reports both documents have to be sign.

- After confirmation all the tests from the Lab, to release the sterilized products, the document once again have to review the quality control manager then sign.

- The production department responsible to record the quantity of products before transferring to the finished goods ware-house and then it could be released to the customer.

- Sampling:

  - Inspection for PVC Granules 		 				:    F-827-i4

  - Inspection for PVC Granules	 					:    F-827-i6

  - Inprocess control for injection molding and extrusion 		:    F-823-i15

  - In-process control for Suction Tube 					:    F-823-i52

  - In-process control for Yankauer Handle				:    F-823-i171

  - In-process control for packaging machine				:    F-823-i4

  - Final Control for Yankauer Suction Set 				:    F-823-F44

  - Final Control for Suction Connecting Tube				:    F-823-F141

- Modification / Documentation change History

| Date | Changes | Revision |
| --- | --- | --- |
| 01.09.2018 | Add point, 5.1, 5.2, 5.5.1, 5.5, 5.6, 7, 8, 9. | 8 |
| 04.01.2022 | Updated Section 5.1 (PVC Mais Material) and Section (Sampling) | 9 |
| 01.09.2023 | 5,6,7,8,9,10,11,13 and 14 updated | 10 |