- Purpose:

Purpose of this procedure is for the production and testing procedure for Epidural set to ensure the good quality products before releasing to the market.

- Scope:

This procedure is applicable for the production and testing procedure for Epidural set

- Definition:

An epidural set is a combination of components which are used to access the epidural space to inject medication for local anesthetics and/or narcotics for relief of pain during surgery.

- Responsibilities:

  - Quality Controller is responsible for sampling & testing.

  - Quality control supervisor are responsible for checking the obtained results.

  - Machine operators and assembly peoples are responsible to produced good quantity and good quality products.

- Procedure:

  - Receiving inspection of Bulk components from supplier:

    - Ensure that received components are correct as per purchased order.

    - Make sure that the received components are as per specification and sizes.

    - Take samples & inspect as per AQL procedure, product description, specification with apply the AQL sampling plans for normal inspection.

    - Inspect the package condition, must be no damaged at all bags be a good sealed and correct component in order to specification.

  - Epidural Catheter bulk components are follows:

- Touhy Needle G-19 x 80mm

- Loss of Resistance Device Syringe 5ml.

- Epidural Catheter

- Catheter Connector

- Catheter Guide

- Epidural Catheter Clamp

- 0.2µ Luer Lock Flat Filter

- Applicator Sponge

- Hypodermic Needle G-21

- Hypodermic Needle G-25

- Gauze Swab White (5 x 5cm)

- Gauze Swab White (7.5 x 7.5cm)

- Adhesive Dressing

- Adhesive Tape

- Bandage

- Paper Drape

- Fenestrated Drape

- Luer lock syringe 3ml.

- Luer lock Syringe 5ml.

- Povidone iodine

- Surgical Drape

- Hospital Wrapping sheet

- Hips Tray-2

    - Inspection & Test Procedure for Epidural set components:

- Tuohy Needle and Stylet:

Quality inspector is responsible to make sure that the needle should be smooth surface, curved tip, with printed graduation marks, accurate in dimension in length, ID & OD, must be strong gluing with holder no pull out, no damage, no bend, freely movement of Stylet in needle, strong gluing of Stylet with holder and free from irregularities defects.

- Loss Resistance Syringe:

The syringe should be catheter tip, clear and transparent with printed graduation scale 5ml with light blue plunger, loss resistance movement of plunger in barrel, good fitting of gasket with barrel, no damage, no crack and no dust particle and black spot.

- Luer lock Flat filter:

It should be clear and transparent flat filter, with male luer lock and female port, good fitting of filter housing, no leakage, no crack, no damage, no dust particle and black spot.

- Gauze Swab:

Gauze should be white and cotton type, accurate in dimension (length and width) straight cutting edges, no black spot, no dust particle and should be free from stains and dust particle.

- Hypodermic Needle:

Needle should be smooth surface, sharp tip, no tip bend, no damage, no blockage, accurate in dimension, strong gluing with needle holder, no pull out and free from irregularities defects.

- Applicator Sponge:

It should be soft clear and dust free, good and adhesive with hooks no pull-out.

- 3ml & 5ml Luer Lock syringes:

Syringe should be clear and transparent, no flashing, no black spot, no dust particle, no crack, no damage, clear printed graduation scale, smooth movement of plunger in barrel and free from scratches.

- Epidural Catheter clamp:

Clamp should be free from flashing, black spot good fitting, no loose fitting, strong clamp of catheter, no leakage and without any irregularity’s defects.

- Catheter Connector:

Connector should be no crack, no damage, good fitting with female connector, no loose fitting and free from irregularities defects.

- Catheter Guide:

Catheter guide should be no flashing, smooth movement on catheter, no crack and damage, accurate in dimension without any irregularity’s defects.

- Swab sticks Povidone–Iodine:

The povidone-iodine swab sticks should not dry, good sealing pouch, packed in each pouch 3 pieces, and must be with valid expiry date.

- Fenestrated Drape:

Fenestrated drape should be accurate in dimension length & width, straight cutting edges, no dust particle, no black spot, no holes and stains.

- Paper Drape:

Paper drape should be white soft, accurate in dimension in length and width, good and straight cutting edges, no black spots and dust particles, no holes and free from stains.

- Hips Tray:

Tray should be compartmenting type, accurate in dimension, no crack and damaged, free from dust particles and black spot.

Quality inspector is responsible to record the results into OBM inspection reports of each component.

6. Issuance and receiving of Bulk Components for assembly:

  - Issuance:

- Production controller is responsible to issue the bulk components from bulk store for going on production of epidural set.

- Production controller is responsible to writer the complete information of components in issue slip and also mention exact quantity of the components.

  - Receiving of components:

- Production controller is responsible to make sure that the received components are same as per issue slip.

- Make sure that components are inspected and approved by Quality inspector and inspection tag is available as approved.

- Make sure that all the components are free from damaged crack and without any irregularity’s defects.

- Assembly of Epidural set:

- Production controller is responsible to issue production order for assembly for epidural set.

- Production controller is responsible to arrange the components in specified components trays and each tray’s must be with labels for the identification of components during assembly.

- Assembly personnel are responsible to arrange the epidural tray place the components one by one in designed place in a tray and make sure that all components are available and properly placed and make sure that the components are free from damaged and cracks and without any irregularity’s defects.

- Quality controller is responsible to check the products as per in process control form and make sure that all components are available in tray and meeting the criteria of specification.

- Quality controller is responsible to check the product label and make sure that the information on label are correct and components written on a label are same.

- Quality controller is responsible to record the results in process control form of epidural catheter set during assembly.

- After approval from quality controller, cover the tray by hospital wrap and pack in manual pouch and seal by manual sealing machine.

- Place approved label on the pouch and keep in designated area for final packaging.

- Packaging (Carton):

- Production controller is responsible to issue production order for packaging quantity epidural set.

- Packaging personnel are responsible to arrange carton and check one by one piece during packaging and make sure that pouch is properly sealed and without any irregularity’s defects.

- Packaging personnel are responsible to place the quantity in a carton to follow the packaging order.

- Place approved sticker on the external surface of the carton by glue and make sure it will not come out after dry.

- The big carton must be with label and detail information.

- Repeat the process for all remaining quantity and arrange carton on the pallet then transfer to sterilization section.

- Sterilization:

- The finished products are sterilized as per work instruction WI-7-5-1-1.

- Final control and Release:

- The Quality controller must be check, the unloaded products after sterilization.

- The quality assurance manager or quality control manager have to review the completed cycle parameters and inspection reports both documents have to be sign.

- After confirmation all the test from the Lab, to release the sterilized products, the document once again has to review the quality control manager then sign.

- The production department responsible to record the quantity of products before transferring to the finished goods ware-house and then it could be released to the customer.

- Documentation:

The results should be documented in quality records forms.

- Inspection & Test report for Touhy Needle		:	F-823-T017

- Inspection & Test report for General Products 	:	F-823-T90

- Inspection & Test report for Hypodermic Needle	:	F-823-T37

- Inspection & Test report for Syringes		:	F-823-T61

- In process control for Epidural set			:	F-823-i22

- In process control for manual sealing machine	:	F-823-i4-2

- Final control for Epidural Set				:	F-823-F20

- References:

- BS 6196

- ISO 7864

- ISO 9626

- ISO 80369-7

- ISO 80369-20

- ISO 7886-1

- ISO 22610

- ISO 14079

- Modification / Documentation Change History:

| Sr. | Date | Changes | Revisions |
| --- | --- | --- | --- |
| 1 | 18.05.2021 | First Issue | 00 |
| 2 | 27.02.2023 | Update final control | 01 |
| 3 |  |  |  |