## F-823 i14-5 (Rev.5)
Saudi Mais Co. for Medical Products | [14 empty] | Doc. # |  | : | F-823-i14-5
Quality Record |  | Ref. Doc. (WI-8-2-2-40 / TF-40-PS / TF-40-PDBM) | [12 empty] | Rev. # |  | : | 6
Title | In-process Control for Sodium Bicarbonate (Fresenius 5008 PE Bag Type) | [13 empty] | Date. |  | : | 11.11.2025
[15 empty] | Page # |  | : | 1 / 2
Type.# : ___________________ Date.# : _________________________ | [6 empty] | Instruments Used for Inspection
Lot # : __________________ Shift. # : _________________________
S.N | Work Description |  | Requirement | [5 empty] | Time (Hours) | [7 empty] | Quality Remarks
[5 empty] | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6
1 | Sheets Appearance |  | Clear, Transparent, No black spot, no line marks
2 | Sheets Cutting |  | Smooth and straight cutting, accurate in length as per specification
3 | Bag Welding Status |  | Bag should be proper welded from edges, no holes and gape in sealing
4 | Width of bag welded edges |  | Width required 4mm - 6mm, strong welded edges, no open, no gap and hole
5 | Bag Status |  | Smooth cutting edges, accurate in dimension as per specification, proper sealed edges, no holes, no open
6 | Holder Status |  | Smooth surface, no flashing, no damage, no crack free from irregularities
7 | Inlet filter assembly with holder |  | Proper inserted in holder inlet channel, no damage, no crack, no backout
8 | Hanager Status |  | Hanger should not broken and damage, and good fitting with holder
9 | Sealing Bag with holder |  | Bag should be proper sealed with holder, intact and no gape in sealing
10 | Porous Filter |  | White filter, accurate in dimension as per specification, no damage, no blockage
11 | Porous Filter Tube dimension |  | Smooth surface tube, no damage, no bend, no blockage and dimension should be as per specification
12 | Porous Filter fitting with tube |  | Proper fitted with tube, no loose, no pullout, no damage, no blockage
13 | Tube Assembly with holder |  | Proper inserted in holder outlet channel, good fitting, no loose, no backout, no bend and blockage
14 | Clarity of powder |  | Pure white powder
15 | Formation of Lumps |  | Without Lumps formation
16 | Bag filling Status |  | Accurate filling as per required weight, 650±10g, 750±10g, 840±10g, 900±10g
Saudi Mais Co. for Medical Products |  |  |  |  | Time (Hours) 5: Doc. # | : | F-823-i14-5
Quality Record |  | Ref. Doc. (WI-8-2-2-40 / TF-40-PS / TF-40-PDBM) |  |  | Time (Hours) 5: Rev. # | : | 6
Title | In-process Control for Sodium Bicarbonate (Fresenius 5008 PE Bag Type) |  |  |  | Time (Hours) 5: Date. | : | 11.11.2025
[5 empty] | Time (Hours) 5: Page # | : | 2 / 2
S.N | Work Description |  | Requirement |  | Time (Hours) 11: Time (Hours) | Quality Remarks
[5 empty] | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6
17 | Sealing of holder with plastic cover |  | Strong sealing, no leakage
18 | Bag Leak Test |  | Pressure 0.6bar (±0.2bar)
19 | Leak from holder port |  | Should not leak from ports
20 | Leakage from bag welded edges |  | Should not leak from bag welded edges
21 | Leakage from holder sealing with bag |  | Should be leakage free
22 | Holder fitting status into dialysis M/C |  | Good fitting , no damage, no leakage
23 | Pritned quality of sticker |  | Clear and visible printed, Complete information with UDI QR Code
24 | Packing condition |  | Good & Proper closed with UDI QR Code
Remarks :
Checked by Q.C. ( Name & Sign ) |  |  |  | Reviewed by: Q.C.M. :
Legend :
√ = Acceptable
X = Not Acceptable
R = Read Remarks N/A = Not Applicable
Legend :
√ = Acceptable
X = Not Acceptable
R = Read Remarks N/A = Not Applicable
Vernier caliper
Micro Scope
Micro Meter
Measuring Cylinder
Weight Balance
Cometech Machine
Measuring Scale/ Tape
Leake Testing Device