## F-823 i27 (Rev.7)
| Saudi Mais Co. for Medical Products | [10 empty] | Doc.# | : | F-823-i27
| Quality Record |  | Ref. Doc. (WI-8-2-2-29 / TF-29-PS / TF-29-PDBM) | [8 empty] | Rev.# | : | 7
| Title : In-process Control for 3 Way Stopcock Assembly | [10 empty] | Date# | : | 20.09.2026
[12 empty] | Page# | : | 1/1
| Date : ______________________________ |  | Shift : First Second | Instrument Used For Inspection
| Tube Length :___________________ |  | Lot : __________________________
| Sr | Description of Components | Requirements | Time ( Hours )
|  |  |  | 7 Am | 9 Am | 11 Am | 1 Pm | 3 Pm | 5 Pm | 7 Pm | 9 Pm | 11 Pm | 1 Am | 3 Am | 5 Am | 7 Am
| 1 | Clarity , Appearance | Clear , good appearance & free from any irregularities
| 2 | Dimension of tube | As per Specification ( MM )
| 3 | Tube gluing status | Should be strong gluing no pullout and no excess glue appeared
| 4 | Tube length required | As per Specification (CM)
| 5 | End Plug fitting | Good fitting without leakage
| 6 | Luer lock functioning | Good fitting without leakage
| 7 | Crack test after knob inserted | No crack at all side (upper , down)
| 8 | Status of housing with 2 FLL connector | Should be no crack , no scratch line
| 9 | Knob & housing fitting status | After fitting knob should be smooth movement & should not retracted from housing
| 10 | Knob hole status | Should be open without flashing , no blockage
| 11 | Leak test by air pressure | 1.0 bar , 15 - 20 seconds pressure. ( Free from Leakage )
| 12 | Leak test for liquid pressure | No leakage through by manual test
| 13 | Status movement of knob | Should not retracted from knob
| 14 | Pull test knob + housing | No pulled out
| 15 | Taken hourly leak test with needle Ø 0.40 mm punctured. | Must be detected the small hole from needle punctured , Accurately detected
Remarks:
Checked by Q.C. : | Checked By Q.C First Shift: |  | Checked By Q.C Second Shift: | Time ( Hours ) 9 Am: Reviewed By.
Legend :
√ = Acceptable
X = Not acceptable
R = Read Remarks
NA = Not applicable
Vernier caliper
Micro Scope
Cometech Machine
Measuring Scale/ Tape
Micro Meter
Measuring Cylinder
Leak Testing Device