## F-823 i27 (Rev.4)
Saudi Mais Co. for Medical Products | [10 empty] | Doc. # : F- 823-i27
Quality Record |  | Ref. Doc. (WI-8-2-2-29 / TF-29-PS / TF-29-PDBM) | [8 empty] | Rev. # : 6
Title : | [10 empty] | Date.# : 09.11.2025
| In-process Control for 3Way Stop Cock Assembly | [9 empty] | Page.# : 1 / 1
Type.# : ____________________ |  |  | Date.# : ______________ |  |  |  |  | Instrument Used For Inspection
Lot # : ____________________ |  |  | Shift # : ______________
S.N. | Description of Components | Requirements | Time ( Hours ) | [11 empty] | Quality Remarks
|  |  | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6
1 | Clarity , Appearance | Clear , Good Appearance
2 | Dimension Tube (Manual Assembly) | Ø 2.7 mm x 4.1 mm ± 0.1 mm
3 | Dimension Tube (Automation Assembly) | Ø 3.2 mm x 4.6 mm ± 0.1mm
4 | Tube gluing status | Strong , No gap
5 | Tube length required | 10±1.0cm, 25±1.0cm, 30±1.0cm, 50±1.0cm, 80±1.0cm 100±1.0cm
6 | Appearance of tube after glue | Good , No excess glue appeared
7 | End Plug fitting | Good fitting without leakage
8 | Luer lock functioning | Good fitting without leakage
9 | Crack test after knob inserted | No crack at all side (upper , down)
10 | Status of housing with 2 FLL connector | Should be no crack , no scratch line
11 | Movement status for knob + housing | Smooth movement
12 | Knob part fitting status | Smooth movement
13 | Knob hole | Should be open without flashing , no short shot
14 | Leak test for air pressure | 1.0 bar , 15 - 20 seconds pressure. ( Free for Leakage )
15 | Leak test for liquid pressure | No leakage through by manual test
16 | Status movement of knob | Should not retracted from knob
17 | Life test (Extreme test) continues rotation | Must be continues test up to 3 days
18 | Pull test between parts + tube | No pulled out
19 | Pull test knob + housing | No pulled out
20 | Taken hourly leak test with needle | Must be detected the small hole from
| Ø 0.40 mm punctured. | needle punctured , Accurately detected
Checked by Q.C. : |  |  | Time ( Hours ) 7: Reviewed by Q.C.M. :
Legend :
√ = Acceptable
X = Not acceptable
R = Read Remarks
NA = Not applicable
Vernier caliper
Micro Scope
Cometech Machine
Measuring Scale/ Tape
Micro Meter
Measuring Cylinder
Weight Balance