## F-823-i47 (Rev. 6)
Saudi Mais Co. for Medical Products | [10 empty] | Doc.# | : | F-823-i47
Quality Record |  | Ref. Doc. (WI-8-2-2-44 / TF-44-PS / TF-44-PDBM) | [8 empty] | Rev.# | : | 6
Title: | [10 empty] | Date.# | : | 28.04.2026
| In-process Control for Breathing Filter Assembly | [9 empty] | Page.# | : | 1/1
Size # : ________________________ Type #:_____________________________________________ Shift # : _____________ | [8 empty] | Instruments Used for Inspection
Date # : ______________________ Lot # : ____________________ Prod. Code#: ______________________________
Sr. | Descriptions | Requirements | Time (Hours)
|  |  | 7 | 9 | 11 | 1 | 3 | 5 | 7 | 9 | 11 | 1 | 3 | 5
1 | Clarity , Appearnce | Clear , White colour
2 | Lower Housing Condition | No crack, No damaged with good fitting for Endotracheal connector
3 | Upper housing condtion (Cover) | No crack, No damaged with good fitting for breathing circuit adaptor
4 | Bacterial Filter Required size (Adult) | Non woven filter Dia. 59±1mm
5 | Humidifying papper required Dimension (Adult) | Dia. 53±1mm
Width 15±1mm
6 | Bacterial Filter Required size (Pediatric) | Non woven filter Dia. 44±1mm
7 | Humidifying papper required Dimension (Pediatric) | Dia. 30±1mm
Width 15±1mm
8 | Deformity of housing on filter place | No defect at all
9 | Status of humididfying paper | Good position , Not movement after fixed on housing.
10 | Top filter condition (Non woven) | Good position , Not movement after fixed on housing
11 | Fitting status | Good fitting & Correct position.
12 | Appearance after sealing | Good apppearance and no damage all of side
13 | Sharp edges after sealing | No edges & Smooth.
14 | Leak test after sealing by under water, as per AQL sampleing plan. | No leakage from all of the sealing side
15 | End plug status | Good fitting , Good twisted.
Remarks:
Checked by: Q.C. |  |  | Time (Hours) 7: Reviewed by: Q.C.M.
Legend :
√ = Acceptable
X = Not Acceptable
R = Read Remarks
N/A = Not Applicable
Vernier caliper
Micro Scope
Leakage Testing Device
Measuring Scale/ Tape
Measuring Cylinder
Weight Balance
Micro Meter