- Reference:

Mucous Extractor 25 & 40ml.

- Definition:

A mucous extractor is a device which is indicated for use in new-born, infants or small children to clear their nasal passage or for suctioning tracheal/bronchial secretion samples of neonates, children and adults for the purpose of laboratory analysis.

- Scope:

This document describes the specifications of Mucous Extractor for single use delivered STERILE by Saudi Mais Co. for Medical Products.

- Quality Points:

- Transparent & Strong container for clear observation.

- Soft Suction tube with two lateral eyes & open tips

- Connector funnel & with finger control.

- Tube length 45-50cm.

- With anti-reflux valve.

- Made of PVC, PC, PE material.

- Product Code Number:

  - Mucous Extractor 25ml (Funnel Control Type Connector):

| SR. No. | Product Code | Description |
| --- | --- | --- |
|  | 371-0550 | Mucous Extractor with NRV Size CH 05 - Funnel Control Sterile |
|  | 371-0650 | Mucous Extractor with NRV Size CH 06 - Funnel Control Sterile |
|  | 371-0850 | Mucous Extractor with NRV Size CH 08 - Funnel Control Sterile |
|  | 371-1050 | Mucous Extractor with NRV Size CH 10 - Funnel Control Sterile |
|  | 371-1250 | Mucous Extractor with NRV Size CH 12 - Funnel Control Sterile |
|  | 371-1450 | Mucous Extractor with NRV Size CH 14 - Funnel Control Sterile |
|  | 371-1650 | Mucous Extractor with NRV Size CH 16 - Funnel Control Sterile |

  - Mucous Extractor 25ml (Finger Control Type Connector):

| SR. No. | Product Code | Description |
| --- | --- | --- |
|  | 370-0550 | Mucous Extractor with NRV Size CH 05 - Finger Control Sterile |
|  | 370-0650 | Mucous Extractor with NRV Size: CH 06 - Finger Control Sterile |
|  | 370-0850 | Mucous Extractor with NRV Size: CH 08 - Finger Control Sterile |
|  | 370-1050 | Mucous Extractor with NRV Size: CH 10 - Finger Control Sterile |
|  | 370-1250 | Mucous Extractor with NRV Size: CH 12 - Finger Control Sterile |
|  | 370-1450 | Mucous Extractor with NRV Size: CH 14 - Finger Control Sterile |
|  | 370-1650 | Mucous Extractor with NRV Size: CH 16 - Finger Control Sterile |
|  | 370-1600 | Mucous Extractor with NRV Size: CH 16 - with single chamber Sterile |

  - Mucous Extractor 40ml:

| SR. No. | Product Code | Description |
| --- | --- | --- |
|  | 372-1050 | Mucus Extractor Pediatric Size CH 10 - Funnel Control Sterile |
|  | 372-0850 | Mucus Extractor Pediatric Size CH 08 - Funnel Control Sterile. |
|  | 372-40ML | Mucus Extractor Specimen Plastic- For Collection 40ML |
|  | 372-80ML | Mucus Extractor Specimen Plastic- For Collection 80ML |

- Product Specification Table for Suction Tube:

- Product Drawings & Dimensions:

  - Mucous Extractor 25ml

    - Material Specification of Mucous Extractor Funnel Control Connector:

| Parts Name | Parts Descriptions | Materials |
| --- | --- | --- |
|  | Suction Tube | PVC |
|  | Suction Connecting Tube | PVC |
|  | Container | PC |
|  | Container Cap | PE |
|  | NRV | PE + Filter |
|  | Graduation Scale | Black Ink |
|  | Funnel Control Connector | PVC |

  - Mucous Extractor 25ml Finger Control Connector:

    - Material Specification of Mucous Extractor 25ml Finger Control Connector:

| Parts Name | Parts Descriptions | Materials |
| --- | --- | --- |
|  | Suction Tube | PVC |
|  | Suction Connecting Tube | PVC |
|  | Container | PC |
|  | Container Cap | PE |
|  | NRV | PE + Filter |
|  | Graduation Scale | Black Ink |
|  | Finger Control Connector | PVC |

  - Mucous Extractor 40ml:

    - Material Specification of Mucous Extractor 40ml:

| Parts Name | Parts Descriptions | Materials |
| --- | --- | --- |
|  | Suction Connecting Tube | PVC |
|  | Container Cap | PE |
|  | Container | PC |
|  | Suction Connecting Port | PE |
|  | Graduation Scale | Blue Ink |

| Product Attributes |  |
| --- | --- |
| Sr. No. | Parts |
|  | Suction Tube: Suction tubes illustrated are designed for removing small amounts of secretions from the nose, throat or ears & chest. |
|  | Graduation Scale: Graduation scale make by PVC black ink. The function of graduation scale help to measure tube for insertion on patient body. |
|  | Suction Connecting Tube: Suction connecting tube made of PVC material. The function of use for suction connecting tube adjusted with vacuum machine or syringes for sucking the mucus from patient body. |
|  | Funnel Type Connector: Funnel type connector made of PVC material. The function of use for connecting to the catheter tip syringe or vacuum machine. |
|  | Finger Control Type Connector: Finger control type connector made of PVC material. The function of use for connecting to the vacuum machine. |
|  | Container: The container made of PC material. The function of use for store the patient mucus. |
|  | Container Cap: The container cap made of Pe material. The function of use for closed the container. |
|  | Suction tube connecting port: The suction tube connecting port made by PE material. The function of use for connected the suction tube (funnel type) connector. |

- Physical Properties:

  - Visual appearance:

Clear / Good Appearance. No damaged, No Short Shot, No flashing.

  - Suction Tube status:

Good appearance. Clear & transparent. Open tip with two lateral eyes. Free from irregularities.

  - Suction tube connector:

Clear appearance, good status no flushing, No short shot. good fitting with suction tube. No leakage.

  - Container w/cap:

Clear & transparent. no flashing, no crack & no edge broken. Cap good/tight fitting with

container. The tube also good fitting with cap, no leakage & no pullout.

  - NRV (Non-Return Valve) status:

NRV has good appearance. Good fitting on inside the cap. NRV proper working.

  - Dimension status:

Dimensions are accurate are in accordance with specification table.

  - Tube fitting into connector:

The tube into parts fitting is in accordance with standard 8836:2019.

  - Leakage and Flow test

The tube leakage and flow rate test as per standard 8836:2019 Annex D.

- Biological properties:

After sterilization by ethylene oxide, the product fulfils the requirements of EN ISO 10993 for: 				Cytotoxicity,

Irritation,

Sensitization.

- Sterility:

Routine release for sterility is based on the control of physical parameters and a sterility control on biological indicators in accordance with EN 11135-1.

- Product Carton & Box specifications as per (F42-S10):

| SR. No. | Product Descriptions | 34x23x16.5cm |  | 51x35x36cm |
| --- | --- | --- | --- | --- |
|  |  | Pcs/Box | Small Box | Pcs/Carton |
|  |  | Qty. | Qty. | Qty. |
|  | Mucous Extractor (25ml) | 50 | 4 | 200 |
|  | Mucous Extractor (40ml) | 20 | 4 | 80 |

- Packaging and labelling:

Each Mucous Extractor 25 & 40ml is packed in a blister pack to guarantee the sterility of the inside when perforations are not broken.

The single piece label contains the following information:

- The item name and code / reference.

- For single use, ETO sterilized, read instruction, store temperature.

- Lot number and expiry date, production date.

- If latex contain or latex free.

- CE marking 1639.

- Manufacturer name and address.

- Authorized European Representative Name and address.

The label on the shipping big carton bears at least the following information:

- Item name and code / reference.

- Lot number.

- Sterilization date.

- For single use, ETO sterilized, read instruction, store temperature.

- CE marking 1639.

- Manufacturer name and address.

- Authorized European Representative Name and address.

- Quantity per carton.

-And the cartons contain the fallowing symbol; keep side up – upside down prohibited.

- Do not expose to direct sunlight, store at room temperature.

- Handle with care.

- Use no hooks.

- Fragile.

- Do not stock more than 8 cartons.

- Expiry:

The expiry date is (05) years from its Production date.

- Documents Changing History:

| Date | Changing History | Revision No. |
| --- | --- | --- |
| 22.09.2014 | Added Doc. History table | 06 |
| 06.02.2023 | Added technical drawings | 07 |

| Prepared By: Q.A. Coordinator: | Checked By: Q.C.M: | Approved By: Q.A.M: |
| --- | --- | --- |