QUALITY CONTROL SYSTEMS  /  EVIDENCE-BASED TECHNICAL REVIEW

# 3-Way Stopcock with Extension Tube — Revised In-Process Control Record Strategy

3-Way Stopcock with Extension Tube

| Issue status | PROPOSED CONTROLLED DRAFT — NOT APPROVED |
| --- | --- |
| Date prepared | 22 Oct 2026 |
| Document type | Word process-control report / companion to six independent XLSX forms |
| Controlled source | WI-8-2-2-29 Rev.8 |
| Assessment basis | IN-PROCESS ONLY — not final inspection/release |
| Use restriction | For QC/QA review only. No product acceptance or release permitted on draft alone. |

Evidence rule: “Required” is not “available”. No gap is verified closed without approved documentation and objective records.

# 1. Operational record structure

| ID | Record title | Original split preserved | Applicability |
| --- | --- | --- | --- |
| F-823-i27 | 3-Way Stopcock Assembly | 15 original checks retained, stopcock control | Primary |
| F-823-i15 | Moulding & extrusion | 21 original legends retained with component/line identification | Shared |
| F-823-i4-7 | Neil Pack packaging | 3 source settings and 16 final checks | Shared |
| F-823-i4-1 | Zip Lock bag machines | Original blow/sealing machine grouping preserved | Shared |
| F-823-i13 | CVC Set | Separate independent CVC record | Separate / not superseded |
| F-823-i13-2 | CVC Complete Kit | Independent two-page CVC record | Separate / not superseded |

# 2. Field requirements for revised operational record

Product/part and approved BOM/specification revision; lot, production work order, date, machine/shift, material lot and operator/QC.

Characteristic plus ORIGINAL COMPANY acceptance criterion (with unresolved criteria visibly indicated); approved verification route and controlled specification ID.

Actual time-stamped observation/measurement, equipment ID/calibration status, accept/reject/pending, remarks, sample identity and applicable evidence reference.

Signature/review fields remain blank for actual authorized review. A proposed draft or blank test report cannot count as executed evidence.

No change to approved sample interval, process limits or AQL before QC/QA/Engineering authorization.

# 3. Gap-to-form mapping

| Gap | Priority | Operational form IDs | Change needed / alternative evidence |
| --- | --- | --- | --- |
| 3WS-G01 | P1 | i27, i15 | Add connector classification, approved drawing ID/rev, controlled dimensions and actual results or type-qualification ID; distinguish tool approval from hourly appearance screening. |
| 3WS-G02 | P1 | i15 | Material grade/revision, supplier/lot, qualified modulus test report or verified CoA, approval result and molded lot linkage. |
| 3WS-G03 | P1 | i27 | Screen: observed air leakage, actual pressure/time, equipment ID and lot/time. Qualification: test method/ref connector, pressure (kPa), hold (s), measured leakage (Pa·m³/s), report ID, Pass/Fail. |
| 3WS-G04 | P1 | i27 | Define source route (qualification vs routine), test pressure (kPa), duration (s), water observation (falling drop Y/N), apparatus ID and report. |
| 3WS-G05 | P1 | i27 | Record approved type qualification ID, reference connector and change-control approval; qualification record includes pressure, hold, measured leakage and verdict. |
| 3WS-G06 | P1 | i27, i15 | Type-test report, stress exposure ID, post-stress leakage, affected part/design revision, result and approval. |
| 3WS-G07 | P1 | i27 | Connection class L1/L2; target force (N); applied force (N); hold (s); separation Y/N; reference connector; type-test report. |
| 3WS-G08 | P1 | i27 | Record connector type, torque (N·m), hold (s), disengagement result, method and qualification report ID. |
| 3WS-G09 | P1 | i27 | Qualified testing: connector design ID, reference connector ID, applied torque, hold, overriding Y/N, pass/fail and report. |
| 3WS-G10 | P2 | i27 | Tubing bond: approved specification ID/rev; minimum force (N) [OPEN]; applied (N); hold (s); observed separation; result. Keep knob retention check separate. |
| 3WS-G11 | P2 | i27 | Line/machine mode; actual screening pressure/time; leak signal; challenge test, equipment ID; record ID; batch/time/operator; disposition. |
| 3WS-G12 | P1 | i4-7, i27 | Reference EO load ID, sterilization record ID, packaging seal-validation references, and independent authorized release sign-off on designated controlled record. |
| 3WS-G13 | P2 | i4-7, i4-1 | Controlled process form ID/rev; validated parameter/specification ID; actual temperature/time/width; machine/shift/lot; sealing/cut result and signature. |

# 4. Test frequency / evidence ownership

| Frequency route | Control owner | Record / requirement |
| --- | --- | --- |
| At designated source time / shift | QC / Production | Appearance, line screening, dimensions and product count; preserve source schedule and approved SOP. |
| Incoming/supplier | Supplier Quality / QC | Purchased accessory certificate, grade, drawing, component lot and approved supplier dossier. |
| Type/qualification | Engineering / QA | Standard method, report and reference connector for qualified product design; no invented routine lot frequency. |
| Periodic / change driven | Engineering / QA | Risk/change assessment and qualifications after tooling/material/sterilization/packaging changes. |
| After EO | QA / Authorized release | Cycle/BI where applicable, final inspection and formal disposition, not an in-process checkbox. |

# 5. Original requirement preservation & approval restrictions

The companion Excel forms retain the source checks and their controlled identities. Any process target without verified approved engineering/supplier specification is shown as a SOURCE COMPANY CRITERION — TECHNICAL CONFIRMATION PENDING. All files are DRAFT candidates; their actual implementation requires controlled revision, training and line qualification where applicable.

# 6. Known outstanding records

OPEN — Supplier product qualification/type-test records

OPEN — Controlled product drawings, DD-02-PS / TF-29-PS and BOM

OPEN — ISO 80369-20 for detailed test methods when applicable

OPEN — EO validation and individual load release evidence

OPEN — Current company controlled form index and exact approved sampling interval

OPEN — Source F-823-i27A where automation is used