# 3-Way Stopcock | Revised In-Process Control Record

MANUAL + AUTOMATED ASSEMBLY • DRAFT REPLACEMENT FOR F-823-i27 / F-823-i27A

Source records: F-823-i27 Rev.6 (20 items), F-823-i27A Rev.5 (35 stations)

Controlling WI: WI-8-2-2-29 Rev.8; proposed revised WI is NOT approved

Use: Manufacturing IPC only - NOT Final Inspection, Sterility or Final Release

DRAFT FOR TECHNICAL REVIEW • No supplied source has been modified. Actual test conformity and product release are not established.

# A. Production / Sampling / Specification Header

Production date: __________  Shift: ______  Product code/configuration: ______________________

Route (tick): [ ] Manual  [ ] Automation  Production order: __________  Lot: __________

Tube nominal/size: __________  Machine / Line: __________  Cavity: __________  Material lot(s): __________

TF-29-PS Rev (verified): __________  TF-29-PDBM Rev: ______  Machine recipe Rev: ______

Inspection interval / 3h population: __________  Sampling plan SOP-8-2-6 Rev.9 / specific approved plan: __________

Sampling level: _____  AQL: _____  Code: _____  Sample n: _____  Ac: _____  Re: _____

Instrument ID(s): __________  Calibration validity: __________  Controlled log / file ID: __________

IMPORTANT: Conditions identified OPEN TECHNICAL INPUT REQUIRED must be approved before use in an issued controlled form. Nominal 10 cm length shows source conflict (WI +/-0.5 cm versus forms +/-1 cm). Do not pass a result using the wider form tolerance pending approved resolution.

# B. Manual Assembly - Reconstructed IPC Checklist (20 preserved characteristics)

| No. | Characteristic | Acceptance / approved reference | Actual / Result + unit | Time / sample | Decision | Evidence / remark |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Clarity , Appearance | Clear , Good Appearance | ________ | _______ | P / F / N/A / Hold | _________ |
| 2 | Dimension Tube (Manual Assembly) | Manual ID / OD separately: 2.7 / 4.1 mm +/-0.1 mm subject to confirmed drawing | ________ | _______ | P / F / N/A / Hold | _________ |
| 3 | Dimension Tube (Automation Assembly) | N/A for manual route; automation ID / OD 3.2 / 4.6 +/-0.1 mm per source WI subject to approved drawing | ________ | _______ | P / F / N/A / Hold | _________ |
| 4 | Tube gluing status | Strong , No gap | ________ | _______ | P / F / N/A / Hold | _________ |
| 5 | Tube length required | 10 cm: current WI 10 +/-0.5 cm (spec confirmation PENDING); 25/30/50/80/100 cm +/-1 cm if approved | ________ | _______ | P / F / N/A / Hold | _________ |
| 6 | Appearance of tube after glue | Good , No excess glue appeared | ________ | _______ | P / F / N/A / Hold | _________ |
| 7 | End Plug fitting | Good fitting without leakage | ________ | _______ | P / F / N/A / Hold | _________ |
| 8 | Luer lock functioning | Good fitting without leakage | ________ | _______ | P / F / N/A / Hold | _________ |
| 9 | Crack test after knob inserted | No crack at all side (upper , down) | ________ | _______ | P / F / N/A / Hold | _________ |
| 10 | Status of housing with 2 FLL connector | Should be no crack , no scratch line | ________ | _______ | P / F / N/A / Hold | _________ |
| 11 | Movement status for knob + housing | Smooth movement | ________ | _______ | P / F / N/A / Hold | _________ |
| 12 | Knob part fitting status | Smooth movement | ________ | _______ | P / F / N/A / Hold | _________ |
| 13 | Knob hole | Should be open without flashing , no short shot | ________ | _______ | P / F / N/A / Hold | _________ |
| 14 | Leak test for air pressure | Company water-bath IPC: 1.0 bar, 15-20 s; no observed leak | ________ | _______ | P / F / N/A / Hold | _________ |
| 15 | Leak test for liquid pressure | No leakage through by manual test | ________ | _______ | P / F / N/A / Hold | _________ |
| 16 | Status movement of knob | Should not retracted from knob | ________ | _______ | P / F / N/A / Hold | _________ |
| 17 | Life test (Extreme test) continues rotation | Link valid 3-day endurance report / Pending; separate from hourly measurement | ________ | _______ | P / F / N/A / Hold | _________ |
| 18 | Pull test between parts + tube | No separation under approved force (N) and duration (s) [OPEN TECH INPUT] | ________ | _______ | P / F / N/A / Hold | _________ |
| 19 | Pull test knob + housing | No separation under approved force (N) and duration (s) [OPEN TECH INPUT] | ________ | _______ | P / F / N/A / Hold | _________ |
| 20 | Taken hourly leak test with needle Ø0.40 mm punctured | 0.40 mm needle-punctured challenge detected and reject routing confirmed; hourly source WI | ________ | _______ | P / F / N/A / Hold | _________ |

# C. Automation Assembly - Reconstructed 35 Station IPC Checklist

| Station | Characteristic | Acceptance / controlled setting | Actual / Reading | Time / cavity | Decision | Evidence / remark |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Mainbody loading unit (Housing Parts) | Clear , Good appearance | ________ | ________ | P / F / Hold | ________ |
| 2 | Mainbody loading rotation 90⁰ | Should be rotate at 90⁰ not pulled out | ________ | ________ | P / F / Hold | ________ |
| 3 | Silicon unit for mainbody | Should be present of lubricant into mainbody | ________ | ________ | P / F / Hold | ________ |
| 4 | Knob insertion unit | Should be accurate of lubricant while insertion | ________ | ________ | P / F / Hold | ________ |
| 5 | First knob status after insrted | Should be the knob perfected inserted on the housing | ________ | ________ | P / F / Hold | ________ |
| 6 | Second knob status after insrted | Should be the knob perfected inserted on the housing | ________ | ________ | P / F / Hold | ________ |
| 7 | Presence first knob checking unit | Accurate and good inserted | ________ | ________ | P / F / Hold | ________ |
| 8 | Presence second knob checking unit | Accurate and good inserted | ________ | ________ | P / F / Hold | ________ |
| 9 | Pusher rod unit for first line nozzle side | Should be strong , no breaking | ________ | ________ | P / F / Hold | ________ |
| 10 | Pusher rod unit for second line nozzle side | Should be strong , no breaking | ________ | ________ | P / F / Hold | ________ |
| 11 | T-stopper of loading knob on first line | Should be not breaked | ________ | ________ | P / F / Hold | ________ |
| 12 | T-stopper of loading knob on second line | Should be not breaked | ________ | ________ | P / F / Hold | ________ |
| 13 | T-stopper presence | Should be presence of T- stopper | ________ | ________ | P / F / Hold | ________ |
| 14 | T-stopper presence Tighting | Good fitting into housing parts | ________ | ________ | P / F / Hold | ________ |
| 15 | Assembled rotation unit on first line | Good movement , smoothness in rotation at 90⁰ | ________ | ________ | P / F / Hold | ________ |
| 16 | Assembled rotation unit on second line | Good movement , smoothness in rotation at 90⁰ | ________ | ________ | P / F / Hold | ________ |
| 17 | Tighting T-stopper (second option) | Smooth fitting | ________ | ________ | P / F / Hold | ________ |
| 18 | Leak test unit | No detected leak; machine nominal 2.4 bar x 0.2 s; record actual recipe/log and verified reject event | ________ | ________ | P / F / Hold | ________ |
| 19 | Gluing unit on mainbody (Nozzle side) | Good gluing and No excess gluing | ________ | ________ | P / F / Hold | ________ |
| 20 | Tubing cutting unit on first line | Tube cut 10 +/-0.5 cm per current WI (form +/-1 cm conflict); approval/spec required | ________ | ________ | P / F / Hold | ________ |
| 21 | Tubing cutting unit on second line | Tube cut 10 +/-0.5 cm per current WI (form +/-1 cm conflict); approval/spec required | ________ | ________ | P / F / Hold | ________ |
| 22 | Tubing presence checking unit | Perfect inserted all the way without gap | ________ | ________ | P / F / Hold | ________ |
| 23 | Gluing unit on internal tubing | Good gluing and No excess gluing | ________ | ________ | P / F / Hold | ________ |
| 24 | Male luerlock loading on first line | Perfect inserted all the way without gap | ________ | ________ | P / F / Hold | ________ |
| 25 | Male luerlock loading on second line | Perfect inserted all the way without gap | ________ | ________ | P / F / Hold | ________ |
| 26 | Pull test on top end | No separation under approved joint force/time [OPEN TECHNICAL INPUT] | ________ | ________ | P / F / Hold | ________ |
| 27 | Pull test on bottom end | No separation under approved joint force/time [OPEN TECHNICAL INPUT] | ________ | ________ | P / F / Hold | ________ |
| 28 | Leak test unit in whole assembled | No detected leak; machine nominal 2.4 bar x 0.2 s; record actual recipe/log and verified reject event | ________ | ________ | P / F / Hold | ________ |
| 29 | End cap loading unit | Perfect inserted at all | ________ | ________ | P / F / Hold | ________ |
| 30 | End cap presence checking | Not pulledout | ________ | ________ | P / F / Hold | ________ |
| 31 | Knob rotation to home position | Smooth movement while rotation | ________ | ________ | P / F / Hold | ________ |
| 32 | T-stopper injection unit on first line | Accurate and good injected | ________ | ________ | P / F / Hold | ________ |
| 33 | T-stopper injection unit on second line | Accurate and good injected | ________ | ________ | P / F / Hold | ________ |
| 34 | Good products (Final product) | Challenge article passes intended good/reject routing, with count reconciliation | ________ | ________ | P / F / Hold | ________ |
| 35 | Rejected products (Rejected) | Challenge article passes intended good/reject routing, with count reconciliation | ________ | ________ | P / F / Hold | ________ |

# D. Quantitative IPC Measurement / Leak Logging

Repeat the following rows for every interval and sample in the approved sampling plan. Attach an equipment log when it contains the same traceable fields.

| Date/time | Sample / cavity | Test / unit | Approved value / limit | Actual value | Instrument / recipe ID | P/F | Printout / evidence ID |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ________ | ________ | Tube ID / OD / length | ________ | ________ | ________ | ___ | ________ |
| ________ | ________ | Tube ID / OD / length | ________ | ________ | ________ | ___ | ________ |
| ________ | ________ | Tube ID / OD / length | ________ | ________ | ________ | ___ | ________ |
| ________ | ________ | Tube ID / OD / length | ________ | ________ | ________ | ___ | ________ |
| ________ | ________ | Tube ID / OD / length | ________ | ________ | ________ | ___ | ________ |
| ________ | ________ | Tube ID / OD / length | ________ | ________ | ________ | ___ | ________ |

# E. Detector/Reject Challenge and Machine Settings

| Time | Positive challenge ID | Machine leak pressure / hold (actual) | Detector result | Expected station | Actual station | Reconciled reject qty | P/F / NCR |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ______ | _______ | _______ bar / _____ s | Detected / Not | 34 / 35 | ______ | ______ | _______ |
| ______ | _______ | _______ bar / _____ s | Detected / Not | 34 / 35 | ______ | ______ | _______ |
| ______ | _______ | _______ bar / _____ s | Detected / Not | 34 / 35 | ______ | ______ | _______ |
| ______ | _______ | _______ bar / _____ s | Detected / Not | 34 / 35 | ______ | ______ | _______ |

# F. Separate Connector Qualification / Alternative Evidence References (not automatically hourly)

| Test route | Approved ISO application part | Controlled test procedure | Test report / linked batch | Result / applicable limit | Scope review |
| --- | --- | --- | --- | --- | --- |
| ISO 80369-20 Annex B - pressure decay | __________ | WI-8-2-6-05 | __________ | ______ Pa.m3/s / ______ | P / F / Pending |
| ISO 80369-20 Annex C - falling drop | __________ | WI-8-2-6-06 | __________ | Pressure ___; hold ___; drop Y/N | P / F / Pending |
| Joint qualification / 3-day knob life | Approved company protocol | __________ | __________ | Force/cycles/duration ______ | P / F / Pending |

# G. Nonconforming Output, Containment and Review

Defect / first fail time and station: __________________  Last-good verification time: __________________

Affected production quantity and cavity/lot: __________________  Material hold location / On-Hold ID: __________________

NCR / Deviation ID: ______________  Immediate containment: ____________________________________________

Corrective action / recheck evidence: __________________  Authorization to restart (function/date/time): __________

Sampling or alternate test report ID: _______________  Final disposition reference: ________________________

Performed by QC: _____________________  Date/time: ____________  Reviewed by QCM / delegated authority: __________________

IMPORTANT: QC in-process Pass is not QA final release; formal final inspection/sterilization release remains separate.