# IN-PROCESS CONTROL CLOSURE REGISTER

Sterile Amnion Hook | Evidence and source verification register | No ISO closure asserted

CONTROLLED DRAFT • Proposed wording only. No supplied record demonstrates completed operational checks, approvals or product conformity.

# 1. Closure rules and current status

This register tracks findings generated by comparing the supplied WI and blank F-823-i5 form. It is not an ISO clause register because no applicable ISO standard was supplied. All gap corrective text appears only in unapproved draft documents; none can be verified closed without approved technical source, authorizations and executed objective evidence.

Status definitions: DRAFT ADDRESSED = proposed corrective text written but not authorized; OPEN — TECHNICAL INPUT = required controlled specification or rule absent; OPEN — EVIDENCE = approved executed record not verified; OPEN — REFERENCE = document master verification outstanding; VERIFIED CLOSED = only after source, authorization and appropriate executed evidence are reviewed (zero verified closed here).

# 2. Gap closure matrix

| Gap ID / risk | Current issue | Source trace | Proposed documentary coverage | Required objective evidence | Verified evidence | ISO / documentary status | Operational status | Function owner |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GAP-001 ¦ P1 | Hook formation, tip integrity and bend/strength criteria | WI p.2, §§5.2.1 and 5.3; F-823-i5, row 3, p.1 | Proposed revised WI/form draft; NOT approved | approved TF-26-PS/drawing and technical decision on strength check; completed in-process log, relevant disposition records and approval of revised controls. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — TECHNICAL INPUT / APPROVAL / EVIDENCE | QC / Engineering / QA |
| GAP-002 ¦ P2 | Dimensional measurement evidence | WI p.2, §5.3; p.3, §6; F-823-i5, row 2, p.1 | Proposed revised WI/form draft; NOT approved | controlled TF-26-PS/drawing; approved measurement instructions; instrument identification/calibration record; filled dimensional results. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — TECHNICAL INPUT / APPROVAL / EVIDENCE | QC / Engineering |
| GAP-003 ¦ P2 | Incomplete explicit coverage of WI visual defects | WI p.2, §§5.2.1 and 5.3; F-823-i5, rows 1 and 4, p.1 | Proposed revised WI/form draft; NOT approved | controlled visual defect specification/classification where applicable; completed field-level visual checks and NCR evidence for failures. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — TECHNICAL INPUT / APPROVAL / EVIDENCE | QC / Engineering |
| GAP-004 ¦ P2 | Two-hour frequency not evidenced unambiguously | WI p.2, §5.2.1; F-823-i5, column header “Time (Hours)”, p.1 | Proposed revised WI/form draft; NOT approved | signed completed production/shift logs and actual check times demonstrating interval compliance, or deviation/disposition records. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — APPROVAL / EXECUTED TIME LOG | QC / Production |
| GAP-005 ¦ P2 | AQL and sampling plan identification | WI p.2, §5.2.1; F-823-i5, p.1, all four checks | Proposed revised WI/form draft; NOT approved | approved AQL and/or in-process sampling control plan, completed sample/decision records. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — TECHNICAL INPUT / APPROVAL / EVIDENCE | QC / QA |
| GAP-006 ¦ P2 | Process-run traceability is incomplete on this form | WI p.2, §5.2.1; F-823-i5, header, p.1 | Proposed revised WI/form draft; NOT approved | controlled batch/production record or completed new form tying each check to machine/run; verification of record linkage. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — ALTERNATIVE RECORD VERIFICATION | Production / QC |
| GAP-007 ¦ P2 | Failure escalation, product hold and disposition recording | WI p.2, §§5.2.1 and 5.3; F-823-i5, legend and Remarks, p.1 | Proposed revised WI/form draft; NOT approved | approved NC/hold procedure, completed hold/NCR and traceable disposition/re-inspection evidence. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — TECHNICAL INPUT / APPROVAL / EVIDENCE | QC / QA / Production |
| GAP-008 ¦ P2 | Instrument / method and objective-record linkage | WI p.2, §5.3; p.3, §6; F-823-i5, “Instruments Used for Inspection” column, p.1 | Proposed revised WI/form draft; NOT approved | approved test/measurement method, instrument control records and signed measured-result worksheet. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — TECHNICAL INPUT / APPROVAL / EVIDENCE | QC / Metrology |
| GAP-009 ¦ P3 | Unresolved WI reference suffix on controlled form | WI p.4, §10; WI all-page header; F-823-i5, header “Ref. Doc”, p.1 | Proposed revised WI/form draft; NOT approved | authorized document master list, referenced WI-8-2-2-26-1 if it exists, TF-26-PDBM and approved form/WI reference verification. | Not supplied / not verified | ISO: not assessed. Draft addressed only. | OPEN — REFERENCE VERIFICATION | Document Control / QA |

# 3. Non-gap checks and review boundaries

| Control referenced in WI | Current verification | Decision |
| --- | --- | --- |
| Incoming raw material — F-823-T1 | Not supplied | Excluded from this in-process-form closure review. |
| Packaging in-process — F-823-i4 | Not supplied | No negative assessment of packaging controls. |
| Sterilization — WI-7-5-1-1 | Not supplied | No sterilization validation/conformity statement. |
| Final release — F-823-F39 | Not supplied | No final-release conformity statement. |
| TF-26-PS specification | Mentioned WI p.3 §6 and form header | OPEN: retrieve controlled revision and actual limits. |
| Form Ref. Doc. WI-8-2-2-26-1 / TF-26-PDBM | Header visible; sources unavailable | OPEN: verify valid identifiers and applicability. |

# 4. QC Manager closure authorization

For each item, close only after the responsible function confirms: (1) approved technical/control source; (2) approved WI and form revision; (3) applicable completed record or verified alternate evidence; (4) documented disposition of observed deviations; (5) correct mapping to lot/run, date and signature. If evidence is missing, retain OPEN status. No closure approval is recorded in this draft.

Manager decision / technical input: ____________________________________________________________________________________________

Reviewed by / date: _____________________________   Approved by (per site authority) / date: _____________________________