# ISO APPLICABILITY & CLOSURE REGISTER

Breathing Circuit Components / Breathing Filter / Potential HME — Conditional evidence register

DRAFT FOR TECHNICAL REVIEW • Evidence-based assessment • 15 October 2026

Document status: Not approved. No evidence of actual product conformity or operational closure was supplied.

This document is NOT a certificate. "No evidence supplied" is not the same as "test failed". No test performance is inferred. Actual WI, approved specification and product classification are pending. Conditional ISO requirements remain OPEN VERIFICATION until applicability is resolved.

# 1. Standard-level register

| Requirement ID | Clause and pages | ISO Says (paraphrase) | Control route | Gap ID | Actual verified evidence | Status | Outstanding action |
| --- | --- | --- | --- | --- | --- | --- | --- |
| BSF-01 | ISO 23328-1:2003 Cl.1; PDF9 / printed1 | BSF salt-aerosol method scope, separate from non-filtration aspects. | Conditional - confirm BSF intended use | GAP-008 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |
| BSF-02 | ISO 23328-1:2003 Cls.3.2-3.6; PDF10-11 / printed2-3 | Ambient 23±2 °C, 60±15% RH, 96±10 kPa; adult test flow 30 L/min, pediatric 15 L/min; unconditioned and conditioned states; documented sample-size rationale. | Conditional qualification | GAP-008 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |
| BSF-03 | ISO 23328-1:2003 Cls.4-5; PDF12 / printed4 | PV=100×(cP/cC); efficiency=100−PV. Report device identification, manufacture/lot, quantity tested and sample-by-sample efficiency in each state. No general minimum % in Part 1. | Conditional qualification | GAP-008 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |
| BSF-04 | ISO 23328-1:2003 normative Annex A; PDF13-14 / printed5-6 | Condition according to intended use; maximum recommended duration or 25±1 h if unstated; test within 5 min after conditioning. | Conditional qualification | GAP-008 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |
| HME-01 | ISO 9360-1:2000 Cls.1, 5.1-5.3; PDF11-13 / printed1-3 | HME scope; qualifying patient-port cone connection and accessory ports; sterile HME packaging according to referenced packaging standard. | Conditional HME | GAP-009 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |
| HME-02 | ISO 9360-1:2000 Cls.6.1-6.5; PDF13-25 / printed3-15 | HME moisture-loss, pressure-drop (30/60/90 L/min), gas-leakage (7±0.5 kPa), compliance test methods. Test-method values, not generally mandated product acceptance limits. | Conditional HME | GAP-009 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |
| HME-03 | ISO 9360-1:2000 Cls.7.1-7.4; PDF25-26 / printed15-16 | Applicable HME marking, packaging declarations, performance figures, intended tidal-volume range and supplier-provided information. | Conditional HME | GAP-009 | None supplied; Pending | Documentary: provisional only Operational: OPEN - EVIDENCE | QA to determine applicability; retrieve controlled qualification/type/supplier records and reconcile acceptance criteria. |

# 2. Company in-process documentation findings

| Gap | Priority | Issue | Required evidence / closure criteria | Verified evidence | Current closure |
| --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 | Incorrect WI baseline and unverified process route | Required: matching approved WI, approved TF-44-PS/TF-44-PDBM, effective revision register and cross-reference. Route: Document control / production. Evidence: Pending. Closure: verify product-specific approved WI controls and revise this draft against actual WI. | None supplied beyond blank photographed forms | Open - Evidence |
| GAP-002 | P1 | Breathing-circuit leak-check method is undefined in the form | Required: approved leak-test method + product acceptance criterion, calibration, filled data for representative lot, acceptance/disposition. Route: In-process; evidence pending. Close only after WI, form and actual records are approved/verified. | None supplied beyond blank photographed forms | Open - Technical Input |
| GAP-003 | P1 | Breathing-filter sealing-water test is insufficiently defined | Required: approved test method and AQL/control plan, verified equipment and completed tested-lot seal evidence. Route: In-process. Not proof of filtration efficiency. | None supplied beyond blank photographed forms | Open - Technical Input |
| GAP-004 | P2 | Bonding/retention requirement is subjective | Required: controlled bonding specification, validation/testing evidence, calibrated device if quantitative, sampled completed records. Route: In-process / qualification as defined in approved plan. | None supplied beyond blank photographed forms | Open - Technical Input |
| GAP-005 | P2 | Swivel fit, motion and O-ring function lack objective acceptance | Required: approved connector drawings, gauge selection, functional test criteria, records and disposition. Route: In-process. | None supplied beyond blank photographed forms | Open - Technical Input |
| GAP-006 | P2 | Sampling and time checkpoints cannot be validated | Required: approved process sampling plan, WI frequency reference, dated signed in-process records showing actual samples and corrections. Route: In-process. | None supplied beyond blank photographed forms | Open - Evidence |
| GAP-007 | P2 | Actual dimensions and drawing tolerances are incompletely controlled | Required: approved drawing and specs + calibrated measurements in completed records. Route: In-process. | None supplied beyond blank photographed forms | Open - Evidence |
| GAP-008 | P2 | Breathing-system filter performance record not linked (conditional) | Required: valid NaCl test report (original/conditioned), manufacturer sample rationale, approved performance specification and scope-to-product mapping. Route: qualification / periodic / supplier per approved plan. | None supplied beyond blank photographed forms | Open - Evidence |
| GAP-009 | P2 | HME applicability and verification are not demonstrated (conditional) | Required: approved intended use, applicable product classification, HME qualification records and approved labeling/packaging links if applicable; N/A decision with rationale otherwise. Route: qualification / labeling / QA. | None supplied beyond blank photographed forms | Open - Evidence |
| GAP-010 | P2 | Product variants and N/A decisions not controlled | Required: approved variant matrix/BOM and filled records with correct applicable checks. Route: In-process / changeover. | None supplied beyond blank photographed forms | Open - Evidence |
| GAP-011 | P2 | Time-box check marks do not preserve objective measurements | Required: approved traceable form/worksheet and executed controlled sample records; reviewed, with instrument validity verified. Route: In-process. | None supplied beyond blank photographed forms | Open - Approval |
| GAP-012 | P3 | Non-measurable and inconsistent visual language | Required: approved visual defect catalogue, training evidence where needed, and traceable checks. Route: In-process. | None supplied beyond blank photographed forms | Open - Technical Input |
| GAP-013 | P2 | Referenced specifications and revision validity unverified | Required: document master list, controlled applicable WI/specs/drawings and effective-date proof. Route: document control. | None supplied beyond blank photographed forms | Open - Evidence |

# 3. Closure rules

• Verified Closed requires the approved controlled source, applicability confirmation and applicable executed objective evidence actually verified, with appropriate QA/QC authorization.

• Documentary Covered - Operational Evidence Pending is allowed only after an approved, source-verified requirement and complete corresponding revised record exist. Current unapproved drafts do NOT meet this condition.

• Not Applicable with Justification requires approved technical scope and intended-use evidence, not simply a reviewer choice.

• No ISO compliance percentage or 100% declaration is supportable from the submitted collection; product-family scope and effective controlled references are incomplete.

• For ISO 23328-1, test results require appropriate conditioned/unconditioned report with method parameters and an approved device-specific acceptance claim; do not invent a universal minimum efficiency.

• For ISO 9360-1, measured/disclosed HME performance parameters do not inherently imply a universal standard-imposed limit. Assess each parameter against an approved intended-use specification, where appropriate.

# 4. Source register

SRC-01 | BS EN ISO 9360-1:2009 (ISO 9360-1:2000), 30 PDF pages | Applies only if a supplied device is an HME in the scope of Clause 1; not to all breathing circuit components.

SRC-02 | BS EN ISO 23328-1:2008 (ISO 23328-1:2003), 22 PDF pages | Salt-aerosol filtration test method for breathing system filters; no universal minimum efficiency specified by Part 1.

SRC-03 | WI-8-2-2-26 Rev. 5, Sterile Amnion Hook, 01 Sep 2023, 4 pages | Different product; excluded from breathing circuit WI comparison; separate limited observation only.

SRC-04 | In-process control photo, corrugated breathing circuits and accessories, PDF photo p. 1 | Rev. 4, 07 Oct 2025; form number cropped / not reliably legible; 17 characteristics.

SRC-05 | F-823-i49 Rev. 2, In-process Control for Swivel Connector, photo p. 2 | Dated 01 Mar 2024; 10 characteristics.

SRC-06 | F-823-i47 Rev. 6, In-process Control for Breathing Filter Assembly, photo p. 3 | Dated 28 Apr 2026; 15 characteristics.

SRC-07 | F-823-i39 Rev. 5, In-process Control for Catheter Mount, photo p. 4 | Dated 30 May 2024; 11 characteristics.