# REVISED IN-PROCESS CONTROL FORMS

Four proposed controlled templates • NOT final inspection/release

DRAFT FOR TECHNICAL REVIEW • Evidence-based assessment • 15 October 2026

Document status: Not approved. No evidence of actual product conformity or operational closure was supplied.

Every row retains the observed control; additions request measurable results, justified N/A, sample identity and approved method. Fields labelled OPEN TECHNICAL INPUT REQUIRED must be completed by Engineering/QC/QA before controlled issuance.

# A. Corrugated Breathing Circuits & Accessories

SRC-04 | photographed p.1 | Rev.4 dated 07 Oct 2025 | form number NOT fully legible

| Identification | Entry / controlled reference |
| --- | --- |
| Product code / variant | ____________________________________________________________ |
| Lot / order / date / shift / machine | ____________________________________________________________ |
| Source WI / spec / BOM & revisions | WI-8-2-2-44-1 Rev.__ / TF-44-PS Rev.__ / TF-44-PDBM Rev.__ / controlled verification __ |
| Sampling plan / frequency / trigger | ____________________________________________________________ |
| Inspector / equipment / calibration / HOLD-NCR link | ____________________________________________________________ |

Sampling / time-grid: each designated check must include actual clock time, lot, sample quantity and result. In the fields below enter numeric values for measurable characteristics or a defined defect outcome for visual checks; attach worksheet if more space is required. Do not enter unapproved numeric acceptance limits.

| # | Characteristic | Source-form requirement / qualification | Actual test entry | Decision / trace |
| --- | --- | --- | --- | --- |
| 1 | Corrugated Tubing | Good appearance; no holes, black spots, sink marks or fading. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 2 | Bending tube | Flexible, no damage/crack/blockage. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 3 | Tube ribs | Smooth, without roughness. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 4 | Tube length | Source selections: 150 / 160 / 180 / 200 / 250 cm. Controlled tolerance pending. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 5 | Elbow connector with end cap | Good fit; no cracks or holes. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 6 | Elbow-to-Y fitting | Good fit; not loose. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 7 | Gluing tube + adaptor | Strong bond; no pull-out. Quantitative criterion pending. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 8 | Y connector parts + caps | Correct fit; no damage or cracks. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 9 | Mask condition | Good fit; inflatable valve intact, no damage/leak. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 10 | Adaptor fitting with Y connector | Tight; no looseness, damage or cracks. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 11 | HME filter status | Smooth; sealed, free of cracks, damage, black spots and particles (HME applicability pending). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 12 | HME filter size | Neonate / pediatric / adult selection; controlled specification pending. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 13 | Breathing bag status | Clear silicone bag; connector fit; no holes/leak; source color light green. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 14 | Breathing bag size | According to controlled specification (pending). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 15 | Standard male connector status | Smooth; source connector 15/22 mm; correct fit; free of black spots/cracks/damage. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 16 | CO2 sampling line / connector | Clear; dimension to approved drawing; glued, no pull-out, leak or black spots. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 17 | Leak test | Source: "Should be no Leakage". Test condition/limit must be approved before execution. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |

## Checkpoint sign-off and disposition

Checked by QC: __________________   Date / time: ____________  Reviewed by QC Manager/authorized reviewer: __________________   Date: ____________

# B. Swivel Connector

F-823-i49 | photographed p.2 | Rev.2 dated 01 Mar 2024

| Identification | Entry / controlled reference |
| --- | --- |
| Product code / variant | ____________________________________________________________ |
| Lot / order / date / shift / machine | ____________________________________________________________ |
| Source WI / spec / BOM & revisions | WI-8-2-2-44-1 Rev.__ / TF-44-PS Rev.__ / TF-44-PDBM Rev.__ / controlled verification __ |
| Sampling plan / frequency / trigger | ____________________________________________________________ |
| Inspector / equipment / calibration / HOLD-NCR link | ____________________________________________________________ |

Sampling / time-grid: each designated check must include actual clock time, lot, sample quantity and result. In the fields below enter numeric values for measurable characteristics or a defined defect outcome for visual checks; attach worksheet if more space is required. Do not enter unapproved numeric acceptance limits.

| # | Characteristic | Source-form requirement / qualification | Actual test entry | Decision / trace |
| --- | --- | --- | --- | --- |
| 1 | Clarity/appearance | Clear, good appearance (approved workmanship catalogue required). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 2 | Swivel condition | No cracks/damage; smooth surface, not defective. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 3 | Surface irregularities | No flashing, short shot, black spot or white spot. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 4 | Swivel type | Disposable / reusable - select applicable type. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 5 | Connector diameter / fit | Source: 22 mm connector for endotracheal adaptor; fit to drawing. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 6 | Suction connector | Source: 15 mm connector; fitting to applicable adaptor ID/OD per specification. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 7 | Rotation status | Smooth rotation; no sharp edge/flash (functional method pending). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 8 | O-ring | Black O-ring; smooth movement; no leakage. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 9 | Connector with cap | Elastomeric cap; correct fit; OD/ID per approved specification. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 10 | Condition of product | All parts conform to approved condition/defect catalogue. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |

## Checkpoint sign-off and disposition

Checked by QC: __________________   Date / time: ____________  Reviewed by QC Manager/authorized reviewer: __________________   Date: ____________

# C. Breathing Filter Assembly

F-823-i47 | photographed p.3 | Rev.6 dated 28 Apr 2026

| Identification | Entry / controlled reference |
| --- | --- |
| Product code / variant | ____________________________________________________________ |
| Lot / order / date / shift / machine | ____________________________________________________________ |
| Source WI / spec / BOM & revisions | WI-8-2-2-44-1 Rev.__ / TF-44-PS Rev.__ / TF-44-PDBM Rev.__ / controlled verification __ |
| Sampling plan / frequency / trigger | ____________________________________________________________ |
| Inspector / equipment / calibration / HOLD-NCR link | ____________________________________________________________ |

Sampling / time-grid: each designated check must include actual clock time, lot, sample quantity and result. In the fields below enter numeric values for measurable characteristics or a defined defect outcome for visual checks; attach worksheet if more space is required. Do not enter unapproved numeric acceptance limits.

| # | Characteristic | Source-form requirement / qualification | Actual test entry | Decision / trace |
| --- | --- | --- | --- | --- |
| 1 | Clarity/appearance | Clear; white colour. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 2 | Lower housing condition | No cracks/damage; correct fit for endotracheal connector. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 3 | Upper housing condition | No cracks/damage; correct fit for breathing-circuit adaptor. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 4 | Adult bacterial filter media | Nonwoven diameter 59±1 mm (source-form value; verify drawing). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 5 | Adult humidifying paper | Diameter 53±1 mm; width 15±1 mm (source-form values). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 6 | Pediatric bacterial filter media | Nonwoven diameter 44±1 mm (source-form value). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 7 | Pediatric humidifying paper | Diameter 30±1 mm; width 15±1 mm (source-form values). | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 8 | Housing deformation at filter position | No defect at filter seat. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 9 | Humidifying paper position | Correctly positioned; does not move after fit. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 10 | Top nonwoven filter condition | Correctly positioned; does not move after fit. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 11 | Fitting status | Correct fit / correct position. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 12 | Appearance after sealing | No damage on any side; approved seal appearance. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 13 | Sharp edges after sealing | No sharp edges; smooth finish. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 14 | Under-water seal leak test | No leak from sealing side under controlled, approved pressure/time/sample plan. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 15 | End plug status | Correct fit and twist/engagement. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |

## Checkpoint sign-off and disposition

Checked by QC: __________________   Date / time: ____________  Reviewed by QC Manager/authorized reviewer: __________________   Date: ____________

# D. Catheter Mount

F-823-i39 | photographed p.4 | Rev.5 dated 30 May 2024

| Identification | Entry / controlled reference |
| --- | --- |
| Product code / variant | ____________________________________________________________ |
| Lot / order / date / shift / machine | ____________________________________________________________ |
| Source WI / spec / BOM & revisions | WI-8-2-2-44-1 Rev.__ / TF-44-PS Rev.__ / TF-44-PDBM Rev.__ / controlled verification __ |
| Sampling plan / frequency / trigger | ____________________________________________________________ |
| Inspector / equipment / calibration / HOLD-NCR link | ____________________________________________________________ |

Sampling / time-grid: each designated check must include actual clock time, lot, sample quantity and result. In the fields below enter numeric values for measurable characteristics or a defined defect outcome for visual checks; attach worksheet if more space is required. Do not enter unapproved numeric acceptance limits.

| # | Characteristic | Source-form requirement / qualification | Actual test entry | Decision / trace |
| --- | --- | --- | --- | --- |
| 1 | Clarity/appearance | Clear, good appearance. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 2 | Irregularities | No short shot, flashing or cracks. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 3 | S-bore / extendable / corrugated tube | Smooth; no holes/leaks/flash. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 4 | Swivel adaptor / ET tube connector | Good fitting; smooth rotation. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 5 | Swivel adaptor / Y connector | Good fit and rotation; no pull-out. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 6 | Y connector / outer straight connector | Good fit; smooth rotation. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 7 | Extensible tube at both connectors | Good fit / strong glue; no pull-out or crack. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 8 | Y swivel connector port cap | Tight; no leak or break. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 9 | Tube bonding / 15-22 mm connector | Strong bond; no pull-out/crack/damage. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 10 | Tube bonding / swivel connector | Strong bond; no pull-out/crack/damage. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |
| 11 | Gluing with connectors | Strong bond; no pull-out. | Time __ Sample ID __ Actual / Unit __ Instrument / Method __ | P / F / Pending / N/A N/A reason __ Evidence/NCR __ |

## Checkpoint sign-off and disposition

Checked by QC: __________________   Date / time: ____________  Reviewed by QC Manager/authorized reviewer: __________________   Date: ____________

# E. Non-routine verification evidence link (not routine in-process measurements)

| Evidence route | Applicability to product | Approved reference / qualification scope | Review outcome |
| --- | --- | --- | --- |
| ISO 23328-1 NaCl BSF test | BSF Y / N / Pending; selected adult/pediatric __ | Report ID/rev __; test & conditioner __; n __; agreed efficiency criterion __ | Verified / Pending / N/A reason __ |
| ISO 9360-1 HME qualification | HME Y / N / Pending; VT ≥250 mL intended __ | Test package (moisture loss / pressure drop / leakage / compliance) __; labeling __ | Verified / Pending / N/A reason __ |
| Approved seal/leak/retention methods | Applies per configuration/controlled plan | Method ID/rev __; pressure/time/limit/n __; equipment __ | Verified / Pending __ |
| Document control and BOM | Every affected in-process product | Effective WI / drawing / BOM / form revisions __ | Verified / Pending __ |