SAUDI MAIS  /  QUALITY CONTROL

ISO CLOSURE / OPEN-INPUT REGISTER

Requirement-by-requirement evidence status, not a simulated compliance certificate

| Product family | HME + breathing system filters (circuit controls noted separately) |
| --- | --- |
| Standards | BS EN ISO 9360-1:2009 (identical to ISO 9360-1:2000); BS EN ISO 23328-1:2008 (identical to ISO 23328-1:2003) |
| Related WI | WI-8-2-2-44 Rev. 8 (21 Oct 2025) |
| Finding count | 22 gaps; no verified execution evidence supplied |
| Status | OPEN — documentary proposal, not operational closure |

DRAFT FOR QC/QA REVIEW  •  Evidence not supplied shall not be presumed to exist.

# Closure rules

Verified Closed requires approved source control, test/inspection evidence, correct configuration/lot applicability, completed review/approval and verified results. Documentary Covered is permitted only once controlled document revisions have been approved; these four deliverables are unapproved proposals. “Open — Evidence” means actual executed source data absent; “Open — Technical Input” means acceptance/configuration/method unresolved. N/A must have product-specific documented justification.

# ISO clauses / requirement coverage

| Req ID | ISO clause/page | What must be shown | Linked gap(s) | Evidence record requested | What was verified here | Operational status |
| --- | --- | --- | --- | --- | --- | --- |
| H01 | ISO 9360-1 §1, PDF11/ISO1 | Scope ≥250 mL | GAP-019 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Conditional — configuration missing |
| H02 | §§5.1–5.2, PDF12–13/ISO2–3 | Port geometry & non-intermateable accessory port | GAP-005 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Partly described by WI; evidence open |
| H03 | §5.3, PDF13/ISO3 | Sterile package per ISO 11607 if sterile | GAP-009 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Open; supporting ISO not provided |
| H04 | §6.1–6.2, PDF13,21–23/ISO3,11–13 | Moisture-loss test/conditions | GAP-001 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| H05 | §6.3/Table4, PDF23,25/ISO13,15 | Pressure drop pre/post use | GAP-002 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| H06 | §6.4, PDF23/ISO13 | Gas leakage at 7±0.5kPa in mL/min | GAP-003 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated by bubble test |
| H07 | §6.5, PDF23,25/ISO13,15 | Compliance in mL/kPa | GAP-004 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| H08 | §§7.1–7.4, PDF25–26/ISO15–16 | Marking, IFU, reported characteristics | GAP-008 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| F01 | ISO 23328-1 §§1–2, PDF9/ISO1 | Breathing-system filter scope | GAP-019 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Conditional |
| F02 | §§3.1–3.4, PDF10/ISO2 | NaCl test conditions/apparatus/conditioning | GAP-006 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| F03 | §§3.5–3.6, PDF11/ISO3 | Sample rationale; controlled test | GAP-006/007 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| F04 | §4/Table1, PDF12/ISO4 | PV/efficiency and adult/pediatric flows | GAP-006 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| F05 | §5, PDF12/ISO4 | Required BSF test-report information | GAP-007 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |
| F06 | Annex A, PDF13–15/ISO5–7 | Normative clinical-use conditioning | GAP-006 | Approved test/IFU/product document + lot/configuration/revision linkage | No executed evidence submitted | Not demonstrated |

# Detailed Gap Closure Register

| Gap ID | P | Issue | Control route | Required evidence | Actual verified evidence | Status | Next closure gate |
| --- | --- | --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 | Moisture-loss performance not traceably established | Qualification / product file | Approved moisture-loss qualification reports incl. Table 2 conditions/calibration checks and results; approved specification/IFU; design and lot linkage; change-control and periodic review criteria. Closure requires confirmed applicable reports and approval, not blank templates. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-002 | P1 | Pressure-drop performance and declared results missing | Qualification / periodic | Approved pre/post-use pressure-drop report, apparatus calibration, device-specific approved limits/claims, change control and link to released design. | NOT PROVIDED; proposed form/WI only | Open — Evidence / Approval | Technical approval and source evidence verification pending |
| GAP-003 | P1 | Underwater bubble check cannot replace quantitative HME leakage | Qualification + routine in-process | Approved quantitative ISO §6.4 pre/post-use report and approved device acceptance; approved validated underwater method and batch records with actual results, sampling and disposition. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-004 | P2 | HME compliance and internal-volume information not traceable | Qualification / product file | Approved test calculation and drawings, compliance test data with corrections, product labeling/IFU, design-revision linkage. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-005 | P1 | HME conical connector conformance not demonstrated | Incoming/first article + process + qualification | Approved ISO 5356-1 conformance assessment, controlled drawings/port configurations, gauge qualifications/calibration and lot sampling records. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-006 | P1 | Breathing filter filtration efficiency unverified against salt test | Qualification / external laboratory | Actual NaCl challenge reports with unused/conditioned results and test setup data; documented flow and aerosol; approved efficiency claims; approved controlled route and design/lot traceability. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-007 | P2 | Filter challenge sampling and formal ISO test-report content undefined | Qualification / documentation | Signed manufacturer sampling rationale; final controlled ISO test report and reports register; evidence scope/revision match. | NOT PROVIDED; proposed form/WI only | Open — Evidence / Approval | Technical approval and source evidence verification pending |
| GAP-008 | P2 | HME IFU/marking and declared performance checks incomplete | Packaging / labeling / controlled record | Approved IFU, label artwork, labeling verification, declarations supported by reports; documentation of conditional applicability. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-009 | P1 | Sterile HME packaging validation evidence not established | Packaging validation / sterilization / final release | Approved ISO 11607 and sterilization validations, packaging inspection/seal strength/visual records as applicable, traceable sterilization/release evidence. No operational closure without these. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-010 | P1 | Ultrasonic weld control absent from in-process HME form | Routine in-process / setup | Approved validated weld window, signed setup and process readings, equipment calibration/qualification, defects and containment records. | NOT PROVIDED; proposed form/WI only | Open — Approval / Evidence | Technical approval and source evidence verification pending |
| GAP-011 | P2 | HME media-width acceptance differs between WI and form | Engineering specification / routine process | Approved product drawings, engineering change assessment resolving discrepancy, revised controlled WI/form, measured in-process evidence. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Approval | Technical approval and source evidence verification pending |
| GAP-012 | P2 | Underwater HME in-process test method and AQL data incomplete | Routine in-process | Approved method validation/sampling plan, filled sample and lot decision records, equipment/test setup traceability and segregation if failures. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-013 | P1 | Breathing circuit leak-test parameters conflict internally | Routine process + method qualification | Controlled reconciled WI and machine WIs, validated leak thresholds per design, calibrated tester logs, actual sample test and approved disposition. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Approval | Technical approval and source evidence verification pending |
| GAP-014 | P1 | Joint pull-off acceptance is not measurable in circuit form | Routine in-process / product test | Approved criterion and source confirmation, signed force-test readings, calibrated load cell, sampling basis, in-process dispositions. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-015 | P2 | Circuit lengths conflict and tolerances are missing | Engineering / routine in-process | Approved drawings/item master for all variants, reconciled WI/form, measurement records by lot. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Approval | Technical approval and source evidence verification pending |
| GAP-016 | P2 | Dimensional checks in WI omitted from circuit form | Routine in-process | Approved dimensions by configuration, calibrated instruments, actual measurement sheets and controlled engineering rationale for N.A. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-017 | P2 | Hourly form grid versus required checking frequency unresolved | Routine in-process / QC plan | Approved per-operation sampling/control plan, recorded checks with timestamps and deviation handling. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Approval | Technical approval and source evidence verification pending |
| GAP-018 | P2 | Circuit HME status lacks evidence and configuration linkage | Routine in-process / material traceability | Completed HME source in-process records and lot traceability; approved configuration/qualification matrix; circuit assembly observations. | NOT PROVIDED; proposed form/WI only | Open — Evidence / Approval | Technical approval and source evidence verification pending |
| GAP-019 | P1 | Product size claims exceed/require ISO applicability determination | Design / regulatory / qualification | Approved product intended-use statement/IFU, applicability matrix, correct verification evidence for each age/tidal-volume design, clinical/Regulatory approval. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-020 | P2 | Form reference numbers conflict with supplied WI | Document control | Document-control register and approved change/revision history; verification of WI-8-2-2-44-1 existence, controlled source for photo F-823-i39. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Approval | Technical approval and source evidence verification pending |
| GAP-021 | P2 | Catheter-mount numeric characteristics absent from photographed form | Routine in-process / controlled-copy verification | Current controlled F-823-i39 source, released dimensional/tensile specifications, calibrated measurement records and change approval. | NOT PROVIDED; proposed form/WI only | Open — Technical Input / Evidence | Technical approval and source evidence verification pending |
| GAP-022 | P2 | Binary checks lack accountable sample/results and nonconformance trail | Routine in-process / QA records | Completed signed process logs with actual values, sampled quantities, NCR/hold records and traceability. Blank redesigned forms remain documentary drafts only. | NOT PROVIDED; proposed form/WI only | Open — Approval / Evidence | Technical approval and source evidence verification pending |

# Manager decision and source access

1. Obtain product-code/lot/size map; signed drawings and claims; authoritative forms. 2. Reconcile WI/form inconsistent process settings. 3. Assign QA/QC/Engineering responsibility for qualification evidence and label validation. 4. Verify source ISO 5356-1, ISO 11607, BS EN 5367 and machine-specific WI before technical closure. 5. Reassess revision after approvals and verified implementation.

No ISO verification is marked closed on this evidence set. Source editions assessed are those supplied; current required regulatory editions were not researched or verified.