SAUDI MAIS  /  QUALITY CONTROL

REVISED IN-PROCESS QUALITY CONTROL FORMS

Breathing Circuit + HME Filter + Catheter Mount | NOT Inspection & Test Reports

| Origin | F-823-i8 Rev. 4 (07 Oct 2025); F-823-i47 Rev. 6 (28 Apr 2026); F-823-i39 Rev. 5 (30 May 2024; photographed copy) |
| --- | --- |
| Stage | IN-PROCESS ASSEMBLY CONTROL ONLY |
| WI linkage | WI-8-2-2-44 (proposed revision; parent references pending check) |
| Status | Editable master templates; unapproved / not executed |

DRAFT FOR QC/QA REVIEW  •  Evidence not supplied shall not be presumed to exist.

# Record-use instructions

Use A=acceptable, X=not acceptable, R=see remarks, N/A=not applicable (record rationale). Enter actual time and the numeric result for all measurable characteristics in the supplemental numeric log. Retain the original form’s hourly/two-hour headings as history; exact required frequency and sampling plan must be approved per the WI §13. A completed in-process record does not constitute final inspection, release or ISO qualification proof.

In all templates, enter Product/Code/Type/Size/Revision, production lot, manufacturing order, date/shift, check time, sample count, QC name and instrument IDs. Use the separate attached test/results log for repeat readings and any failed units. No blanks may be treated as acceptance.

# FORM A — In-process Control: HME / Breathing Filter Assembly

Baseline: F-823-i47 Rev. 6 (28 Apr 2026)   |   DRAFT CONTROL STATUS: not approved   |   Gap links: 001–012, 017–020, 022

| Product code / configuration | Lot / manufacturing order | Date / shift | Parent WI rev | QC control plan / revision |
| --- | --- | --- | --- | --- |
| __________________ | __________________ | __________________ | PENDING | __________________ |

Allowed result codes: A / X / R / N/A. Fill numeric results in supplemental log. Any X -> hold affected output, record NCR and escalation; no automatic release.

## Characteristics (original rows 1–15 retained; additions 16–17)

| # | Characteristic | Acceptance / source or open field | Method / equipment | Frequency / evidence route | Actual check / disposition |
| --- | --- | --- | --- | --- | --- |
| 1 | Clarity / appearance | Clear / white colour; no visible foreign matter | Visual | Per approved assembly plan | Time ___ / Result ___ / QC ___ |
| 2 | Lower housing | No cracks/damage; ETT connector fit per approved design | Visual / fit / gauge ref | Per approved plan | Time ___ / Result ___ / QC ___ |
| 3 | Upper housing cover | No cracks/damage; circuit-adaptor fit | Visual / fit / gauge ref | Per approved plan | Time ___ / Result ___ / QC ___ |
| 4 | Adult bacterial filter media | Ø59 ±1 mm, correct approved variant | Caliper (ID logged) | During applicable assembly | Time ___ / Result ___ / QC ___ |
| 5 | Adult humidifying paper | Ø53 ±1 mm; width 15 mm nominal / 15±1 mm conflict — HOLD for approved spec | Caliper; actual logged | During applicable assembly | Time ___ / Result ___ / QC ___ |
| 6 | Pediatric bacterial filter media | Ø44 ±1 mm, correct approved variant | Caliper | During applicable assembly | Time ___ / Result ___ / QC ___ |
| 7 | Pediatric humidifying paper | Ø30 ±1 mm; width conflict pending drawing | Caliper; actual logged | During applicable assembly | Time ___ / Result ___ / QC ___ |
| 8 | Housing deformation | No visible deformation affecting fit/seal | Visual | Per approved plan | Time ___ / Result ___ / QC ___ |
| 9 | Humidifying paper positioning | Correct position; no movement after fixation | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 10 | Top nonwoven filter positioning | Correct position; no movement after fixation | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 11 | Assembly / seal fit | Correct configuration, fit, and position | Visual / functional | Per approved plan | Time ___ / Result ___ / QC ___ |
| 12 | Post-weld appearance | No visible damage or defective sealing | Visual | Per approved plan | Time ___ / Result ___ / QC ___ |
| 13 | Post-weld edges | Smooth; no sharp edges | Visual/tactile | Per approved plan | Time ___ / Result ___ / QC ___ |
| 14 | Post-weld underwater bubble screen | No bubbles under controlled 60mL syringe/10s method, pressure validation OPEN | Validated underwater setup | AQL/control plan pending | Time ___ / Result ___ / QC ___ |
| 15 | End plug / cap | Proper closure, twisting, and fit; no damage | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 16 | Ultrasonic welder settings | Time 0.5–1.5s; delay 0.5–1.5s; hold 0.5–1.5s; pressure 5.5–6.5kg/cm²; validated recipe required | Machine readout / parameter log | At setup, changes and approved frequency | Time ___ / Result ___ / QC ___ |
| 17 | Component lot and intended variant | Match released BOM; HME family adult/ped verified | BOM/production order | Each setup / change | Time ___ / Result ___ / QC ___ |

## Recurring in-process shift check grid — enter A / X / R / N/A

Use the defined approved control-plan frequency; these printed hour slots represent available shift windows, not compulsory hourly testing. Enter actual time in check log and numeric actual readings in the attached results log. All X values require NCR / affected-unit reference.

| # | Characteristic | 7 | 9 | 11 | 1 | 3 | 5 | 7 | 9 | 11 | 1 | 3 | 5 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Clarity / appearance | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 2 | Lower housing | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 3 | Upper housing cover | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 4 | Adult bacterial filter media | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 5 | Adult humidifying paper | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 6 | Pediatric bacterial filter media | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 7 | Pediatric humidifying paper | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 8 | Housing deformation | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 9 | Humidifying paper positioning | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 10 | Top nonwoven filter positioning | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 11 | Assembly / seal fit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 12 | Post-weld appearance | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 13 | Post-weld edges | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 14 | Post-weld underwater bubble screen | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 15 | End plug / cap | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 16 | Ultrasonic welder settings | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 17 | Component lot and intended variant | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |

F-823-i47: Actual required frequency (approved plan) ___; sample quantity per check ___; shift start/end ___; unperformed required checks / reason ___; checker initials on linked numeric log ___.

## HME welded process and leak test actual result log

| Time | Sample / unit IDs | Machine / recipe | Weld s | Delay s | Hold s | Pressure kg/cm² | Leak n / fail n | NCR/hold | QC initials |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |

Qualification records are NOT routinely entered as measured ISO results. Reference only: moisture-loss report ___; pressure-drop report ___; ISO 9360 quantitative leak report ___; compliance/internal volume ___; ISO 23328 NaCl filtration ___; connector conformance ___; approved configuration/IFU ___; checked by ___; date ___. OPEN input if unavailable.

Remarks / NCR / affected lots: _______________________________________________________________________    QC checked by: _____________   Reviewed by: _____________

# FORM B — In-process Control: Corrugated Breathing Circuit & Accessories

Baseline: F-823-i8 Rev. 4 (07 Oct 2025)   |   DRAFT CONTROL STATUS: not approved   |   Gap links: 013–018, 020, 022

| Product code / configuration | Lot / manufacturing order | Date / shift | Parent WI rev | QC control plan / revision |
| --- | --- | --- | --- | --- |
| __________________ | __________________ | __________________ | PENDING | __________________ |

Allowed result codes: A / X / R / N/A. Fill numeric results in supplemental log. Any X -> hold affected output, record NCR and escalation; no automatic release.

## Characteristics (original rows 1–17 retained; additions 18–22)

| # | Characteristic | Acceptance / source or open field | Method / equipment | Frequency / evidence route | Actual check / disposition |
| --- | --- | --- | --- | --- | --- |
| 1 | Corrugated tubing | No holes/spotting/defects/sink marks/fade | Visual | Per approved plan | Time ___ / Result ___ / QC ___ |
| 2 | Tube bending | Flexible, no damage/cracks/blockage | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 3 | Tube ribs quality | Smooth, free of roughness | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 4 | Tube length | Source WI 160/180/200 cm; form also 150/250 cm: design configuration/tolerance approval pending | Scale / actual value | Sample per approved plan | Time ___ / Result ___ / QC ___ |
| 5 | Elbow & cap | Sound fit; no cracks/holes | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 6 | Elbow–Y connector fit | No looseness/damage | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 7 | Tube–adaptor glue | Load 45 ±2 N; speed 50 ±5 mm/min; hold 1 min, ≥150mm position; no separation; source confirmation pending | Universal/Cometech tester | Per validated sample plan | Time ___ / Result ___ / QC ___ |
| 8 | Y-parts & caps | Correct fitting; undamaged | Physical/visual | Per approved plan | Time ___ / Result ___ / QC ___ |
| 9 | Mask fit | No damage/leak; valve checks as product specification | Visual/physical | If installed | Time ___ / Result ___ / QC ___ |
| 10 | Y-adaptor fit | Tight; no cracks/damage | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 11 | HME filter status | Surface intact, sealed, no foreign matter + F-823-i47 ref | Visual / record | If installed | Time ___ / Result ___ / QC ___ |
| 12 | HME size | Adult/Pediatric/Neonate only if approved intended-use/design scope | Order + evidence matrix | If installed | Time ___ / Result ___ / QC ___ |
| 13 | Breathing bag fit | As approved bag specification, leak-free/no damage | Visual/physical | If installed | Time ___ / Result ___ / QC ___ |
| 14 | Breathing bag size | Approved product-specific specification required | Order/drawing | If installed | Time ___ / Result ___ / QC ___ |
| 15 | Standard male connector | Correct Ø15/22 geometry per applicable drawing/gauge | Caliper / go-no-go | Per approved plan | Time ___ / Result ___ / QC ___ |
| 16 | CO2 sampling line | Correct drawing, secure adhesive, no leak | Visual/physical | If installed | Time ___ / Result ___ / QC ___ |
| 17 | Circuit leak test | Use qualified machine/design method only; source WI conflicting 6kPa vs 18–32kPa | Leak machine + actual log | Per approved plan | Time ___ / Result ___ / QC ___ |
| 18 | Tube dimensions | ID/OD nominal adult 22×25, pediatric 15×19, neonate 10×15 mm; tolerances pending | Caliper / actual log | Per drawing/applicable | Time ___ / Result ___ / QC ___ |
| 19 | Tube sheet thickness | 0.55–0.70 mm if this tube type; applicability check | Caliper / actual log | Per drawing/applicable | Time ___ / Result ___ / QC ___ |
| 20 | Rib distance | 5.8–6.2 mm if applicable | Caliper / actual log | Per drawing/applicable | Time ___ / Result ___ / QC ___ |
| 21 | Rib thickness | 2.3–2.5 mm if applicable | Caliper / actual log | Per drawing/applicable | Time ___ / Result ___ / QC ___ |
| 22 | HME lot / qualification linkage | Installed unit code/lot and accepted HME assembly/qualification ref | Traceability verification | If HME installed | Time ___ / Result ___ / QC ___ |

## Recurring in-process shift check grid — enter A / X / R / N/A

Use the defined approved control-plan frequency; these printed hour slots represent available shift windows, not compulsory hourly testing. Enter actual time in check log and numeric actual readings in the attached results log. All X values require NCR / affected-unit reference.

| # | Characteristic | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Corrugated tubing | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 2 | Tube bending | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 3 | Tube ribs quality | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 4 | Tube length | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 5 | Elbow & cap | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 6 | Elbow–Y connector fit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 7 | Tube–adaptor glue | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 8 | Y-parts & caps | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 9 | Mask fit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 10 | Y-adaptor fit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 11 | HME filter status | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 12 | HME size | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 13 | Breathing bag fit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 14 | Breathing bag size | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 15 | Standard male connector | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 16 | CO2 sampling line | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 17 | Circuit leak test | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 18 | Tube dimensions | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 19 | Tube sheet thickness | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 20 | Rib distance | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 21 | Rib thickness | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 22 | HME lot / qualification linkage | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |

F-823-i8: Actual required frequency (approved plan) ___; sample quantity per check ___; shift start/end ___; unperformed required checks / reason ___; checker initials on linked numeric log ___.

## Circuit numeric results and device traceability

| Check time | Sample ID | Item / type | Dimension result & unit | Device/Method ID | Leak pressure / hold / result | Pull test N / speed / hold | P/F / NCR |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |

HME included? Y/N ___; part code ___; HME lot ___; F-823-i47 accepted record ___; filtration / HME qualification applicability check ___; reviewed by _____.

Remarks / NCR / affected lots: _______________________________________________________________________    QC checked by: _____________   Reviewed by: _____________

# FORM C — In-process Control: Catheter Mount (PROPOSED)

Baseline: F-823-i39 Rev. 5 (30 May 2024; photographed copy)   |   DRAFT CONTROL STATUS: not approved   |   Gap links: 020–022

| Product code / configuration | Lot / manufacturing order | Date / shift | Parent WI rev | QC control plan / revision |
| --- | --- | --- | --- | --- |
| __________________ | __________________ | __________________ | PENDING | __________________ |

Allowed result codes: A / X / R / N/A. Fill numeric results in supplemental log. Any X -> hold affected output, record NCR and escalation; no automatic release.

## Characteristics (original photographed rows 1–11 retained; added measurements 12–15)

| # | Characteristic | Acceptance / source or open field | Method / equipment | Frequency / evidence route | Actual check / disposition |
| --- | --- | --- | --- | --- | --- |
| 1 | Appearance | Clear / acceptable per approved drawing | Visual | Per approved plan | Time ___ / Result ___ / QC ___ |
| 2 | Irregularities | No short shots, flashing or cracks | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 3 | Tube condition | Smooth; no holes, leakage or flashing | Visual/physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 4 | Swivel–ETT fitting | Good fitting and smooth rotation | Physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 5 | Swivel–Y fitting | Good fitting, rotatable, no pulling off | Physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 6 | Y–outer connector | Good fitting and smooth rotation | Physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 7 | Extendable tube glued both sides | Strong bond, no detachment or crack | Physical / approved tester | Per approved plan | Time ___ / Result ___ / QC ___ |
| 8 | Y port cap fit | Tight, no leak/breakage | Physical | Per approved plan | Time ___ / Result ___ / QC ___ |
| 9 | Tube glued with 15/22 connector | Strong glue, no separation, damage or cracks | Physical / approved tester | Per approved plan | Time ___ / Result ___ / QC ___ |
| 10 | Tube glued with swivel | Strong glue, no separation, damage or cracks | Physical / approved tester | Per approved plan | Time ___ / Result ___ / QC ___ |
| 11 | All connector gluing status | Approved product-specific strength; no separation | Approved tester | Per approved plan | Time ___ / Result ___ / QC ___ |
| 12 | Tube length — ADDED | 15cm nominal; approved tolerance pending | Scale / actual log | Approved plan | Time ___ / Result ___ / QC ___ |
| 13 | Tube connector ID×OD — ADDED | Nominal 15×22mm; tolerance pending | Caliper / actual log | Approved plan | Time ___ / Result ___ / QC ___ |
| 14 | Connector ID×OD — ADDED | Nominal 15×22mm; tolerance pending | Caliper / actual log | Approved plan | Time ___ / Result ___ / QC ___ |
| 15 | Glue tensile value — ADDED | Validated product-specific N, speed and duration pending; do NOT borrow circuit force | Universal test machine | Approved plan | Time ___ / Result ___ / QC ___ |

## Recurring in-process shift check grid — enter A / X / R / N/A

Use the defined approved control-plan frequency; these printed hour slots represent available shift windows, not compulsory hourly testing. Enter actual time in check log and numeric actual readings in the attached results log. All X values require NCR / affected-unit reference.

| # | Characteristic | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Appearance | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 2 | Irregularities | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 3 | Tube condition | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 4 | Swivel–ETT fitting | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 5 | Swivel–Y fitting | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 6 | Y–outer connector | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 7 | Extendable tube glued both sides | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 8 | Y port cap fit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 9 | Tube glued with 15/22 connector | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 10 | Tube glued with swivel | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 11 | All connector gluing status | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 12 | Tube length — ADDED | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 13 | Tube connector ID×OD — ADDED | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 14 | Connector ID×OD — ADDED | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| 15 | Glue tensile value — ADDED | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |

F-823-i39: Actual required frequency (approved plan) ___; sample quantity per check ___; shift start/end ___; unperformed required checks / reason ___; checker initials on linked numeric log ___.

## Catheter mount dimensional and tensile results log

| Actual time | Sample ID | Length cm | Tube conn. ID/OD mm | Other conn. ID/OD mm | Pull force N | Speed/hold | Method + equipment ID | P/F; NCR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ |

Remarks / master-source verification: ________________________________    F-823-i39 current approved Rev.: PENDING     QC checked by: _____________    Reviewed by: _____________

# APPENDIX — Exact Gap-to-In-Process Form Field Crosswalk

This appendix reproduces the Revised Inspection Text field instructions from the seven-column Master Table; they apply exclusively to in-process records or evidence links, NOT to a standalone incoming/final inspection report. Record design choices, including sampling/measurement frequency, need approved control plans before issue.

## GAP-001 | Moisture-loss performance not traceably established

IN-PROCESS LINK FIELD (not routine moisture testing): HME design/configuration ___; qualified moisture-loss report ID/rev ___; declared tidal-volume range ___; evidence status Valid/Pending/N.A. ___; QC initials ___. Verification performed under approved qualification/periodic plan, not every hourly check.

## GAP-002 | Pressure-drop performance and declared results missing

LINK: qualified pressure-drop report ID/rev ___; design ___; approved limit/reference ___; evidence verified by/at ___; result Ref Valid / Pending. Do not add three routine hourly readings unless an approved control plan requires them.

## GAP-003 | Underwater bubble check cannot replace quantitative HME leakage

ROUTINE F-823-i47: bubble screen, approved WI/method ID ___; test setup/pressure or validated surrogate ___; exposure ___ s; tested ___, failed ___; P/F ___; failed unit disposition ___. SEPARATE QUALIFICATION LINK: ISO 9360 §6.4 report ID ___ and approved mL/min claim/limit ref ___.

## GAP-004 | HME compliance and internal-volume information not traceable

QUALIFICATION EVIDENCE LINK: design/revision ___; compliance report ID ___; internal-volume drawing/test ID ___; approved declaration/IFU ID ___; verified by/date ___.

## GAP-005 | HME conical connector conformance not demonstrated

IN-PROCESS: correct ports/fit by approved go/no-go gauge or visual check ___; gauge ID ___; result ___; drawing rev ___. OTHER EVIDENCE: ISO 5356-1 compliance report ID/rev ___; approved design variant ___.

## GAP-006 | Breathing filter filtration efficiency unverified against salt test

QUALIFICATION LINK ONLY: BSF performance report ID/rev ___; manufacturer claim/specification ___; variant adult/pediatric ___; unchanged-media/design verification ___; reviewer/date ___. Routine control remains physical media position/dimensions.

## GAP-007 | Filter challenge sampling and formal ISO test-report content undefined

QUALIFICATION REFERENCE: documented sample-size rationale ID/rev ___; ISO test report ID ___; configuration/lot/date ___; review valid/pending ___. Keep actual data in the laboratory record.

## GAP-008 | HME IFU/marking and declared performance checks incomplete

OTHER CONTROL: Packaging/label verification record ref ___; approved label/IFU rev ___; ISO §7 matrix ref ___; QC review ___; N.A. rationale ___. Do not convert final packaging controls into hourly assembly checks.

## GAP-009 | Sterile HME packaging validation evidence not established

CONTROLLED RECORD LINK (packaging/sterilization stage): sterile? Y/N ___; approved packaging validation report ___; sealing batch record ___; sterilization record ___; final release record ___.

## GAP-010 | Ultrasonic weld control absent from in-process HME form

ADD WELD SETUP/CHANGE LOG: machine ___; recipe/rev ___; weld time actual ___ s [WI 0.5–1.5]; delay ___ s [0.5–1.5]; hold ___ s [0.5–1.5]; pressure indicated ___ kg/cm² [5.5–6.5]; checked time ___; P/F ___; deviation/disposition ___.

## GAP-011 | HME media-width acceptance differs between WI and form

REVISED PAPER DIMENSIONS: adult Ø53 ±1 mm; pediatric Ø30 ±1 mm; paper width approved drawing specification ___ (currently WI 15 mm nominal versus form 15 ±1 mm: HOLD for resolution); actual ___ mm; caliper ___; P/F ___.

## GAP-012 | Underwater HME in-process test method and AQL data incomplete

UNDERWATER LEAK SCREEN: SOP/control plan ___; lot quantity ___; inspection level/AQL if applicable ___; sample size ___; defects ___; setup validated? ___; test duration 10 s; bubbling Y/N ___; accept/reject ___; NCR ___; checker/date ___.

## GAP-013 | Breathing circuit leak-test parameters conflict internally

CIRCUIT LEAK: machine/model/ID ___; validated method/rev ___; pressure set ___ kPa; actual ___; charge ___ s; test hold ___ s/min; equipment output / leak or decay ___ [approved units]; acceptance ref ___; P/F ___; test time/lot/sample IDs ___.

## GAP-014 | Joint pull-off acceptance is not measurable in circuit form

TENSILE RESULT: equipment/calibration ___; sample ID ___; target 45 ±2 N; observed ___ N; speed 50 ±5 mm/min; actual ___; held 1 min; actual ___; test location ≥150 mm; detached Y/N; Pass/Fail ___; NCR ___.

## GAP-015 | Circuit lengths conflict and tolerances are missing

TUBE LENGTH: item code ___; drawing/rev ___; ordered nominal ___ cm; permitted range ___ cm; measured ___ cm; device ID ___; P/F ___.

## GAP-016 | Dimensional checks in WI omitted from circuit form

ADD DIMENSION LOG: variant ___; ID/OD target & tolerance ref ___, actual ___/___ mm; wall/sheet 0.55–0.70 mm actual ___; rib distance 5.8–6.2 mm actual ___; rib thickness 2.3–2.5 mm actual ___; caliper ID ___; P/F ___; N.A. justification ___.

## GAP-017 | Hourly form grid versus required checking frequency unresolved

HEADER: inspection control plan/rev ___; required frequency ___; trigger ___; sample count per check ___; shift actual start/end ___; per-check actual time ___; missed checks / escalation ___. Hourly slots are time-labels, not evidence of a required hourly frequency.

## GAP-018 | Circuit HME status lacks evidence and configuration linkage

HME INSTALLED? Y/N ___; part/configuration ___; HME lot ___; F-823-i47 record ref ___; qualification status ref ___; intended patient category ___; mating port verified ___; P/F ___. N.A. if no HME installed.

## GAP-019 | Product size claims exceed/require ISO applicability determination

CONFIGURATION / SCOPE CHECK: HME/BSF intended tidal-volume range ___; adult/paediatric/neonate ___; ISO 9360-1 applicable Y/N + rationale ___; qualified performance evidence ___; hold unsupported claim for technical/Regulatory decision.

## GAP-020 | Form reference numbers conflict with supplied WI

FORM HEADER PROPOSED: Parent WI: [DOCUMENT CONTROL TO CONFIRM]; product specification: TF-44-PS rev ___; drawing TF-44-PDBM rev ___; controlled form no./rev/date; superseded-reference reconciliation ID ___.

## GAP-021 | Catheter-mount numeric characteristics absent from photographed form

CAT MOUNT FORM: product/type/lot/time; length nominal 15 cm actual ___ and tolerance ref ___; tube connector 15×22 mm actual ID/OD ___; second connector 15×22 mm actual ___; force criterion approved ref ___ N and actual ___ N; result P/F ___; gauge/tester IDs ___; photo-source status “pending master verification”.

## GAP-022 | Binary checks lack accountable sample/results and nonconformance trail

Each form: legend A=Acceptable, X=Not Acceptable, R=Read Remarks, N/A=Justified; actual result capture log with unit, n tested, n failed, instrument & QC ID; deviation/NCR and affected lot trace; signed review. Keep header lot, shift, date and relevant schedule.