QUALITY CONTROL SYSTEMS  /  EVIDENCE-BASED TECHNICAL REVIEW

# Complete Revised WI — CVC Set & CVC Complete Kit

Central Venous Catheter (CVC) Set / Complete Kit

| Issue status | PROPOSED CONTROLLED DRAFT — NOT APPROVED |
| --- | --- |
| Date prepared | 22 Oct 2026 |
| Document type | Standalone full proposed Work Instruction |
| Controlled source | WI-DD-02 Rev.3 dated 01 Feb 2022; company source 11 pages |
| Assessment basis | ISO 10555-1/-3, 11070, 7864, 7886-1, 80369-7 and 11135 as applicable |
| Use restriction | For QC/QA review only. No product acceptance or release permitted on draft alone. |

Evidence rule: “Required” is not “available”. No gap is verified closed without approved documentation and objective records.

| Document ID | Source revision | Proposed status | Authorized revision / effective date |
| --- | --- | --- | --- |
| WI-DD-02 — proposed revision | Rev.3 | DRAFT FOR QC/QA REVIEW | OPEN — APPROVAL REQUIRED |

## 1. Purpose

Control receiving, injection moulding, extrusion, assembly, in-process monitoring, packaging, EO evidence linkage and release transfer for the applicable sterile CVC configurations.

Define source-backed acceptance and evidence ownership without assigning ISO type tests to every production hour.

## 2. Scope and products

Applies to CVC catheter components and associated approved kit variants produced or assembled under WI-DD-02.

Use F-823-i13 for CVC Set and F-823-i13-2 for CVC Complete Kit as distinct records. Apply F-823-i15 for moulded/extruded components, F-823-i27 for 3-way component assemblies when applicable, and packaging forms only when the equipment/product applicability is approved.

## 3. References and controlled documents

ISO 10555-1:2013+A1:2017; BS EN ISO 10555-3:2013; ISO 11070:2014; ISO 7864:2016; ISO 7886-1:2017; ISO 80369-7:2016; ISO 11135:2014, each only within its applicability scope.

DD-02-PS; DD-02-PDBM; WI-7-5-1-1; approved sampling/inspection SOP; form register. These dependent controlled source revisions are OPEN — TECHNICAL INPUT REQUIRED until verified.

ISO 80369-20 methods and other referenced standards must be obtained via approved access if claimed in test dossiers. Do not infer test conditions from missing references.

## 4. Definitions

CVC: central venous catheter. In-process control: monitored manufacturing check and production record, not a final release certificate.

Qualification/type test: documented design/performance test controlling a product configuration, linked to supplier/design/changes.

OPEN TECHNICAL INPUT REQUIRED: unverified acceptance, sample plan, method or specification requiring technical authorization.

## 5. Responsibilities

Production: approved work order, material issues, machine/assembly setup, process identification and immediate segregation of suspect production.

QC Controller: first-off/start-up checks, interval monitoring and accurately completed designated in-process form with instrument IDs, actual readings and sample traceability.

QC Supervisor / QC Manager: review deviations, test/evidence applicability, and records per controlled authority matrix; no approval signature is pre-populated.

Engineering / Supplier Quality / QA: validate drawings, BOM variants, qualification/type evidence, approved control frequency and EO release/validation linkage.

## 6. Incoming material, moulding and accessory receipt

6.1 Verify raw material supplier identity, PO, material grade/color/lot, packaging condition and approved specification; log receiving report (WI-DD-02 Rev.3 §5.1, p.1).

6.2 Prepare approved mould/die and machine recipe. Production issues and QC examines first-off injection samples (housing, knob, connectors) and performs scheduled controlled monitoring in F-823-i15. Reject flashes, cracks, short shots, black spots, blockages and out-of-spec dimensions (source §5.2, p.2).

6.3 Extrusion: verify approved raw materials and die, run first-off checks; inspect tubular smoothness, cream-white appearance where specified, OD/ID, length, wall features and absence of damage/fish-eyes; record specified interval and equipment. Use F-823-i15 and linked DD-02-PS.

6.4 Guidewire receiving: preserve original straight/J-tip, coating, size/length, condition and handling checks. Verify safety wire/core connection, Annex F fracture, Annex G flexing and Annex H peak tensile qualification through controlled supplier/qualification evidence, not guessed routine frequency [CVC-G06].

6.5 Introducer needle / cannula: verify size, bevel/tip, lack of bending, blockage and secure hub. Apply ISO 11070 Annex I force only to the introducer-needle configuration with size-based criterion [CVC-G07]; verify fitting/union and traceability.

6.6 Hypodermic needle: verify color, dimensions, needle point, hub integrity, packaging and appropriate ISO 7864 Table 2 bond + §4.13 lumen-patency evidence [CVC-G08].

6.7 Dilator: examine nominal size, effective length, finish, lumen and Luer hub fit; retain supplier/qualification peak tensile evidence with ISO 11070 Annex C and applicable Table 2 limits [CVC-G09].

6.8 Syringe / guidewire syringe: verify appearance, markings and piston movement; confirm ISO 7886-1 Annex B/C/D type/dossier evidence where applicable; the standard does not itself prescribe syringe lot release sampling [CVC-G10].

6.9 Fixation devices, clamps, injection sealing caps, 3-way stopcock(s), scalpel, sterile tray and other supplied components: preserve WI-DD-02 Rev.3 §§5.4.4–5.4.7 receiving check items including fit, appearance, closure and specification identity; type-specific performance evidence goes to the relevant approved component file.

6.10 The CVC Complete Kit includes additional approved accessories (e.g. tray, drapes, swabs, forceps, towels, tourniquet, suture) only per the approved variant BOM. Verify part number and quantity against BOM; do not automatically apply a universal item list to every variant [CVC-G14].

## 7. CVC printing, dimensions and design-specific checks

7.1 Print the catheter identity, graduation and company markings according to controlled artwork and approved product specification. Preserve the source 1 cm graduation spacing where it is approved for the product/variant and log measured spacing/check method [CVC-G14].

7.2 Confirm CVC tube length, OD, lumen arrangement, soft tip, the approved color marking and any Murphy eye according to selected model. If radio-detectability is required by documented risk assessment, verify qualified evidence for relevant design and material revision [CVC-G04].

7.3 Shaft/end tensile checks: use ISO 10555-1 §4.6 Table 1 and Annex B for its scope; separately use ISO 10555-3 §4.4 for applicable soft-tip junctions. Never substitute the 3/4/5 N soft-tip criterion for shaft tensile. Do not use “N/min” for static force [CVC-G01].

7.4 Inspect applicable surfaces with minimum 2.5x magnification and report observed extraneous matter/surface injuries [CVC-G03].

## 8. Assembly and in-process monitoring

8.1 Verify issued components against the work order and variant BOM before assembly. Arrange approved components in the tray, load guidewire into dispenser/advancer where applicable and confirm compatibility.

8.2 Execute first-off/startup and prescribed interval checks in individual F-823-i13 (CVC Set) or F-823-i13-2 (CVC Complete Kit), preserving original visual/function items. Do not treat a blank time cell as proof of an executed check.

8.3 Record date, shift, lot, product code/configuration, approved drawing/BOM revision, responsible QC, machine/line, equipment ID, actual measurements and disposition. Each open technical criterion must link to the current approved specification prior to use.

8.4 Reconcile components that require a separate controlled verification route (supplier, qualification, periodic) by exact evidence ID. A missing routine test entry shall not be mistaken for evidence that qualification passed.

8.5 Tensile acceptance: apply CVC-G01/CVC-G02 specific control matrices rather than universal “20 N/min”, retain company 20 N only if independently approved and defined.

8.6 Liquid leak and aspiration-air check: verify design/qualification under ISO 10555-1 Annexes C and D; separate any routine fixture screening from such tests [CVC-G05].

8.7 For flow-rate-declared catheters, record each lumen: claimed and measured ml/min, relevant OD band, 80%/90% threshold under ISO 10555-1 §4.9, and evidence route [CVC-G02].

8.8 Preserve visual controls for syringe clarity/graduations, needle/cannula tip, guidewire condition, fixation and clamp function, secure gluing and integrity of all accessory components. Record actual observations and defects.

8.9 Confirm applicable Luer material/dimensions/type-performance qualification reference per ISO 80369-7, rather than inferring qualification from simple “good fitting” [CVC-G11].

8.10 Complete Kit: verify all quantities, sizes and compatibility against approved BOM variant, in F-823-i13-2 pages 1–2 and component genealogy [CVC-G14].

## 9. Packaging and labeling

Confirm correct approved packaging film/paper, forming/cutting, intact sealing, legible product code, size, lot, manufacture and expiration dates, sterile markings and applicable symbols from the approved artwork.

Use packaging form F-823-i4-7 (Neil Pack) or F-823-i4-1 (Zip Lock blow/seal/cut) only if approved for the actual line/product and relevant validation. Their original company temperature, seal-width and timing values remain subject to controlled recipe verification.

Check sealing/peel condition at specified original sample times, document actual setpoints, observed result, instrument, deviation, and packaging specification reference.

## 10. Sterilization linkage

Identify the CVC product lot and accepted packaging configuration on the formal EO load record, not solely on assembly form.

Verify ISO 11135 §11 acceptance of routine cycle conditions, BI result where used, deviations and applicable release prerequisites against the actual load. Process validation and post-change requalification must be controlled. [CVC-G12]

## 11. Final inspection and release

Perform independent post-sterilization product and packaging checks against the approved final inspection/release protocol.

QC/QA authorized final release requires documented results, approved deviations and reviewed sterilization record. In-process signatures are NOT finished-product release approval.

## 12. Records, traceability and document control

F-823-i13 = CVC Set in-process; F-823-i13-2 = CVC Complete Kit in-process (actual supplied title). Original WI §10 lists i13-2 as “Guidewire Syringe”; document-control reconciliation is required before adoption [CVC-G13].

F-823-i15 = injection/extrusion; F-823-i27 = 3-way assembly; F-823-i4-7 = Neil Pack; F-823-i4-1 = Zip Lock Machines; other company forms remain effective only after approved applicability checking.

All revised forms are independent Draft candidates; do not supersede forms or change approvals without controlled effective revision process.

## ISO qualification / evidence matrix (not automatically per lot)

| Gap | Applicable standard | Control route | Evidence / approval input |
| --- | --- | --- | --- |
| CVC-G01 | ISO 10555-1:2013+A1:2017, 4.6, Table 1, ISO p.5, PDF p.15; ISO 10555-3:2013, 4.4, Table 1, ISO p.2, PDF p.10 | Qualification / periodic controlled evidence | Approved tensile matrix by component and size; Annex B or relevant soft-tip method report; instrument calibration; qualified test records and lot/design linkage. Existing 20 N |
| CVC-G02 | ISO 10555-1:2013+A1:2017, 4.9 and Annex E, ISO p.5, PDF p.15 | Qualification / periodic controlled evidence | Approved DD-02-PS flow claims and Annex E test reports by lumen; validated frequency decision. Status: OPEN — TECHNICAL INPUT / EVIDENCE. |
| CVC-G03 | ISO 10555-1:2013+A1:2017, 4.4, ISO p.4, PDF p.14; ISO 11070:2014, 4.3, ISO p.5, PDF p.11 | Qualification / periodic controlled evidence | Approved visual standard, magnifier/instrument suitability, training and completed in-process records. Status: OPEN — EVIDENCE. |
| CVC-G04 | ISO 10555-1:2013+A1:2017, 4.2, ISO p.4, PDF p.14 | Qualification / periodic controlled evidence | Risk assessment, qualified radio-detectability test, materials traceability and change-control review. Status: OPEN — EVIDENCE. |
| CVC-G05 | ISO 10555-1:2013+A1:2017, 4.7.1, 4.7.2 and Annexes C/D, ISO p.5, PDF p.15 | Qualification / periodic controlled evidence | Approved Annex C/D methods and reports, acceptance decisions, equipment verification, traceability. Status: OPEN — EVIDENCE. |
| CVC-G06 | ISO 11070:2014, 8.3–8.6 and normative Annexes F/G/H, ISO p.9, PDF p.15 | Qualification / periodic controlled evidence | Approved guidewire design/control plan; actual Annex F/G/H reports or verified supplier evidence; approved disposition. Status: OPEN — EVIDENCE. |
| CVC-G07 | ISO 11070:2014, 5.4.2 and normative Annex I, ISO p.6, PDF p.12 | Qualification / periodic controlled evidence | Approved needle specification and Annex I test record/supplier evidence. Status: OPEN — EVIDENCE. |
| CVC-G08 | ISO 7864:2016, 4.12 Table 2 and 4.13 Table 3, ISO pp.7–9, PDF pp.13–15 | Qualification / periodic controlled evidence | Approved ISO 7864 control route, actual bond/patency evidence, supplier verification and lot traceability. Status: OPEN — EVIDENCE. |
| CVC-G09 | ISO 11070:2014, 9.3.3, 6.3 Table 2 and Annex C, ISO pp.7,10, PDF pp.13,16 | Qualification / periodic controlled evidence | Approved dilator spec, tensile force qualification and evidence linked to controlled configuration. Status: OPEN — EVIDENCE. |
| CVC-G10 | ISO 7886-1:2017, 13.1, 13.2 and normative Annexes B/C/D, ISO p.10, PDF p.18 | Qualification / periodic controlled evidence | Approved syringe technical specification, actual design tests and accepted-supplier evidence. Status: OPEN — EVIDENCE. |
| CVC-G11 | ISO 80369-7:2016, 4.2–4.3, 5, 6.1–6.6 and Annexes B/C, ISO pp.4–6, PDF pp.11–13 | Qualification / periodic controlled evidence | Approved Luer design/type-test dossier including applicable ISO 80369-7 clauses and ISO 80369-20 test methods (latter not supplied). Status: OPEN — EVIDENCE / MISSING SOURCE. |
| CVC-G12 | ISO 11135:2014, 11.1–11.3, ISO p.23, PDF p.35 | Qualification / periodic controlled evidence | Validated process/qualification, individual load cycle record, BI evidence as applicable, deviation decisions, documented authorized release. Status: OPEN — EVIDENCE. |
| CVC-G13 | Document control reconciliation; source WI-DD-02 Rev.3, p.10 §10 vs F-823-i13-2 Rev.1 header; no ISO clause claimed | Document control | Approved document master list, change request, revision/effective-date confirmation, verified supersession if any. Status: OPEN — APPROVAL. |
| CVC-G14 | ISO 11070:2014, 10 and Table 4, ISO pp.10–11, PDF pp.16–17 | Qualification / periodic controlled evidence | Approved BOM and packaging drawing per set, reconciliation checklist and completed kit records. Status: OPEN — TECHNICAL INPUT / EVIDENCE. |

# 13. Original CVC dimensional variants (REFERENCE ONLY)

Source: WI-DD-02 Rev.3, pp.7–8, Table 1. Listed examples reproduced as original company reference values for reconciliation; they are NOT approved acceptance limits in this draft until checked against controlled DD-02-PS and current BOM. Full source table remains authoritative pending Document Control.

| Model | Tube length | Catheter OD×ID (mm) | Dilator | Dilator OD×ID (mm) | Dilator L (mm) | Needle | Guidewire OD mm × L cm |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 3 FR Single | 6 cm | 0.99×0.62 | 5 FR | 1.86×0.96 | 60 | 22/20 G | 0.45×45 |
| 4 FR Single | 10 cm | 1.36×0.85 | 5 FR | 1.86×0.96 | 60 | 22/20 G | 0.45×45 |
| 5 FR Single | 10 cm | 1.72×1.10 | 6 FR | 1.98×1.00 | 60 | 22/20 G | 0.45×45 |
| 6 FR Single | 10 cm | 1.98×1.30 | 6 FR | 2.00×1.00 | 60 | 22/20 G | 0.45×45 |
| 4 FR Double A | 10 cm | 1.40 OD | 5 FR | 1.86×0.96 | 74 | 22/20 G | 0.45×45 |
| 4 FR Double B | 10 cm | 1.40 OD | 6 FR | 1.98×1.00 | 50.48 | 22/20 G | 0.45×45 |
| 4 FR Triple | 10 cm | 1.54 OD | 6 FR | 2.00×1.00 | 70.70 | 22/20 G | 0.45×45 |
| 5 FR Double | 15 cm | 1.76 OD | 6 FR | 2.00×1.00 | 73 | 22/20 G | 0.55×45 |
| 5 FR Triple | 15 cm | 1.90 OD | 8 FR | 2.50×1.50 | 105 | 22/20 G | 0.55×45 |
| 7 FR Double / 9FR | 30 cm | 2.40 OD | 9 FR | 2.75×1.60 | 104 | 16/17 G | 0.90×45 |
| 7 FR Double / 10FR | 30 cm | 2.40 OD | 10 FR | 2.93×2.00 | 103 | 16/17 G | 0.90×45 |
| 7 FR Triple / 9FR | 30 cm | 2.56 OD | 9 FR | 2.75×1.60 | 103 | 16/17 G | 0.90×45 |
| 7.5 FR Double | 20 cm | 2.56 OD | 10 FR | 2.75×1.60 | 103 | 16/17 G | 0.90×45 |
| 8 FR Double | 20 cm | 3.00 OD | 9 FR | 2.80×1.60 | 103 | 16/17 G | 0.90×45 |
| 7 FR Triple / 10FR | 30 cm | 2.56 OD | 10 FR | 2.92×2.00 | 103 | 16/17 G | 0.90×45 |
| 11 FR Double | 20 cm | 3.90 OD | 12 FR | 3.32×2.00 | 155 | 16/17 G | 0.90×45 |
| 11.5 FR Single | 24 cm | 3.90 OD | 12 FR | 3.32×2.00 | 155 | 16/17 G | 0.90×45 |
| 12 FR Double | 20 cm | 4.05 OD | 12 FR | 3.93×2.50 | 152 | 16/17 G | 0.90×45 |

# 14. Nonconforming product / escalation

## 14.1 Handling

For failed acceptance, out-of-spec readings, missing test evidence or unapproved specification, mark the status FAIL or PENDING, segregate implicated output according to approved hold workflow and notify Supervisor/QA. Record time interval, lots, last conforming check and re-inspection/restart authorization.

Production shall not erase failures or substitute an unapproved visual judgement for a quantitative criterion.

# 15. Revision History

| Version | Effective date | Change description | Status |
| --- | --- | --- | --- |
| 3 | 01 Feb 2022 | Original controlled source WI (not modified) | Baseline |
| DRAFT | 22 Oct 2026 | Source-driven process reconciliation, ISO test routing and in-process form mapping; linked CVC-G01 to G14 | Not approved |