# CHEST TUBE KIT

# ISO 20697 CLOSURE REGISTER

Requirements, route, proposed coverage and objective evidence

Prepared for: QC Manager | Prepared: 19 October 2026

Status: TECHNICAL REVIEW DRAFT — approval and operational evidence pending

Product specification TF-64-PS / TF-64-PDBM not supplied; do not operationalize provisional limits.

Assessment boundary: the source form F-823-i133 is IN-PROCESS ASSEMBLY CONTROL, not a finished-product Inspection & Test Report. Other test methods use linked controlled records.

ISO citation note: printed ISO pages below are derived from the published 2018 text. Official PDF page numbers have NOT been checked; the 2017 DIS preview is a draft and is not used to establish conformity.

# Register status & approval rule

No supplied signed final specification, qualification record, executed lot record or official PDF page audit establishes verified closure. All 25 requirement checkpoints are OPEN for evidence and/or applicability. “Documentary Covered” in this register denotes proposed draft route ONLY; it is not verification of approved implementation or actual product conformity.

Coverage: 25 ISO subclause checkpoints mapped; 21 identified distinct gaps.

Verified closed: 0

Awaiting applicability/technical input: All entries require a signed product configuration/risk matrix; no N/A is asserted.

| Req ID | ISO clause / page | ISO Says (accurate paraphrase) | Applicability / decision | Related Gap | Required evidence / route | Actual verified evidence | Documentary coverage | Operational status | Outstanding action / closure |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-01 | ISO 20697:2018 §4 ¦ printed p.3 Official PDF p. not verified | Safe access, maintenance of drainage; retention only when present | Applicable catheter; retention conditional | GAP-011 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-02 | ISO 20697:2018 §5.1 ¦ printed p.3 Official PDF p. not verified | Documented risk-management analysis and supporting file | Applicable | GAP-012 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-03 | ISO 20697:2018 §5.2 ¦ printed p.4 Official PDF p. not verified | Risk-appropriate biological safety evaluation | Applicable patient-contact components | GAP-012 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-04 | ISO 20697:2018 §5.3 ¦ printed p.4 Official PDF p. not verified | Detectability when required by risk assessment | Risk/claim conditional | GAP-007 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-05 | ISO 20697:2018 §5.4 ¦ printed p.4 Official PDF p. not verified | External effective-length surface free of hazardous defects and extraneous matter | Applicable catheter | GAP-007 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-06 | ISO 20697:2018 §5.5.1–5.5.3 ¦ printed pp.4–5 Official PDF p. not verified | OD nominal designation (±1 Fr) and effective-length unit; no standard length tolerance | Applicable catheter | GAP-007 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-07 | ISO 20697:2018 §5.5.4 ¦ printed p.5 Official PDF p. not verified | Nominal balloon volume in mL if balloon is present | Conditional — feature unverified | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-08 | ISO 20697:2018 §5.6 ¦ printed p.5 Official PDF p. not verified | Avoid misconnection by appropriate risk assessment | Applicable if interfaces/connectors | GAP-011 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-09 | ISO 20697:2018 §5.7 ¦ printed p.6 Official PDF p. not verified | MRI hazard evaluation if applicable | Conditional — environment/claims | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-10 | ISO 20697:2018 §5.8 ¦ printed p.6 Official PDF p. not verified | Appropriate sterilization standards and SAL 10^-6 for sterile device | Applicable marketed sterile kit | GAP-018 | Validated sterilization process / lot release | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-11 | ISO 20697:2018 §6.1 / Annex A informative ¦ printed p.6 / p.11 Official PDF p. not verified | Risk-based catheter kink performance; no default fixed kink threshold | Applicable placement performance | GAP-011 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-12 | ISO 20697:2018 §6.2 / Annex B ¦ printed p.6 / p.13 Official PDF p. not verified | Corrosion evaluation for exposed metallic catheter components when relevant | Conditional — catheter metal | GAP-016 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-13 | ISO 20697:2018 §6.3 / Annex C ¦ printed p.6 / p.14 Official PDF p. not verified | Function-preserving deformation resistance under maximum specified suction | Conditional — negative pressure | GAP-013 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-14 | ISO 20697:2018 §6.4.1 / Table 1 / Annex D ¦ printed pp.6–7 / p.15 Official PDF p. not verified | External connection tensile force by OD band: 5 N or 15 N | Applicable connection existence/OD | GAP-014 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-15 | ISO 20697:2018 §6.4.2 / Table 2 / Annex E ¦ printed p.7 / p.16 Official PDF p. not verified | Catheter/junction tensile force by OD band: 10 N or 20 N | Applicable tubing/junction configuration | GAP-014 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-16 | ISO 20697:2018 §6.5 / Annex F ¦ printed p.7 / p.18 Official PDF p. not verified | Collection device impact resistance; 2% maximum vacuum loss where suction applies | Conditional — collecting/suction device | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-17 | ISO 20697:2018 §6.6 / Annex G ¦ printed p.7 / p.20 Official PDF p. not verified | Flow rate at least 80% or 90% of stated claim when flow claimed | Conditional — flow claimed | GAP-015 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-18 | ISO 20697:2018 §6.7 / Annex H informative ¦ printed p.7 / p.22 Official PDF p. not verified | Retention strength risk-based for non-balloon retention features | Conditional — feature | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-19 | ISO 20697:2018 §6.8 / Annex I ¦ printed p.7 / p.24 Official PDF p. not verified | Balloon safety / no leakage or obstructed drainage holes | Conditional — balloon | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-20 | ISO 20697:2018 §6.9 / Annex J/K ¦ printed p.8 / pp.27,29 Official PDF p. not verified | Balloon integrity / volume maintenance depending on compliant type | Conditional — balloon | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-21 | ISO 20697:2018 §6.10 / Annex L ¦ printed p.8 / p.31 Official PDF p. not verified | Specified balloon traction resistance if applicable | Conditional — balloon | GAP-020 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-22 | ISO 20697:2018 §6.11 / Annex M ¦ printed p.8 / p.35 Official PDF p. not verified | No air leak under manufacturer-defined maximum suction | Conditional — negative pressure | GAP-013 | Qualification / approved risk / biological evaluation / supplier evidence | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-23 | ISO 20697:2018 §7.1 ¦ printed p.9 Official PDF p. not verified | Durability / legibility of marking | Applicable where marked | GAP-019 | Packaging/label / IFU controlled record | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-24 | ISO 20697:2018 §7.2 ¦ printed p.9 Official PDF p. not verified | Necessary package/device labelling and relevant optional/conditional details | Applicable | GAP-019 | Packaging/label / IFU controlled record | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |
| REQ-25 | ISO 20697:2018 §7.3 ¦ printed pp.9–10 Official PDF p. not verified | IFU contents when separate IFU is supplied | Conditional on IFU | GAP-019 | Packaging/label / IFU controlled record | Pending — none supplied | Draft route proposed; not approved | Open | Open — Technical input / evidence / approval; NOT VERIFIED CLOSED |

Source note: Although the published 2018 text is web-verifiable, the official ISO PDF has not been supplied. Published text is not a substitute for official controlled master verification in the approved quality system.

# Documentary / implementation / operational closure — separate gates

| Gate | Closure condition | Current state |
| --- | --- | --- |
| 1. ISO applicability | Official final PDF, product intended-use and model-specific requirement matrix approved. | OPEN |
| 2. Document readiness | All relevant corrected WI clauses and F-823-i133 fields reconciled with signed BOM/spec. | Draft proposed — no approval |
| 3. Document approval | QC/QA/Engineering/Document Control approve controlled revisions and training. | OPEN |
| 4. Actual qualification | Current product and accessories covered by objective applicable ISO reports. | OPEN |
| 5. Lot implementation | Executed incoming/in-process/packaging/sterilization/finished lot records. | OPEN |
| 6. Final release | Approved final-release record with signed deviations and test dispositions. | OPEN |

# Gap action reconciliation

| Gap ID | Priority | Owner / action | Evidence required | Closure status |
| --- | --- | --- | --- | --- |
| GAP-001 | P2 | Document Control — Unreconciled controlled WI reference | Approved document master list, supersession history, and controlled WI/form linkage. No ev | Open — Approval |
| GAP-002 | P1 | Engineering / Production / QC — Conflicting kit BOM and included instruments | Signed current product-specific BOM, approved assembly specification, engineering change r | Open — Technical Input |
| GAP-003 | P1 | Engineering / QA / Supplier Quality — Latex material contradiction | Approved component spec, supplier certificate and assessed labeling/risk file; reconciled | Open — Technical Input |
| GAP-004 | P1 | QA / Production / Supplier Quality — Lidocaine pack description, quantity and expiry | Approved BOM, product label/authorization, supplier documentation and component traceabili | Open — Technical Input |
| GAP-005 | P2 | Engineering / QA — Suture length / unit inconsistency | Approved suture specification and supplier package label clarifying unit and size; revisio | Open — Technical Input |
| GAP-006 | P2 | Engineering / QC — Gauze description / dimensional inconsistency | Approved gauze component specification, approved dimensional drawing, supplier evidence, v | Open — Technical Input |
| GAP-007 | P1 | Engineering / QC / QA — Catheter size / visibility / surface criteria incomplete | Approved product drawing and gauges, calibration IDs, completed dimensional records, radio | Open — Technical Input |
| GAP-008 | P2 | QC / Production / QA — Missing in-process result codes, time accountability and samples | Approved risk-based control plan, completed signed first-piece and shift records, deviatio | Open — Technical Input |
| GAP-009 | P2 | Warehouse / QC / Production — Incoming release and supplier-to-kit traceability | Approved receiving records/OBM inspection, lot traceability, completed issue slip and appr | Open — Evidence |
| GAP-010 | P2 | Engineering / Production / QC — Surgical instrument / protective items not reconciled | Current signed BOM and supplier component acceptance evidence; reconciled part names; comp | Open — Technical Input |
| GAP-011 | P1 | Engineering / QA — Qualification evidence for safe access, kink and drainage performance | Approved risk file, engineering verification and/or clinical performance evidence, configu | Open — Evidence |
| GAP-012 | P1 | QA / Regulatory / Engineering — Unverified risk management and biological evaluation | Approved product-specific risk management, ISO 10993-1-based evaluation and any required s | Open — Evidence |
| GAP-013 | P1 | Engineering / QA — Negative-pressure integrity verification not linked | Authorized intended-use/label, signed max suction, completed Annex C/M report or approved | Open — Technical Input |
| GAP-014 | P1 | Engineering / QA — Connection and catheter tensile force evidence not linked | Signed design verification plan and Annex D/E measurement worksheets with gauge length/dia | Open — Technical Input |
| GAP-015 | P1 | Engineering / QA — Declared catheter flow performance not verified | Approved product claims, completed Annex G report, dimensional specs, risk review and QA a | Open — Technical Input |
| GAP-016 | P2 | Engineering / QA — Metallic catheter components / corrosion applicability unresolved | Risk review, approved materials, Annex B test if applicable, instrument incoming inspectio | Open — Technical Input |
| GAP-017 | P1 | Production / QA / QC — Sterile barrier / seal validation linkage and checks incomplete | Approved pouch specification, seal validation, sealer maintenance/calibration, actual pack | Open — Evidence |
| GAP-018 | P1 | QA / QCM — Sterilization and final release evidence not traceable | Validated sterilization process, actual batch-cycle and lab results, authorized review, F- | Open — Evidence |
| GAP-019 | P2 | Regulatory / QA / QC — Label/IFU acceptance and variant traceability incomplete | Approved finished-label artwork, IFU, translation/regulatory review, executed packaging/la | Open — Evidence |
| GAP-020 | P2 | Engineering / QA / Regulatory — Unresolved design applicability: balloon, retention, collection device | Approved product specification/BOM, signed risk-based applicability matrix and tests if fe | Open — Technical Input |
| GAP-021 | P2 | QC / QA / Production — Nonconforming product hold & process disposition missing | Executed NCR/reject/segregation records, documented disposition and approved re-verificati | Open — Approval |

# Approval / evidence verification

QC/QA Technical review: __________________  Date: ______________

Engineering/configuration confirmation: _____________  Date: ______________

Document Control reconciliation: ________________  Date: ______________

Actual objective evidence checked by: __________________  Date: ______________

ISO PDF verified print-to-PDF pagination by: __________________  Date: __________

# Sources

# Source / evidence rules

Uploaded WI: WI-8-2-2-64 Rev.01, date 15.12.2022, pages 1–4. Read directly from the supplied .doc; section/page links refer to document printed pages.

Uploaded assembly record: F-823-i133 Rev.04, date 10.10.2025, sheet F-823-i133 (Rev 4), one printed page, exact Excel row references indicated.

ISO 20697:2018 (published edition), verified via publicly accessible published-text rendering; printed pages verified from contents. Official final-edition PDF not supplied; PDF page offsets remain UNVERIFIED. The accessible sample PDF is ISO/DIS 20697:2017 (draft), not an approved substitute.

Referenced but NOT supplied: TF-64-PS, TF-64-PDBM, approved product BOM, WI-7-5-1-1, F-823-i4-2, F-8-2-4, component reports and qualification / biocompatibility / packaging validation evidence.

The 7-column master table preserves the required titles. “Inspection Text” columns refer only to the supplied IN-PROCESS CONTROL record in this assignment. An ISO test not suitable for routine inspection is linked to a separate controlled qualification record.

Published ISO text: ISO 20697:2018 — https://www.theiteh.com/catalog/standards/iso/6aeb145d-52d7-4f24-866e-f2320e7f7646/iso-20697-2018

Official ISO scope/edition: https://www.iso.org/standard/68855.html