# CHEST TUBE KIT

# REVISED IN-PROCESS CONTROL FORM

Editable assembly quality record • replaces NO current controlled form without approval

Prepared for: QC Manager | Prepared: 19 October 2026

Status: TECHNICAL REVIEW DRAFT — approval and operational evidence pending

Product specification TF-64-PS / TF-64-PDBM not supplied; do not operationalize provisional limits.

Assessment boundary: the source form F-823-i133 is IN-PROCESS ASSEMBLY CONTROL, not a finished-product Inspection & Test Report. Other test methods use linked controlled records.

ISO citation note: printed ISO pages below are derived from the published 2018 text. Official PDF page numbers have NOT been checked; the 2017 DIS preview is a draft and is not used to establish conformity.

# A  Identification, conditions & authorisation

| Field | Enter / verify actual record | Field | Enter / verify actual record |
| --- | --- | --- | --- |
| Proposed record ID | F-823-i133 — supersession PENDING | Current verified source | F-823-i133 Rev.04 |
| WI number/revision | __________________ [GAP-001] | TF-64-PS / TF-64-PDBM revisions | __________________ |
| Kit name / product code | __________________ | Product configuration / size | __________________ |
| Production order / batch | __________________ | Lot number / date | __________________ |
| Shift / production line | __________________ | Supervisor / sample ID | __________________ |
| Control plan ID/rev | __________________ | Sample size / frequency | OPEN TECHNICAL INPUT REQUIRED |
| First-piece QC status | PASS / FAIL / HOLD / PENDING | Time/inspector / initial sample | __________________ |
| Lot receiving clearance | OBM / accepted record: ______ | BOM / approved revision | __________________ |

Code legend: P = Pass against verified approved spec; F = Fail; H = Hold / pending investigation; NA = approved Not Applicable with reason; blank is NOT proof of verification. Shift-time columns in original were 07, 08, 09, 10, 11, 12, 01, 02, 03, 04, 05, 06. This draft uses detailed event entries so time and sample remain traceable without assuming hourly sampling.

# B  Component verification checklist (19 original checkpoints retained)

CONTROL: items/designation from original Rev.04 copied for recognition; items with contradictions must be reconciled against a signed BOM/product spec BEFORE use. Listed text is NOT independently approved acceptance criterion. [GAP-002 to GAP-010]

| # | Description | Existing Rev.04 requirement (NOT approved revision) | New verified specification / criterion | Actual quantity / unit | Lot / expiry | Observation / sample & time | Result | Initials / equipment ID |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Clarity , Appearance | Existing source row 10: Clear and Good appearance | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 2 | Chest Tube Catheter and Trocar | Existing source row 11: Size # 08 clear and soft catheter, printed graduation scale, with X-ray line, trocar no damaged | SPEC RECONCILIATION REQUIRED | ________ | ________ | ________ | P / F / H / NA | ________ |
| 3 | Pair of Large Gloves | Existing source row 12: Size Large, Latex free and powder free | SPEC RECONCILIATION REQUIRED | ________ | ________ | ________ | P / F / H / NA | ________ |
| 4 | 20ml ampule Lidocaine HCL 1% (4pcs) | Existing source row 13: not < 5ml each ampule, not broken , clean and clear, with valid expiry date | SPEC RECONCILIATION REQUIRED | ________ | ________ | ________ | P / F / H / NA | ________ |
| 5 | 25 Ga. x 1-1,2" Needle | Existing source row 14: Sharp, no damage, no blockage | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 6 | Needle curved kelly clamp (16cm) | Existing source row 15: Smooth Surface, good function and strong claming | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 7 | Needle curved kelly clamp (14cm) | Existing source row 16: Smooth Surface, good function and strong claming | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 8 | Silk suture | Existing source row 17: 0.30 inches ( 75cm ) non absorbable | SPEC RECONCILIATION REQUIRED | ________ | ________ | ________ | P / F / H / NA | ________ |
| 9 | 4" x 4" Gauze Pads | Existing source row 18: Dimension 38*18cm,soft and white | SPEC RECONCILIATION REQUIRED | ________ | ________ | ________ | P / F / H / NA | ________ |
| 10 | Suture Needle Holder (12cm) | Existing source row 19: Smooth surface, no roughness, good clamp, no rust particles | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 11 | Skin Marker Pen | Existing source row 20: Smooth surface, should not dry | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 12 | Ruler | Existing source row 21: Plastic type 15cm ruler, printed scale | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 13 | Petroleum Gauze | Existing source row 22: Pure vaseline gauze, smooth, sticky | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 14 | Mayo Scissor (14cm) | Existing source row 23: Smooth surface, sharp edges of both blade, no rust particles | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 15 | Scalpel Blade #11 | Existing source row 24: Scalpel with sharp stainless steel blade | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 16 | 10ml Syringe with G:23 x 1-1,2" needle | Existing source row 25: clear, no leakage strong fitting with sharp needle, no damaged, no crack | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 17 | Fenestrated Drape Adhesive type(60 * 60cm) | Existing source row 26: Non woven , low lint , neat and clean, accurate in dimension | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 18 | Compartment Tray | Existing source row 27: Good and Secured, no damaged, no crack | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 19 | Hospital Wrap | Existing source row 28: Accurate in dimension, straight cutting edges, no holes, no stains | Verify controlled spec / BOM | ________ | ________ | ________ | P / F / H / NA | ________ |
| 20 | Surgical Tape — if confirmed in approved BOM | WI p.2 §5.9.1; missing from form | Verify approved spec and qty; do NOT assume supplied | ________ | ________ | ________ | P / F / H / NA | ________ |

For every item: manufacturer-approved criteria and actual observations take priority over historic uncontrolled descriptions. Accepted sample size, criteria and frequency depend on the approved control plan. Additional BOM lines may be inserted after Engineering review.

# C  Event / time control log

| Date & time | Control type (first-piece / in-process) | Sample / tray / lot ID | Inspection characteristic & actual result | Outcome P/F/H/NA | QC initial / instrument |
| --- | --- | --- | --- | --- | --- |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |

# D  Evidence links — external verification (NOT hourly tests)

| Verification / process | Applicable? Y/N/Pending | Actual approved record ID/revision | Review status |
| --- | --- | --- | --- |
| Approved BOM / TF-64-PS / TF-64-PDBM | PENDING | ________________ | PENDING |
| Incoming lot release / supplier CoC | PENDING | ________________ | PENDING |
| ISO 20697 risk/biocompatibility file | PENDING | ________________ | PENDING |
| ISO 20697 performance/type test file | PENDING | ________________ | PENDING |
| Packaging control (F-823-i4-2 or replacement) | PENDING | ________________ | PENDING |
| Sterilization cycle / laboratory test | PENDING | ________________ | PENDING |
| Final release (F-8-2-4) | PENDING | ________________ | PENDING |

# E  Nonconformity / hold & disposition

Observation / deviation: _________________________________________________________

Affected lot/tray/sample/quantity: _________________________________________________

Date/time identified: _________________   P / F / HOLD: __________________________

NCR / Reject reference: ________________________________________________________

Immediate containment / segregation: ______________________________________________

Authorised disposition / re-verification: ___________________________________________

Resumption allowed by / date: ___________________________________________________

# F  Signature and review

QC checked by: ____________________  Signature: ______________ Date/Time: __________

QCM reviewed by: _________________  Signature: ______________ Date/Time: __________

Assembly status: PASS / HOLD / REJECT / PENDING  |  No automatic final release.

Final product release is a separate authorised process on F-8-2-4 and required batch records.