# ISO / DOCUMENT CONTROL CLOSURE REGISTER

Clean Cart A + C • source-derived gap controls • ISO scope not yet established

Coverage summary: ISO clause-by-clause assessment NOT PERFORMABLE; no official applicable standard or controlled technical specifications supplied. The entries below are WI/form reconciliation findings. "Documentary Covered" here means a proposal exists, NOT a verified controlled document or evidence of product conformity.

| Control status | Count |
| --- | --- |
| Verified ISO clauses | 0 - ISO source not supplied |
| Gaps established by supplied WI/form records | 13 |
| P1 / P2 / P3 | 7 / 6 / 0 |
| Verified Closed | 0 |
| Operational ISO closure | NOT ASSESSABLE |

# 1. ISO applicability gate

| Question | Current answer | Required closure |
| --- | --- | --- |
| Which ISO edition applies to cleaning cartridges? | OPEN. ISO 23500-4:2024 publication scope covers haemodialysis concentrates; applicability to cleaning cartridges not established. | Approved Regulatory/QA applicability memorandum and official standards as relevant. |
| Exact normative clauses / pages / annexes? | Not provided; no ISO clause ID or page asserted in this report. | Supply official, applicable ISO standard/edition and trace clauses against product design. |
| Applicable company approved specifications / risk controls? | BOM, material grades, filter/leak/seal methods, label artwork and sampling plans unavailable. | Engineering/QA-approved controlled technical file and methods. |
| Operational conformity evidence? | No completed production or incoming records / qualification reports provided. | Approved completed test reports, applicable lot records, actual review signatures. |

# 2. Gap-level closure register

| ID / scope | Requirement / issue | Documentary proposal | Evidence required | Verified status |
| --- | --- | --- | --- | --- |
| GAP-001 A + C | Wrong chemical named in visual checks | Candidate revised WI + A/C in-process form; NOT approved | Approved A and C BOM/material specifications, incoming material lot and release documents, completed identity check, QA-approved WI/form. | OPEN - approval / input / evidence |
| GAP-002 C | Citric-acid hydration state unresolved | Candidate revised WI + A/C in-process form; NOT approved | Approved C BOM/specification plus documented comparison/decision for monohydrate vs anhydrous; supplier CoA, design/risk evaluation if changed. | OPEN - approval / input / evidence |
| GAP-003 A | A fill-weight acceptance conflict | Candidate revised WI + A/C in-process form; NOT approved | Approved A fill specification and change decision; revised controlled form and completed calibrated-scale results from representative approved production; deviation resolution if needed. | OPEN - approval / input / evidence |
| GAP-004 C | C fill-weight method and tolerance unverified | Candidate revised WI + A/C in-process form; NOT approved | Approved C mass specification and validated method; calibration proof and filled in-process weight record; approved disposition of current mismatch. | OPEN - approval / input / evidence |
| GAP-005 A + C | Uncontrolled multi-size options on common form | Candidate revised WI + A/C in-process form; NOT approved | Approved variant matrix, separate A/C acceptance limits, traceable controlled form or approved conditional selection, change-control record. | OPEN - approval / input / evidence |
| GAP-006 A + C | Leak verification lacks objective method | Candidate revised WI + A/C in-process form; NOT approved | Approved and validated/verified leak method/limits, QC results per sampling plan, traceable instrument status, deviation dispositions. | OPEN - approval / input / evidence |
| GAP-007 A + C | Filter-retention functional check undefined | Candidate revised WI + A/C in-process form; NOT approved | Approved filter method and acceptance criteria; completed controlled evidence of retention and assembly correctness, where required qualification study. | OPEN - approval / input / evidence |
| GAP-008 A + C | Seal/fit assessment insufficiently defined | Candidate revised WI + A/C in-process form; NOT approved | Approved cap/filter drawing and defect catalogue, method verification records, completed inspection evidence; documented route between in-process and qualification tests. | OPEN - approval / input / evidence |
| GAP-009 A + C | Powder condition and contamination records incomplete | Candidate revised WI + A/C in-process form; NOT approved | Approved material visual criteria, lot CoA/incoming records, sampled production checks and nonconformance evidence, if any. | OPEN - approval / input / evidence |
| GAP-010 A + C | Label/UDI and unsupported claims not aligned | Candidate revised WI + A/C in-process form; NOT approved | Approved country/market labelling decision, released artwork and UDI specification, completed printing checks, scan log where applicable. | OPEN - approval / input / evidence |
| GAP-011 A + C | Sampling and measured results not traceable | Candidate revised WI + A/C in-process form; NOT approved | Approved in-process control plan, calibrated device list, completed traceable batch controls, training and review records. | OPEN - approval / input / evidence |
| GAP-012 A + C | Incoming AQL plan not identifiable | Candidate revised WI + A/C in-process form; NOT approved | Approved receiving sampling procedure, signed receiving report F-823-T1 (or controlled successor), CoA, disposition log and material traceability. | OPEN - approval / input / evidence |
| GAP-013 A + C | Record and revision linkage not reconciled | Candidate revised WI + A/C in-process form; NOT approved | Controlled document master list, revision change request, approved source crosswalk, accessible TF specs, training, completed separate final conformity release records as applicable. | OPEN - approval / input / evidence |

# 3. Previously existing requirements reviewed without separate gap count

| WI item | Source | Review position |
| --- | --- | --- |
| Product purpose, scope and responsibilities | M02 / M03 p.1 secs.1-4 | Existing baseline carried into reconstructed WIs; actual execution not verified. |
| Receiving physical condition and lot verification | M02 / M03 p.1 sec.5.1 | Retained; sampling and material records require approval/evidence (GAP-001, 009, 012). |
| Mould installation and cooling / hydraulic checks | M02 / M03 p.2-3 sec.5.3 | Retained; settings depend on approved setup specification not supplied. |
| Appearance, cracks and sharps controls | M02 / M03 p.3 secs.5.4/5.5 | Retained with proposed objective visual/record linkage; no product-specific criteria invented. |
| Conformity checks and release records | M02 / M03 p.4-5 secs.5.7/6 | Source references retained but external report content not verified. |
| Sodium bicarbonate powder bag form F-823-i24-1 | Excel rows 36-66 | Excluded: unrelated separate record and product family. |

# 4. Verification instructions and closure definitions

• Verified Closed: signed approval(s), traceable executed evidence and relevant acceptance demonstration actually verified.

• Documentary Covered - Operational Evidence Pending: controlled documentation approved but test/lot evidence missing. NOT assigned to these unapproved candidate drafts.

• Open - Technical Input: no approved specification, acceptance criterion, method or applicability disposition.

• Open - Approval: proposed controls exist but have not been reviewed and released by authorized personnel.

• Open - Evidence: applicable test/record is required but completed objective evidence was not available.

• N/A: supported by an approved and referenced scope/applicability decision; do not infer solely from wording or this audit.

# Source traceability and review limitations

WI M02 Rev.1, date 10.06.2020: Clean Cart A, printed pages 1-5, user-uploaded PDF pages 9,10,8,7,6 respectively.

WI M03 Rev.2, date 10.06.2020: Clean Cart C, printed pages 1-5, user-uploaded PDF pages 5,4,3,2,1 respectively.

F-823-i24 Rev.6, date 19.10.2025: Clean Cart in-process record, Excel sheet rows 2-32. A separate, unrelated sodium bicarbonate powder-bag form F-823-i24-1 Rev.2 occupies rows 36-66 and is excluded.

External contextual source only: U.S. FDA Baxter K232467 (2023), https://www.accessdata.fda.gov/cdrh_docs/pdf23/K232467.pdf, especially 510(k) summary; it is not the company approved specification.

ISO scope-screening only: ISO 23500-4:2024 published scope, https://www.iso.org/cms/live/live/en/sites/isoorg/contents/data/standard/08/43/84371.html?browse=tc. The standard addresses dialysis-fluid concentrates and its applicability to cleaning cartridges is NOT established.

NOT AVAILABLE: any complete official applicable ISO edition, source ISO PDF pages, company BOMs/technical files, product/material specifications, AQL plan, test procedures, controlled approval list, F-823-T1, F-823-T90, F-823-15, F-823-F57, or completed lots/evidence.