# CONTROLLED DOCUMENT CHANGE REGISTERS

BEFORE → PROBLEM → ACTION → AFTER → EVIDENCE | All changes are proposed drafts

STATUS: PROPOSED CONTROLLED DRAFT  |  NO SIGNATURES OR RELEASE AUTHORIZATION

Source basis: ISO 16628:2022; ISO 5361:2016 invoked clauses/annexes; conditional ISO 8836:2014; scanned WI Rev 00 and non-matching inspection report photo.

# 1. Revised WI change register

| CHANGE / GAP / LINKED REQ | ORIGINAL LOCATION + EXISTING WI TEXT | EXACT PROBLEM | ACTION | REVISED WI TEXT (COMPLETE PROPOSAL) | NEW SECTION | EVIDENCE |
| --- | --- | --- | --- | --- | --- | --- |
| W-C001 GAP-001 R001, R003, R005 | WI p.1 §5.1 and pp.2–10 Receiving COA and visual inspection listed; no ISO 18190 applicability table. | Referenced requirements unverified | ADD | Before release of this WI revision, QA shall obtain ISO 18190:2016 Clauses 4, 5, 6, 7, 8 and 9; record applicability, approved control location, evidence and owner for each applicable requirement. Any unresolved item shall remain OPEN and may not be declared compliant. | 3. References / 7. External standards control | ISO 18190 clause assessment and approved technical file |
| W-C002 GAP-002 R002 | WI pp.1–11 No alternative type-test validation clause identified. | No substitution approval path | ADD | Where a specified ISO type-test method is replaced by an alternative, Engineering and QA shall document equivalent safety, validate the alternative against the referenced ISO method, approve the validation before use, and retain records. | 9. Test strategy | Approved comparative validation |
| W-C003 GAP-003 R004 | WI pp.1–11 No ISO 18562-1 pathway reference identified. | Biocompatibility not documented | ADD | QA shall control an evaluation and test strategy for breathing-gas-pathway biocompatibility in accordance with ISO 18562-1. Evidence and acceptance shall be governed by the approved assessment; ISO 18562-1 requirements remain pending until the standard and assessment are reviewed. | 7. Materials / 9. Qualification | Approved ISO 18562 evaluation/report |
| W-C004 GAP-004 R006, R007, R008 | WI p.2 §§5.1.2–5.1.3 Designated Fr based on averaged axes; major/minor axis tolerance ±0.5 mm. | Raw nominal dimensions and reconciliation absent | MODIFY | Measure and record tracheal-segment major and minor outer axes separately against the approved size-specific drawing; each measurement shall meet its approved nominal ±0.5 mm. Calculate/verify the declared whole-number French size from the controlled design rule. Treat the standard’s “should” size statement as guidance until Engineering documents the interpretation. | 9.1 Dimensional verification | Approved drawing and dimensions worksheet |
| W-C005 GAP-005 R009, R010, R021 | WI p.5 §5.1.25 and p.4 valves Connector declared compliant with ISO 5356-1; valves visually inspected. | No normative verification records | MODIFY / ADD | Verify machine-end 15-mm conical male connector conformity against ISO 5356-1 and cuff-inlet compatibility with ISO 80369-7 L1 male Luer-slip using approved design evidence. Check and reconcile referenced ISO 5361:2016 §5.6.5 requirements. Rejection limits shall come only from the approved referenced standards. | 9.3 Connector inspection / 9.6 inflation | Connector compliance technical file |
| W-C006 GAP-006 R011, R012 | WI p.3 §§5.1.5–5.1.6 and pp.7–8 9.0 kPa or 1.5× diameter, first reached; whole system immersion ≥10 s with no bubbles. | No corresponding report/result evidence | KEEP / MODIFY | For each applicable tracheal and bronchial cuff, record the selected inflation endpoint (9.0 kPa/90 cmH2O or 1.5× measured cuff diameter, whichever occurs first), actual reached pressure/diameter, inflation-source disconnect, entire-system immersion and observation for at least 10 s. Accept only with zero bubbles; log cuff identity separately. | 9.4.1 Cuff leakage | Cuff leak worksheet with readings and sample IDs |
| W-C007 GAP-007 R013, R014, R015 | WI p.3 §§5.1.7–5.1.8; p.8 §§7.7–7.8 Cuff diameter incorrectly refers ISO 5361 Annex C; herniation incorrectly Annex E. | Wrong annexes; prerequisites missing | MODIFY | Measure both cuff diameters using ISO 5361:2016 Annex B: stabilize at (2.0±0.1) kPa at (23±2)°C for 5 min; record four diameters 45° apart and calculate their mean (mm, two decimals). Compare each to its marked nominal within ±15%, disclosing the widest diameter for a noncircular bronchial cuff. Test herniation per ISO 5361:2016 Annex D, individually inflating each cuff; do not allow extension beyond the nearest edge of the appropriate patient-end segment. | 9.4.2 Cuff dimensions / 9.4.3 herniation | Separate cuff diameter / herniation worksheets |
| W-C008 GAP-008 R016 | WI p.3 cuffs Cuffs described as smooth, with no burst or leak. | No documented sharp-edge risk link | LINK / ADD | Inspect the cuff attachment and transition for sharp edges; link the product-specific risk analysis and approved visual acceptance standards. Record outcome for each applicable cuff. | 9.4.4 Surface transition | Risk-management file and inspection result |
| W-C009 GAP-009 R017, R018, R019, R020, R022, R023, R024 | WI pp.3–4 §§5.1.9–5.1.13 Visual tube/pilot/valve instructions, without complete dimensions or deflation checks. | ISO 5361:2016 §5.6 not completely incorporated | MODIFY / ADD | Control cuff inflation-tube dimensions, separation angle, pilot/deflation indicator, intentional deflation, connector interface, S3 length conditional cases and any separation-distance mark per ISO 5361:2016 §§5.6.1–5.6.8 as invoked by ISO 16628:2022 §6.6. Use the component-specific approved drawing and record quantitative actuals. | 9.5–9.6 Inflation subassemblies | Inflation-system verification worksheet |
| W-C010 GAP-010 R025, R026, R027 | WI p.10 §§8–9 Seal, sterilization carton stamp and labels checked. | Referenced ISO 18190 Clauses 7–9 not verified | LINK / ADD | Packaging, sterile supply and manufacturer information shall be released only against approved specifications and the documented applicability of ISO 18190:2016 Clauses 7–9. Record packaging integrity, sterilization-release link where applicable, and controlled artwork/IFU revision; do not equate carton chamber stamp with sterile-release evidence. | 13. Packaging / 14. Sterilization / 15. Release | Approved package/sterilization/label technical file |
| W-C011 GAP-011 R028, R029 | WI p.5 §5.1.26 / p.9 §7.25 Legible at 1m ±10mm, 215 ±5lx, rub 15s wet cloth. | Pre-exposure test and visual acuity absent | MODIFY | For marking qualification, first expose markings for the intended-use cumulative exposure duration to applicable IFU-listed drugs/chemicals and simulated contact substances identified in the risk file. Then verify legibility with vision acuity 1 at (1m±10mm) and (215±5lx); rub 15s without undue pressure with distilled-water-wet cloth. Marking shall remain legible. | 9.7.1 Marking durability | Conditioning rationale, time/illumination readings, result |
| W-C012 GAP-012 R030, R031, R033, R034 | WI p.2 graduation; pp.5,9 marking Graduation and marking described generally. | Mandatory mark list incomplete | ADD | Verify on each applicable tube: whole-number Fr; bronchial OD in mm prefixed “Br” (verified using ISO 16628 Annex B); effective IDs in mm for bronchial AND tracheal lumina (verified per Annex C); length marks at 2-cm intervals from the bronchial tip at the specified location; and the correct “right” or “left” identification. Record each result individually. | 9.7.2 Device markings | Five distinct marking lines with actual values |
| W-C013 GAP-013 R032 | WI p.2 §5.1.3 Generic tube IDs and ODs required as specification. | No ISO16628 effective-ID method | ADD | Determine maximum effective inside diameter of the bronchial and tracheal lumina using ISO 16628:2022 Annex C: calibrated three-bead set (diameter accuracy ±0.01 mm, length accuracy ±0.1 mm), spaced by 1–2 mm knots and weighted (100±5 g). Record maximum passing diameter for each lumen, expressed downward to the nearest 0.2 mm. Verify marked values match approved claims. | 9.2.2 Effective lumen ID | Bead-set verification records |
| W-C014 GAP-014 R035 | WI p.5 §5.1.27 Entire bronchial machine end, bronchial cuff and pilot balloon blue. | Preserve correct current control and separate result field | KEEP / LINK | The machine end of the bronchial segment, the bronchial cuff, and its associated pilot balloon shall each be entirely blue. Retain the distinct company control that tracheal-side elements are clear, subject to approved configuration specification. | 9.8 Appearance / Color | Component-level blue/clear inspection records |
| W-C015 GAP-015 R036, R037, R038 | WI p.10 §8 Generic package label content specified. | ISO16628 additional labels not enumerated | ADD | In addition to approved company packaging data, verify that the individual pack or insert gives content description, controlled major and minor tracheal outside axes, and both tracheal and bronchial cuff marked diameters. Compare these with design specification and applicable device markings. | 13.2 Packaging marking | Approved label artwork proof and recorded check |
| W-C016 GAP-016 R039, R040 | WI p.5 §5.1.24 Bronchial tip curve/cutting smoothness inspected. | Margin-of-safety advisory not documented | LINK | Engineering shall document consideration of bronchial segment length, bronchial cuff length/position, bronchial tip bevel, and relevant right-side anatomy in the approved risk file. The “should” statements in ISO 16628 Clauses 6.4.3/6.7 are advisory; no fabricated numeric acceptance limit shall be added. | 8. Design risk / 9.8 Tip | Approved design/risk rationale |
| W-C017 GAP-017 R041, R042 | WI p.5 §5.1.29; p.10 §7.28 Suction catheter visual and ISO8836 Table2 pull force included. | BOM component scope and ISO8836 method absent | MODIFY / LINK | If product configuration includes a separately classifiable respiratory suction catheter, Supplier Quality/Engineering shall document ISO 8836 applicability and verify attached components using applicable §8.1.1 Table 2 and Annex B controls. The test shall not be applied to every endobronchial assembly connector without technical justification. | 9.9 Optional suction catheter | BOM approval and component axial force worksheet |
| W-C018 GAP-018 | WI p.11 §10; provided End.pdf p.1 WI lists Endobronchial report reference; provided report photograph is Endotracheal Tube. | Cannot compare an actual Endobronchial report row by row | ADD / LINK | QC shall use only the verified current Endobronchial Inspection & Test Report revision linked to this WI. Until obtained and approved, the accompanying proposed form shall remain DRAFT and no claim regarding completeness of the current controlled Endobronchial report is permitted. | 16. Records / 17. Document control | Confirmed controlled form and revision |
| W-C019 GAP-019 | WI p.4 §§5.1.14–5.1.15 duplicate 5.1.14 and 5.1.15 appear twice for separate components. | Non-unique reference addresses | MODIFY | Assign unique section identifiers in the proposed WI and preserve the original page/section mapping in the change-control register. Original WI remains authoritative until approved supersession. | All / Revision history | Revision approval and mapping |
| W-C020 GAP-020 | WI pp.1,6–7,10 Visual general AQL, physical special AQL, two-hour packaging QC and SOP-8-2-6 cited. | Sampling plan and Ac/Re not available | KEEP / LINK | Maintain existing company requirements for visual/general and physical/special sampling and two-hour packaging checks, subject to the applicable approved sampling SOP/plan. Each lot/interval shall record lot quantity, inspection level, AQL, sample size, Ac, Re, randomization, evidence and disposition; no nominal level or AQL shall be assumed. | 11. Sampling / 15. Final release | Approved SOP/plan and completed sampling records |

# 2. Revised inspection report change register

Before column cannot be completed for actual Endobronchial report because it was not provided. The Endotracheal report image is not treated as a baseline. All proposed inspection change actions are ADD/LINK and are subject to subsequent true row-by-row redline against verified controlled source.

| CHANGE / GAP | CURRENT REPORT FIELD / ROW | PROBLEM | ACTION | PROPOSED ROW / REVISED REPORT TEXT | CONTROL LOCATION | EVIDENCE |
| --- | --- | --- | --- | --- | --- | --- |
| I-C001 GAP-018 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Product/form identification Correct Endobronchial form ID, rev, code, left/right, WI rev, lot/batch and controlled spec | Proposed IR row 01 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C002 GAP-020 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Sampling & disposition Lot quantity, plan ID/rev, inspection level, AQL, code, n, Ac, Re, drawn IDs, disposition | Proposed IR row 02 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C003 GAP-004 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Major tracheal OD axis Nominal from drawing ±0.5 mm; actual in mm; equipment ID | Proposed IR row 03 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C004 GAP-004 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Minor tracheal OD axis Nominal from drawing ±0.5 mm; actual in mm; equipment ID | Proposed IR row 04 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C005 GAP-004 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | French designated size Whole number; measured axes / controlled size rule / printed Fr | Proposed IR row 05 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C006 GAP-005 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | 15 mm conical connectors ISO5356-1 conformance evidence ID and approved result | Proposed IR row 06 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C007 GAP-005 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Inflation inlet ISO80369-7 L1 male Luer-slip compatibility record | Proposed IR row 07 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C008 GAP-006 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Tracheal cuff leakage 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 | Proposed IR row 08 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C009 GAP-006 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Bronchial cuff leakage 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 | Proposed IR row 09 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C010 GAP-007 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Tracheal cuff diameter ISO5361 Annex B; claimed mm; measured mean mm; ±15% | Proposed IR row 10 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C011 GAP-007 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Bronchial cuff diameter ISO5361 Annex B; declared/widest mm; mean mm; ±15% | Proposed IR row 11 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C012 GAP-007 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Cuff herniation per cuff ISO5361 Annex D; tracheal/bronchial separate observations | Proposed IR row 12 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C013 GAP-008 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Cuff sharp edges Visual and risk file ID; none at attachments/transitions | Proposed IR row 13 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C014 GAP-009 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Inflation tube / geometry ≤3.0mm OD, ≤10% intrusion where applicable, separation angle ≤45°, S3 conditional | Proposed IR row 14 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C015 GAP-009 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Pilot function & deflation Indication of inflation/deflation, unobstructed intentional deflation | Proposed IR row 15 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C016 GAP-011 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Marking exposure durability Exposures/duration, 1m±10mm, 215±5lx, acuity1, 15s wet-rub result | Proposed IR row 16 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C017 GAP-012 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Fr / Br markings Whole Fr; Br marked OD; Annex B actual mm | Proposed IR row 17 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C018 GAP-013 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Effective IDs / both lumina ISO16628 Annex C; tracheal and bronchial actual and claimed mm | Proposed IR row 18 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C019 GAP-012 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Tip-origin length marks 2cm gradations measured from bronchial tip | Proposed IR row 19 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C020 GAP-012 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Side marking Right or left matches configuration | Proposed IR row 20 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C021 GAP-014 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Bronchial color Bronchial machine end + cuff + pilot each blue | Proposed IR row 21 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C022 GAP-015 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Pack/insert content Description, major and minor axes, tracheal and bronchial cuff diameters | Proposed IR row 22 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C023 GAP-010 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Pack sterile link Pack sealing and sterile-release references where applicable; ISO18190 review pending | Proposed IR row 23 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C024 GAP-001 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | External general/material/design technical file ISO18190 Clauses 4–6 applicability and evidence ID; no implicit PASS | Proposed IR row 24 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C025 GAP-003 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Breathing-gas pathway biocompatibility ISO18562-1 evaluation, approved results and reviewer | Proposed IR row 25 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C026 GAP-010 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Manufacturer information / IFU approval ISO18190 Clause9 assessment, approved IFU/label evidence ID | Proposed IR row 26 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C027 GAP-017 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Optional suction catheter BOM applicability; ISO8836 Table2 force by actual OD if needed | Proposed IR row 27 | Completed inspection row / linked controlled worksheet and evidence ID |
| I-C028 GAP-020 | NOT AVAILABLE FOR VERIFICATION: no product-matching Endobronchial report; photographed Endotracheal form cannot serve as current baseline. | No verified matching report row / quantitative proof. | ADD / LINK | Existing component appearance Smooth tube, blue/clear subassemblies, bonds, connectors, lumen clear, x-ray line, label, stylet, swivel, seal | Proposed IR row 28 | Completed inspection row / linked controlled worksheet and evidence ID |

# 3. Inspection report changes — manager overview

| REPORT ROW | BEFORE | AFTER | REASON / GAP | LINKED WI CHANGE |
| --- | --- | --- | --- | --- |
| Proposed IR row 01 | Current product-matching row not supplied | Product/form identification: Correct Endobronchial form ID, rev, code, left/right, WI rev, lot/batch and controlled spec | GAP-018 | W-C018 |
| Proposed IR row 02 | Current product-matching row not supplied | Sampling & disposition: Lot quantity, plan ID/rev, inspection level, AQL, code, n, Ac, Re, drawn IDs, disposition | GAP-020 | W-C020 |
| Proposed IR row 03 | Current product-matching row not supplied | Major tracheal OD axis: Nominal from drawing ±0.5 mm; actual in mm; equipment ID | GAP-004 | W-C004 |
| Proposed IR row 04 | Current product-matching row not supplied | Minor tracheal OD axis: Nominal from drawing ±0.5 mm; actual in mm; equipment ID | GAP-004 | W-C004 |
| Proposed IR row 05 | Current product-matching row not supplied | French designated size: Whole number; measured axes / controlled size rule / printed Fr | GAP-004 | W-C004 |
| Proposed IR row 06 | Current product-matching row not supplied | 15 mm conical connectors: ISO5356-1 conformance evidence ID and approved result | GAP-005 | W-C005 |
| Proposed IR row 07 | Current product-matching row not supplied | Inflation inlet: ISO80369-7 L1 male Luer-slip compatibility record | GAP-005 | W-C005 |
| Proposed IR row 08 | Current product-matching row not supplied | Tracheal cuff leakage: 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 | GAP-006 | W-C006 |
| Proposed IR row 09 | Current product-matching row not supplied | Bronchial cuff leakage: 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 | GAP-006 | W-C006 |
| Proposed IR row 10 | Current product-matching row not supplied | Tracheal cuff diameter: ISO5361 Annex B; claimed mm; measured mean mm; ±15% | GAP-007 | W-C007 |
| Proposed IR row 11 | Current product-matching row not supplied | Bronchial cuff diameter: ISO5361 Annex B; declared/widest mm; mean mm; ±15% | GAP-007 | W-C007 |
| Proposed IR row 12 | Current product-matching row not supplied | Cuff herniation per cuff: ISO5361 Annex D; tracheal/bronchial separate observations | GAP-007 | W-C007 |
| Proposed IR row 13 | Current product-matching row not supplied | Cuff sharp edges: Visual and risk file ID; none at attachments/transitions | GAP-008 | W-C008 |
| Proposed IR row 14 | Current product-matching row not supplied | Inflation tube / geometry: ≤3.0mm OD, ≤10% intrusion where applicable, separation angle ≤45°, S3 conditional | GAP-009 | W-C009 |
| Proposed IR row 15 | Current product-matching row not supplied | Pilot function & deflation: Indication of inflation/deflation, unobstructed intentional deflation | GAP-009 | W-C009 |
| Proposed IR row 16 | Current product-matching row not supplied | Marking exposure durability: Exposures/duration, 1m±10mm, 215±5lx, acuity1, 15s wet-rub result | GAP-011 | W-C011 |
| Proposed IR row 17 | Current product-matching row not supplied | Fr / Br markings: Whole Fr; Br marked OD; Annex B actual mm | GAP-012 | W-C012 |
| Proposed IR row 18 | Current product-matching row not supplied | Effective IDs / both lumina: ISO16628 Annex C; tracheal and bronchial actual and claimed mm | GAP-013 | W-C013 |
| Proposed IR row 19 | Current product-matching row not supplied | Tip-origin length marks: 2cm gradations measured from bronchial tip | GAP-012 | W-C012 |
| Proposed IR row 20 | Current product-matching row not supplied | Side marking: Right or left matches configuration | GAP-012 | W-C012 |
| Proposed IR row 21 | Current product-matching row not supplied | Bronchial color: Bronchial machine end + cuff + pilot each blue | GAP-014 | W-C014 |
| Proposed IR row 22 | Current product-matching row not supplied | Pack/insert content: Description, major and minor axes, tracheal and bronchial cuff diameters | GAP-015 | W-C015 |
| Proposed IR row 23 | Current product-matching row not supplied | Pack sterile link: Pack sealing and sterile-release references where applicable; ISO18190 review pending | GAP-010 | W-C010 |
| Proposed IR row 24 | Current product-matching row not supplied | External general/material/design technical file: ISO18190 Clauses 4–6 applicability and evidence ID; no implicit PASS | GAP-001 | W-C001 |
| Proposed IR row 25 | Current product-matching row not supplied | Breathing-gas pathway biocompatibility: ISO18562-1 evaluation, approved results and reviewer | GAP-003 | W-C003 |
| Proposed IR row 26 | Current product-matching row not supplied | Manufacturer information / IFU approval: ISO18190 Clause9 assessment, approved IFU/label evidence ID | GAP-010 | W-C010 |
| Proposed IR row 27 | Current product-matching row not supplied | Optional suction catheter: BOM applicability; ISO8836 Table2 force by actual OD if needed | GAP-017 | W-C017 |
| Proposed IR row 28 | Current product-matching row not supplied | Existing component appearance: Smooth tube, blue/clear subassemblies, bonds, connectors, lumen clear, x-ray line, label, stylet, swivel, seal | GAP-020 | W-C020 |

# 4. Test method / verification change register

| CHANGE / REQ | EXISTING TEST METHOD | PROBLEM | ACTION | REVISED TEST METHOD | ACCEPTANCE CRITERIA | EVIDENCE |
| --- | --- | --- | --- | --- | --- | --- |
| T-C001 / R012 | WI pp.3,7–8 cuff immersion method present | No cuff-specific recorded endpoints | KEEP / MODIFY | Record each cuff separately, inflation endpoint, pressure or diameter, ≥10s fully submerged bubble-free result | No bubbles ≥10s | Cuff leakage worksheet |
| T-C002 / R013, R014 | WI cites Annex C for diameter | Wrong Annex and no conditioning/readings | MODIFY | ISO5361 Annex B: 2.0±0.1kPa, 23±2°C, 5min, four 45° readings, mean mm 2dp | ±15% of declared tracheal and bronchial cuff diameter | Cuff diameter worksheet |
| T-C003 / R015 | WI cites Annex E for herniation | Wrong Annex | MODIFY | ISO5361 Annex D; 24h (40±1°C), weight 100g ≥60s, monitor deflation ≥10s; validate product fixture applicability | Neither cuff reaches corresponding distal edge | Cuff herniation protocol |
| T-C004 / R031 | Generic OD visual/dimensional inspection | No ISO Annex B five-wrap method | ADD | ISO16628 Annex B five closely spaced spiral wraps, OD = length/(5π), round up to next 0.5mm | Br marking based on prescribed rounded value | Bronchial OD worksheet |
| T-C005 / R032 | Generic inside diameter dimensional mention | No effective functional lumen ID | ADD | ISO16628 Annex C three-bead string, 100±5g mass, D calibration and 0.2mm increments; two lumina | Actual max passing ID, expressed downward to nearest 0.2mm | Bead method worksheet |
| T-C006 / R028, R029 | Wet-cloth rub described | Exposure conditioning not identified | MODIFY | ISO16628 §9.2.1 IFU/risk-selected exposure then acuity 1, 1m±10mm, 215±5lx, 15s wet rub | Legible at end | Marking validation record |
| T-C007 / R009, R010, R021 | Connector visual fit only | Normative ISO connector methods not verified | LINK | Obtain ISO5356-1/ISO80369-7 criteria and approved compatibility fixtures; technical-file verification | Pass according to externally approved normative criteria | Connector qualification report |
| T-C008 / R017, R018, R019, R020, R022, R023, R024 | Qualitative inflating tube/valve observations | No 5361 §5.6 coverage | ADD | Measure/verify all applicable 5361 §5.6 dimensions/angles, pilot indication, and deflation with conditional S3 cases | Each specified limit met; pending design inputs | Inflation functional worksheet |
| T-C009 / R041 | WI displays ISO8836 Table2 pull forces | Applicability and Annex B method missing | LINK / ADD | If actual qualifying suction catheter: Annex B 23±2°C/50±20%RH 1h, pull 200±20mm/min, Table2 force by OD | ≥5/15/20N by OD category when applicable | Optional component bonding record |

# 5. Other controlled-document change register

| CHANGE | CURRENT / STATUS | PROBLEM | ACTION | PROPOSED CONTROL | RESPONSIBLE DOCUMENT | EVIDENCE |
| --- | --- | --- | --- | --- | --- | --- |
| O-C001 | Product specification and drawings not provided | Nominal axes/cuff/Br/ID absent | LINK / ADD | Approve variant-specific nominal dimensions and claims | Engineering specification | Controlled drawing+approval |
| O-C002 | Referenced ISO standards partially absent | ISO18190/18562/5356-1/80369-7 details not verified | LINK | Obtain applicable original and create clause matrix | Normative reference library | Verified copies and matrix |
| O-C003 | Sampling SOP-8-2-6 not supplied | No validated n/Ac/Re by lot | LINK | Add controlled plan link and lot-level outcome | Sampling Plan / SOP | Approved plan and actual batch sheet |
| O-C004 | Approved marking/pack artwork not supplied | No line-by-line matching | LINK | Align Fr/Br/2 IDs/cuffs/axes/side to actual label | Device labelling master | Approved artwork |
| O-C005 | Sterilization file not supplied | No operational sterile-release proof | LINK | Approval of sterile route under QMS and ISO18190 | Sterilization validation / release | Run and release records |
| O-C006 | Unknown suction catheter BOM applicability | ISO8836 may not apply to finished double-lumen product | LINK / MODIFY | Document component classification and required pull force route | Approved BOM / supplier verification | Component evidence |
| O-C007 | Old Endotracheal form photo only | Report product identity mismatch | LINK | Retrieve exact controlled Endobronchial Report and perform true baseline comparison | Document Control report library | Verified report number / revision |
| O-C008 | WI p.4 duplicate 5.1.14 / 5.1.15 | Ambiguous section links | MODIFY | Unique numbering before QMS issue | WI revision record | Approved change history |

# 6. Linked references and release gates

Each W-C/ I-C/ T-C/ O-C item remains Proposed / Open pending source document reconciliation, approval and objective evidence. Where Source WI location is approximate, it reflects the scanned PRINTED WI page and readable headings. Where source report content is unavailable, the original text is not invented. ISO requirements are paraphrased, not reproduced as large copyrighted passages.

| GATE | PASS CONDITION | CURRENT STATUS |
| --- | --- | --- |
| QG1 | All ISO16628 source references verified | Completed for directly available ISO16628 clauses; externally invoked requirements pending |
| QG2 | WI original citations matched | Scanned 11 pp inspected; original numbering issues found |
| QG3 | Current report actual row/Rev comparison | BLOCKED — matching Endobronchial form missing |
| QG4 | Quantitative criteria / source limits | PARTIAL — drawings and reference standards missing |
| QG5 | Internal draft WI / IR cross-links | PROPOSED / not QMS approved |
| QG6 | Reviewed test records and batch evidence | NOT AVAILABLE FOR VERIFICATION |
| QG7 | QA/QC manager final approval | NOT OBTAINED |