QC SYSTEMS / MEDICAL DEVICES

# ISO Requirement Closure Register

EVIDENCE GATE / REQUIREMENT-LEVEL TRACE | NO VERIFIED CLOSED CLAIMS

| Product | Endobronchial Tube (double-lumen Tracheobronchial Tube) |
| --- | --- |
| Primary standard | ISO 16628:2022 (second edition) |
| Existing work instruction | WI-8-2-2-06-1 / Rev 00 / 05 Dec 2023 |
| Matching inspection report | NOT VERIFIED — supplied form is Endotracheal Tube |
| Document identifier | QC-ISO-CLOSURE-DLT-2026-10-DRAFT |
| Status | CONTROLLED DRAFT FOR TECHNICAL REVIEW — NOT APPROVED |
| Issue date | 09 October 2026 |

Important limitation: source files are photographed scans and standards; no approved lot test data, drawings, ISO 18190 copy, ISO 18562-1 report or matching Endobronchial report was supplied. No operational requirement is claimed closed.

Register design: an approved document/control may support documentary coverage, but a completed, applicable and traceable test/technical file is required for operational closure. The current draft is not itself evidence. Existing standards incorporated from ISO 18190 cannot be declared fully extracted until the standard is provided.

| Requirement ID | ISO clause | ISO Says (paraphrase) | Related gap / change | Documentary coverage | Actual verified evidence | Operational closure status | Next closure action |
| --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-001 | 16628 §4.1 | ISO 18190:2016 clause 4 incorporated | GAP-002 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-002 | 16628 §4.2 | Alternative type tests require validated equivalence if used | GAP-002 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Evidence | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-003 | 16628 §5.1 | ISO 18190 clause 5 incorporated | GAP-002 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-004 | 16628 §5.2 | ISO 18562-1 gas-pathway biological evaluation | GAP-003 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Evidence | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-005 | 16628 §6.1 | ISO 18190 clause 6 incorporated | GAP-002 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-006 | 16628 §6.2 | Whole-number Fr based on tracheal major/minor average | GAP-004 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-007 | 16628 §6.3 | Both tracheal outside axes ±0.5 mm | GAP-004 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-008 | 16628 §6.4.1 | 15 mm machine-end cone and ISO 5356-1 evidence | GAP-007 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Evidence | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-009 | 16628 §6.4.2 | L1 male Luer-slip compatible cuff inlet | GAP-008 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Evidence | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-010 | 16628 §6.4.3 | Bronchial bevel recommended (“should”) | GAP-018 | PROPOSED DRAFT ONLY | None supplied / unverified | Advisory — Risk Review | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-011 | 16628 §6.5.1 | Cuff leak proof at specified limiting endpoint, no bubbles ≥10s | GAP-009 | PROPOSED DRAFT ONLY | None supplied / unverified | Documentary Covered — Operational Evidence Pending (DRAFT only) | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-012 | 16628 §6.5.2 | Both cuff diameters ±15%, Annex B | GAP-010 | PROPOSED DRAFT ONLY | None supplied / unverified | Documentary Covered — Operational Evidence Pending (DRAFT only) | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-013 | 16628 §6.5.3 | Herniation per Annex D, no protrusion | GAP-011 | PROPOSED DRAFT ONLY | None supplied / unverified | Documentary Covered — Operational Evidence Pending (DRAFT only) | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-014 | 16628 §6.5.4 | Cuff transition has no sharp edges; risk file review | GAP-012 | PROPOSED DRAFT ONLY | None supplied / unverified | Documentary Covered — Operational Evidence Pending (DRAFT only) | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-015 | 16628 §6.6 | ISO 5361 §5.6 inflation systems | GAP-013 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-016 | 16628 §6.7 | Bronchial margin of safety design “should”/risk review | GAP-018 | PROPOSED DRAFT ONLY | None supplied / unverified | Advisory — Risk Review | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-017 | 16628 §7 | ISO 18190 clause 7 for sterile products | GAP-002, GAP-022 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-018 | 16628 §8 | ISO 18190 clause 8 packaging | GAP-002, GAP-022 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-019 | 16628 §9.1 | ISO 18190 clause 9 information requirements | GAP-002 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-020 | 16628 §9.2.1 | Marking use exposure, rub, lighting and legibility | GAP-014 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-021 | 16628 §9.2.2(a) | Fr size on tube | GAP-015 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-022 | 16628 §9.2.2(b) / Annex B | Br outside diameter on tube, five-thread OD method | GAP-005, GAP-015 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-023 | 16628 §9.2.2(c) / Annex C | Both effective lumen IDs; three-bead method | GAP-006, GAP-015 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-024 | 16628 §9.2.2(d) | 2 cm mark spacing from bronchial tip | GAP-015 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-025 | 16628 §9.2.2(e) | Right or left tube marking | GAP-015 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-026 | 16628 §9.3 | Bronchial end/cuff/pilot all blue | GAP-016 | PROPOSED DRAFT ONLY | None supplied / unverified | Documentary Covered — Operational Evidence Pending (DRAFT only) | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-027 | 16628 §9.4(a) | Pack/insert description | GAP-017 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Evidence | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-028 | 16628 §9.4(b) | Pack/insert major and minor axis dimensions | GAP-017 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-029 | 16628 §9.4(c) | Pack/insert tracheal and bronchial cuff diameters | GAP-017 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-030 | 8836 §8.1.1 (conditional) | Suction catheter connection detachment force | GAP-019 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |
| REQ-031 | 8836 applicable clauses (conditional) | Suction catheter material/shaft/other scope | GAP-020 | PROPOSED DRAFT ONLY | None supplied / unverified | Open — Technical Input | Obtain approved records; review source, test and configuration; QA sign-off |

# Gap corrective-action closure register

| Gap ID | P | Required approved control & proof | Verification route | Evidence actually verified | Closure status | Owner |
| --- | --- | --- | --- | --- | --- | --- |
| GAP-001 | P2 | Approved product-specific form ID and revision; master document list; QA change-control and lot-linked completed form. Verify exact identities and approved scope. | Document control / QA approval | NONE VERIFIED | Open — Approval | Document Control / QA |
| GAP-002 | P1 | Controlled copy of ISO 18190:2016; approved applicability/trace matrix; linked test, risk, sterile barrier and labelling evidence for each incorporated clause. | Cross-functional regulatory / qualification | NONE VERIFIED | Open — Technical Input | QA / Regulatory |
| GAP-003 | P1 | Approved ISO 18562-1 evaluation and testing/risk rationale; relevant material inventory; configuration linkage; QA approval and validity/review date. | Qualification / change control | NONE VERIFIED | Open — Evidence | QA / Regulatory |
| GAP-004 | P2 | Approved size-specific drawing, master dimensional specification, calibrated results for both axes, Fr designation mapping, resolved written technical interpretation and lot/qualification linkage. | Incoming / routine dimensional + qualification | NONE VERIFIED | Open — Technical Input | Engineering / QC |
| GAP-005 | P1 | Completed Annex B worksheet including winding length, formula, raw/rounded OD, actual Br mark, approved drawing/label artwork, calibration records and unit/sample IDs. | Qualification with controlled production verification | NONE VERIFIED | Open — Technical Input | QC / Engineering |
| GAP-006 | P1 | Annex C test records for BOTH lumens, bead/weight verification, marked diameter artwork, calibrated apparatus and approval to use any defined single-set method. | Type / qualification; controlled routine check | NONE VERIFIED | Open — Technical Input | QC / Engineering |
| GAP-007 | P1 | Approved 15 mm connector test report to ISO 5356-1, controlled drawing, purchased-part approval, change-control scope and assembly inspection record. | Qualification + component/in-process | NONE VERIFIED | Open — Evidence | Engineering / QC |
| GAP-008 | P2 | ISO 80369-7-based compatibility documentation, approved connector components/drawings and functional records for both inflation circuits. | Qualification / supplier and functional | NONE VERIFIED | Open — Evidence | Engineering / Supplier Quality |
| GAP-009 | P1 | Separate tracheal and bronchial cuff test readings, pressure/diameter, elapsed seconds, leak finding, sample ID, equipment traceability and approved sampling plan. | In-process / final per approved control plan | NONE VERIFIED | Documentary Covered — Operational Evidence Pending (DRAFT only) | QC / Production |
| GAP-010 | P1 | Completed Annex B measurements for both cuffs; marked values; calculation worksheet; dimensional specification and calibration evidence. | Qualification / periodic or lot per risk plan | NONE VERIFIED | Documentary Covered — Operational Evidence Pending (DRAFT only) | QC |
| GAP-011 | P1 | Annex D qualified test procedure, tracheal and bronchial results, photographs if relevant, sample IDs and equipment/calibration. | Qualification / periodic per approved risk plan | NONE VERIFIED | Documentary Covered — Operational Evidence Pending (DRAFT only) | QC / Engineering |
| GAP-012 | P2 | Approved risk assessment/design review of cuff transitions; controlled visual criteria; sample inspection findings and disposition for failures. | Risk management + routine visual | NONE VERIFIED | Documentary Covered — Operational Evidence Pending (DRAFT only) | QA / QC |
| GAP-013 | P1 | Controlled drawing including S3 geometry, design-dependent applicability decision, dimensional measurements, deflation functional records, connector interface evidence. | Qualification and defined functional sampling | NONE VERIFIED | Open — Technical Input | Engineering / QC |
| GAP-014 | P2 | Approved use-exposure matrix/IFU; duration rationale; legibility/lighting records; qualification test data and revision/configuration coverage. | Qualification and change control | NONE VERIFIED | Open — Technical Input | QA / Engineering |
| GAP-015 | P1 | Approved left/right artwork; dimension/metered mark verification; actual printed labels; Annex B/C evidence and traceable finished-unit inspections. | Routine visual / qualification metrology | NONE VERIFIED | Open — Technical Input | QC / Engineering |
| GAP-016 | P2 | Completed colour-code inspection rows for all three bronchial features; approved colour standard/artwork and lot linkage. | Routine final visual | NONE VERIFIED | Documentary Covered — Operational Evidence Pending (DRAFT only) | QC |
| GAP-017 | P2 | Approved product label/IFU artworks for all SKUs; label content checklist; packaging print verification; verified ISO 18190 §9 assessment. | Packaging / label approval & routine checks | NONE VERIFIED | Open — Technical Input | QA / Regulatory |
| GAP-018 | P2 | Design drawings and approved risk-management file addressing left/right geometry and rationale. Advisory status clearly distinguished from mandatory clauses. | Design-control assessment | NONE VERIFIED | Open — Evidence | Engineering / QA |
| GAP-019 | P1 | Approved BOM/supplied-catheter decision, applicable ISO 8836 Annex B worksheets and Table 2/3 comparison, calibration, disposition of 5.0 mm case. | Incoming / qualification or supplier | NONE VERIFIED | Open — Technical Input | QC / Supplier Quality |
| GAP-020 | P2 | Product BOM/accessory drawing; supplier conformance and test qualification; ISO 8836 applicability matrix; approved incoming controls. | Supplier / qualification | NONE VERIFIED | Open — Technical Input | QA / Supplier Quality |
| GAP-021 | P2 | Approved SOP-8-2-6 and inspection plan, lot/sample logs, Ac/Re decision records and control-plan approval for nonroutine tests. | Routine / in-process / QA plan | NONE VERIFIED | Open — Technical Input | QC / QA |
| GAP-022 | P1 | Sterilization validation/lot release, sterile barrier qualification and inspection, packaging lot identity, full QC disposition and QA approval. ISO 18190 text also needed. | Sterile release / packaging validation | NONE VERIFIED | Open — Technical Input | QA / Production |
| GAP-023 | P3 | Approved document-change record, corrected unique section numbers, reference check, obsolete-document withdrawal and QA sign-off. | Document control | NONE VERIFIED | Open — Approval | Document Control / QA |

# How to decide closure

• Verified Closed: source clause and applicability established; controls approved/implemented; relevant objective evidence complete and passed; scope, lot and configuration verified; QA acceptance recorded.

• Documentary Covered — Operational Evidence Pending: controlled draft/change describes technical requirements but approvals and/or performed test results remain pending.

• Open — Technical Input: an ISO cross-reference, drawing, acceptance value, approved test method or applicability decision is unavailable.

• Open — Approval: final controlled record code, technical approval, change-control and release decisions not yet signed.

• N/A with Justification: documented product/configuration exclusion with signed decision; do not confuse advisory “should” recommendations with nonapplicable mandatory requirements.

# Mandatory outstanding inputs

1. ISO 18190:2016 full controlled standard and clause crosswalk. 2. Matching approved Endobronchial Tube Inspection & Test Report with complete identifier/revision. 3. Approved size- and side-specific drawings, label artwork and BOM. 4. Approved sampling SOP/control plan. 5. ISO 18562-1 technical evaluation and risk file. 6. Actual ISO 16628 Annex B/C, cuff leak/diameter/herniation, connector and marking qualification records. 7. Conditional suction catheter supplier scope/tests. 8. Sterilization, barrier and final release documentation.