# ENDOBRONCHIAL TUBE

ISO 16628:2022 | QC Manager Compliance Review & Proposed Reconstruction

STATUS: PROPOSED CONTROLLED DRAFT  |  NO SIGNATURES OR RELEASE AUTHORIZATION

Source basis: ISO 16628:2022; ISO 5361:2016 invoked clauses/annexes; conditional ISO 8836:2014; scanned WI Rev 00 and non-matching inspection report photo.

# 1. Executive summary

CURRENT ASSESSMENT: PARTIAL / NOT VERIFIED FOR PRODUCT CONFORMITY. The uploaded 11-page Endobronchial WI addresses selected dimensional, cuff, colour and visual controls. Two normative cuff-annex citations are incorrect (ISO 5361 Annex C/E instead of B/D), the uploaded inspection form depicts a different product (Endotracheal Tube), and referenced ISO 18190 / ISO 18562 source texts and supporting results are not supplied. A proposed WI/report/closure framework is prepared, but the documents and product remain unapproved/unverified.

Applicability rationale: ISO 16628:2022 directly governs the Double-Lumen Tracheobronchial Tube. ISO 5361:2016 excludes endobronchial tubes as a primary product standard but is invoked selectively by ISO 16628 (including §§5.6, Annex B and Annex D). ISO 8836:2014 is restricted to an actual suction-catheter subcomponent if the product BOM confirms this classification; its wholesale application to the double-lumen product would be technically unsound.

| INDICATOR | VERIFIED STATUS / BASIS |
| --- | --- |
| Target product | Endobronchial (Tracheobronchial / DLT) |
| Primary ISO | ISO 16628:2022 |
| Existing WI | WI-8-2-2-06-1 Rev 00 (2023-12-05); 11 numbered pages |
| Existing matching report | NOT AVAILABLE FOR VERIFICATION |
| Received other-product report | Endotracheal Tube form, End.pdf PDF p.1 |
| Core requirements (incl. conditional clauses) | 38 |
| Guidance rows / optional component rows | 2 / 2 |
| Audited gaps P1/P2/P3 | 14 / 6 / 0 |
| WI change register | 20 |
| Inspection changes | 28 |
| Documentary specified-control coverage | 24/38 — 63.2% (proposed-draft definitions only) |
| Verified operational closure | NOT ESTABLISHED — no reviewed actual product/lot evidence supplied |
| Technical inputs / externally controlled evidence | 14 normative items open / evidence reviews pending |
| Manager decisions | 7 |

Percentage rule: denominator excludes the two “should”-level guidance controls and two ISO 8836 component-dependent controls. “Documentary specified-control coverage” counts only audit rows whose proposed control is described with sufficient content for drafting, regardless of its unapproved state. It is NOT a product compliance score or percent verified testing. No requirement is VERIFIED CLOSED.

# 2. Document identification & control

| ITEM | RESULT / PROVENANCE |
| --- | --- |
| Main standard | ISO 16628:2022 (2nd edition), provided; PDF 20 pp.; printed 13 pp. |
| Referenced standard ISO 5361 | ISO 5361:2016 (3rd ed.), provided; direct scope excludes Endobronchial Tube (Clause 1); invoked content selected. |
| Supplemental component source | BS EN ISO 8836:2014, provided; suction-catheter subcomponent applicability pending BOM review. |
| WI scan | End.pdf pp.2–12; WI-8-2-2-06-1 Rev 00; 05.12.2023; “Inspection Procedure for Endobronchial Tube.” |
| Incorrectly matched report | End.pdf p.1 “Inspection & Test Report for Endotracheal Tube”; photo has incomplete/cropped identifier; Rev 4 visually apparent, unverified details. |
| Endobronchial report per WI references | WI printed p.11 references an “Inspection and test report for Endobronchial tube” (form identifier appears F-823-T98 in scan); actual controlled form NOT AVAILABLE FOR VERIFICATION. |
| Referenced controlled sampling SOP | SOP-8-2-6 cited by WI p.1; copy and revision not supplied. |
| Product drawings / specifications / BOM / labels / CoA | NOT AVAILABLE FOR VERIFICATION. |
| Referenced ISO 18190:2016, ISO 18562-1, ISO 5356-1, ISO 80369-7 | Referenced contents, current controlled versions, test evidence NOT AVAILABLE FOR VERIFICATION. |
| Approval signatures on historical WI | Signed images visible; no approval inferred for newly proposed documents. |
| Proposed document numbers | For drafting only; Document Control to allocate final controlled form IDs/revisions. |

# 3. Top management findings

| PRIORITY | FINDING | EFFECT / REQUIRED ACTION |
| --- | --- | --- |
| P1 | GAP-001 External general requirements not verified | Acquire controlled ISO 18190:2016 and create clause-by-clause compliance assessment. |
| P1 | GAP-003 Gas pathway biocompatibility evidence missing | Obtain applicable ISO 18562-1 and approved evaluation/test report. |
| P1 | GAP-005 Connector conformity route incomplete | Add technical-file verification and control compatibility evidence. |
| P1 | GAP-006 Cuff leak record inadequately demonstrated | Maintain existing test; require cuff-specific conditions and measurements in controlled record. |
| P1 | GAP-007 Cuff Annex citations incorrect | Correct citations, restore exact method prerequisites and dedicated worksheets. |
| P1 | GAP-008 Sharp-edge risk traceability not explicit | Link risk analysis, documented design check and inspection. |
| P1 | GAP-009 Referenced cuff inflation system requirements incomplete | Add configuration-gated dimensional and functional verification. |
| P1 | GAP-010 ISO 18190 links for sterile supply, packaging and labelling unverified | Acquire ISO 18190; map to sterilization, packaging and label/IFU validation. |
| P1 | GAP-011 Marking exposure validation omitted | Establish exposure selection from IFU and risk file; execute marking durability qualification. |
| P1 | GAP-012 Required shaft markings not explicitly enumerated | Add per-attribute ISO 16628 §9.2.2 fields and numeric verifications. |
| P1 | GAP-013 Effective inside diameter method missing | Add Annex C method/worksheet for both lumina with verified beads and mass. |

# 4. ISO master requirement & current-gap matrix

ISO Says entries below are technical paraphrases tied to PDF page/printed page, not reproduced standard paragraphs. WI location refers to PRINTED WI pagination (End.pdf PDF page = printed p.+1). “Current Report Says” is unavailable for the matching Endobronchial report; the photographed Endotracheal form is not used as evidence of current Endobronchial report fields.

| REQ / REF | ISO Says (verified paraphrase) | Existing WI Says / location | GAP / P | PROPOSED CONTROL / STATUS |
| --- | --- | --- | --- | --- |
| R001 ¦ 4.1 ISO16628 PDF p.8 / printed p.2 | Applicable general requirements of ISO 18190:2016 Clause 4 apply; exact subrequirements cannot be substantiated without that source. | WI pp.1–11: no traceable ISO 18190 Clause 4 matrix. | GAP-001 ¦ P1 | Design / risk management OPEN — TECHNICAL INPUT |
| R002 ¦ 4.2 ISO16628 PDF p.8 / printed p.2 | If an alternative type-test method is selected, equivalent safety must be established and the alternative validated against the specified method. | WI pp.1–11: no procedure for alternative method approval. | GAP-002 ¦ P2 | Validation / technical file DOCUMENTARY COVERED |
| R003 ¦ 5.1 ISO16628 PDF p.8 / printed p.2 | Applicable materials requirements of ISO 18190:2016 Clause 5 apply; detailed obligations require access to the referenced standard. | WI p.1 receiving / CoA check; material-standard compliance not defined. | GAP-001 ¦ P1 | Incoming / Supplier / technical file OPEN — TECHNICAL INPUT |
| R004 ¦ 5.2 ISO16628 PDF p.8 / printed p.2 and p.9 | Tracheobronchial tubes must be evaluated/tested according to ISO 18562-1 for breathing-gas-pathway biocompatibility. | WI pp.1–11: no ISO 18562-1 technical-file reference identified. | GAP-003 ¦ P1 | Biocompatibility qualification OPEN — TECHNICAL INPUT |
| R005 ¦ 6.1 ISO16628 PDF p.9 / printed p.3 | Applicable design requirements of ISO 18190:2016 Clause 6 apply; review the referenced source for actual limits/methods. | WI pp.2–10 design characteristics; no linked ISO18190 review. | GAP-001 ¦ P1 | Design verification OPEN — TECHNICAL INPUT |
| R006 ¦ 6.2 ISO16628 PDF p.9 / printed p.3 | Designated French/Charrière size is expressed as a whole number and derived from average tracheal-segment major/minor outside axis dimensions; its stated ±1 mm phrase uses “should” and needs documented interpretation before use as a mandatory lot-rejection rule. | WI p.2 §5.1.2 includes whole Fr and averages major/minor axis; no actual values recorded. | GAP-004 ¦ P2 | Design / dimensional verification DOCUMENTARY COVERED |
| R007 ¦ 6.3 ISO16628 PDF p.9 / printed p.3 | Major-axis dimension of the tracheal segment outside cross-section is subject to ±0.5 mm tolerance relative to its controlled nominal dimension. | WI p.2 §5.1.3 mentions ±0.5 mm; nominal/drawing not supplied. | GAP-004 ¦ P2 | Dimensional verification OPEN — TECHNICAL INPUT |
| R008 ¦ 6.3 ISO16628 PDF p.9 / printed p.3 | Minor-axis dimension of the tracheal segment outside cross-section is subject to ±0.5 mm tolerance relative to its controlled nominal dimension. | WI p.2 §5.1.3 mentions ±0.5 mm; nominal/drawing not supplied. | GAP-004 ¦ P2 | Dimensional verification OPEN — TECHNICAL INPUT |
| R009 ¦ 6.4.1 ISO16628 PDF p.9 / printed p.3 | Machine-end tracheobronchial connectors must be 15 mm conical male connectors conforming to ISO 5356-1; technical-file conformance is required. | WI p.5 §5.1.25 references ISO 5356-1 without a test/technical-file link. | GAP-005 ¦ P1 | Connector design verification OPEN — TECHNICAL INPUT |
| R010 ¦ 6.4.2 ISO16628 PDF p.9 / printed p.3 | Cuff inflation inlet connectors must be compatible with the L1 male Luer-slip small-bore connector specified in ISO 80369-7; verify in the technical file. | WI pp.3–4 mentions valves and pilot balloons, not verified ISO 80369-7 compatibility. | GAP-005 ¦ P1 | Connector verification / technical file OPEN — TECHNICAL INPUT |
| R011 ¦ 6.5.1 ISO16628 PDF p.9 / printed p.3 | If cuffs are fitted, each cuff must be permanently integrated with the tube and capable of inflation without leakage. | WI pp.3–4 cuff and pilot balloon attachment discussed. | GAP-006 ¦ P1 | Cuff functional verification DOCUMENTARY COVERED |
| R012 ¦ 6.5.1 ISO16628 PDF pp.9–10 / printed pp.3–4 | Cuff leak challenge: inflate up to 9.0 kPa (90 cmH2O) OR to 1.5 times measured cuff diameter, whichever is reached first; seal and disconnect the inflation source, immerse entire system, inspect ≥10 s, and accept only if no bubbles. | WI p.3 §§5.1.5–5.1.6 includes challenge and 10-s bubble check; actual results fields unverified. | GAP-006 ¦ P1 | Routine / approved risk-based strategy DOCUMENTARY COVERED |
| R013 ¦ 6.5.2 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.32 / printed p.26 | Tracheal cuff resting diameter must fall within ±15% of the declared value when measured per ISO 5361:2016 Annex B. | WI p.3 §5.1.7 cites incorrect ISO 5361 Annex C. | GAP-007 ¦ P1 | Verification / approved frequency DOCUMENTARY COVERED |
| R014 ¦ 6.5.2 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.32 / printed p.26 | Bronchial cuff resting diameter must fall within ±15% of declared value when measured per ISO 5361:2016 Annex B; a non-round bronchial cuff must disclose its widest diameter. | WI p.3 §5.1.7 cites Annex C, although wide diameter language is present. | GAP-007 ¦ P1 | Verification / approved frequency DOCUMENTARY COVERED |
| R015 ¦ 6.5.3 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.36 / printed p.30 | Under ISO 5361:2016 Annex D cuff-herniation test, only the cuff under test is inflated and neither cuff may extend beyond nearest patient-end edge of its respective segment. | WI p.3 §5.1.8 cites incorrect ISO 5361 Annex E. | GAP-007 ¦ P1 | Type / qualification verification DOCUMENTARY COVERED |
| R016 ¦ 6.5.4 ISO16628 PDF p.10 / printed p.4 | Cuff attachment transitions and outer surfaces must be free from sharp edges; demonstrate through review of risk-management documentation. | WI p.3 cuff surface mentions general appearance; risk-file link absent. | GAP-008 ¦ P1 | Design risk file / visual control DOCUMENTARY COVERED |
| R017 ¦ 6.6 → 5361 5.6.1 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | Where fitted, inflation-tube OD ≤3.0 mm; point of separation on concave side if curved; inflation lumen intrusion ≤10% of internal diameter; dimensions per referenced table/figures. | WI pp.3–4 describes the inflation tube qualitatively only. | GAP-009 ¦ P1 | Design verification OPEN — TECHNICAL INPUT |
| R018 ¦ 6.6 → 5361 5.6.2 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | Angle between inflation tube and tracheal tube at separation must not exceed 45°. | WI pp.3–4 contains no angular limit. | GAP-009 ¦ P2 | Dimensional / technical file DOCUMENTARY COVERED |
| R019 ¦ 6.6 → 5361 5.6.3 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | Inflation tube must have a pilot balloon and/or another inflation/deflation indicator. | WI p.4 §§5.1.11–5.1.13 includes pilot balloons. | GAP-009 ¦ P2 | Routine / incoming DOCUMENTARY COVERED |
| R020 ¦ 6.6 → 5361 5.6.4 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | Cuff intentional deflation must not be obstructed by inflation tube, valve or other closure device. | WI p.4: valve integrity but no deflation test. | GAP-009 ¦ P1 | Functional verification DOCUMENTARY COVERED |
| R021 ¦ 6.6 → 5361 5.6.5 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | Inflation inlet must accept the applicable male 6% taper fitting under referenced 5361 Clause 5.6.5; reconcile with ISO 16628 Clause 6.4.2 requirement for ISO 80369-7 L1 male compatibility. | WI pp.3–5 mentions valves, with no interface verification. | GAP-005 ¦ P1 | Connector qualification OPEN — TECHNICAL INPUT |
| R022 ¦ 6.6 → 5361 5.6.6 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | If inflation valve/closure is absent, inflation tube free S3 length must be ≥40 mm. | WI pp.3–4: no S3 dimension. | GAP-009 ¦ P2 | Conditional drawing verification DOCUMENTARY COVERED |
| R023 ¦ 6.6 → 5361 5.6.7 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | If closure device exists, S3 ≥10 mm unless pilot balloon and inflation valve are integral. | WI pp.3–4: no S3 dimension or exception decision. | GAP-009 ¦ P2 | Conditional drawing verification DOCUMENTARY COVERED |
| R024 ¦ 6.6 → 5361 5.6.8 ISO16628 PDF p.10 / printed p.4; ISO5361 PDF p.20 / printed p.14 | If inflation-tube separation distance from patient end is marked, actual dimension must equal mark ±10 mm. | WI pp.3–5 does not establish if that distance is marked. | GAP-009 ¦ P2 | Conditional marking verification DOCUMENTARY COVERED |
| R025 ¦ 7 ISO16628 PDF p.11 / printed p.5 | Sterile-supplied tracheobronchial tubes must satisfy ISO 18190:2016 Clause 7, details not independently verifiable without ISO 18190. | WI p.10 §9 refers to post-sterilization carton chamber stamps only. | GAP-010 ¦ P1 | Sterilization validation / release OPEN — TECHNICAL INPUT |
| R026 ¦ 8 ISO16628 PDF p.11 / printed p.5 | Packaging for tracheobronchial tubes must meet ISO 18190:2016 Clause 8 requirements, which must be retrieved and verified. | WI p.10 §8 checks sealed pouch and label but lacks referenced technical requirements. | GAP-010 ¦ P1 | Packaging qualification OPEN — TECHNICAL INPUT |
| R027 ¦ 9.1 ISO16628 PDF p.12 / printed p.6 | Applicable manufacturer-information requirements of ISO 18190:2016 Clause 9 apply; source must be obtained for full gap analysis. | WI pp.5,10 names marking/pack labels without controlled IFU review. | GAP-010 ¦ P1 | Label / IFU regulatory review OPEN — TECHNICAL INPUT |
| R028 ¦ 9.2.1 ISO16628 PDF p.12 / printed p.6 | Markings must remain durable/legible after representative exposure to materials they contact in intended use for the expected exposure duration, including listed drugs/chemicals and applicable simulated fluids. | WI p.5 §5.1.26 includes reading/rubbing step but omits representative-use chemical conditioning. | GAP-011 ¦ P1 | Marking validation / qualification OPEN — TECHNICAL INPUT |
| R029 ¦ 9.2.1 ISO16628 PDF p.12 / printed p.6 | After exposure, check marking legibility at 1 m ±10 mm with visual acuity 1 and 215 ±5 lx; hand-rub without undue pressure for 15 s using distilled-water-wet cloth; remain legible. | WI p.5 §5.1.26 partly includes 1 m, 215 lx, 15 s and water cloth, but not acuity/pre-exposure. | GAP-011 ¦ P1 | Marking validation / qualification DOCUMENTARY COVERED |
| R030 ¦ 9.2.2(a) ISO16628 PDF p.12 / printed p.6 | Device shaft must display whole-number Charrière/French size designation. | WI p.2 §5.1.2 explains Fr but no label-specific inspection field. | GAP-012 ¦ P2 | Visual / labelling DOCUMENTARY COVERED |
| R031 ¦ 9.2.2(b) ISO16628 PDF p.12 / printed p.6; Annex B PDF p.16 / printed p.10 | Device shaft must state bronchial segment outside diameter in mm with prefix Br; verify using ISO 16628 Annex B. | WI pp.5,9 marks generally; not the Br numeric measurement/mark linkage. | GAP-012 ¦ P1 | Dimensional qualification and visual DOCUMENTARY COVERED |
| R032 ¦ 9.2.2(c) ISO16628 PDF p.12 / printed p.6; Annex C PDF pp.17–18 / printed pp.11–12 | Device shaft must show effective inside diameters (mm) for BOTH tracheal and bronchial lumen; determine using ISO 16628 Annex C bead method. | WI p.2 dimensional table only refers generic ID/OD, no effective ID bead method. | GAP-013 ¦ P1 | Type / qualification verification DOCUMENTARY COVERED |
| R033 ¦ 9.2.2(d) ISO16628 PDF p.13 / printed p.7 | Length graduations appear at 2 cm intervals, measured from bronchial tip, on the machine end of the tracheal segment. | WI p.2 notes accurate graduation only; origin/interval unspecified. | GAP-012 ¦ P2 | Visual / dimensional DOCUMENTARY COVERED |
| R034 ¦ 9.2.2(e) ISO16628 PDF p.13 / printed p.7 | Device shaft must be labelled right or left, matching the manufactured configuration. | WI p.5 §5.1.27 discusses colors; right/left identification missing. | GAP-012 ¦ P1 | Visual / label verification DOCUMENTARY COVERED |
| R035 ¦ 9.3 ISO16628 PDF p.13 / printed p.7 | Entire machine end of bronchial segment, bronchial cuff and corresponding pilot balloon must be blue. | WI p.5 §5.1.27 explicitly describes bronchial blue. | GAP-014 ¦ P2 | Visual / routine DOCUMENTARY COVERED |
| R036 ¦ 9.4(a) ISO16628 PDF p.13 / printed p.7 | Individual package or insert must describe package contents. | WI p.10 §8 requires label product name/code/lot etc, but does not explicitly check contents description. | GAP-015 ¦ P2 | Label / packaging DOCUMENTARY COVERED |
| R037 ¦ 9.4(b) ISO16628 PDF p.13 / printed p.7 | Individual package or insert must include the major-axis and minor-axis dimensions of the tracheal segment. | WI p.10 §8 does not state major/minor axis on packaging. | GAP-015 ¦ P1 | Label artwork approval DOCUMENTARY COVERED |
| R038 ¦ 9.4(c) ISO16628 PDF p.13 / printed p.7 | Individual package or insert must state tracheal and bronchial cuff diameters as declared for Clause 6.5.2. | WI p.10 §8 does not state both cuff diameters on packaging. | GAP-015 ¦ P1 | Label artwork approval DOCUMENTARY COVERED |
| R039 ¦ 6.4.3 (should) ISO16628 PDF p.9 / printed p.3 | Bronchial patient end is recommended to have a bevel; this is guidance wording, not a standalone “shall” obligation. | WI p.5 §5.1.24 includes bronchial tip curve visual inspection. | GAP-016 ¦ P3 | Design evaluation DOCUMENTARY COVERED |
| R040 ¦ 6.7 (should) ISO16628 PDF p.10 / printed p.4 | Bronchial segment length, cuff location/length, bevel and right upper-lobe considerations are recommended inputs to design margin-of-safety assessment. | WI p.5 mentions tip curve but no margin-of-safety risk reference. | GAP-016 ¦ P2 | Risk management / design validation DOCUMENTARY COVERED |
| R041 ¦ 8836:2014 7.3.2, 8.1.1 ISO8836 PDF p.20 / ISO printed p.9 and PDF p.23 / printed p.12 | IF a separate suction catheter component qualifies as a suction catheter under ISO 8836, verify securely attached connector per Annex B and Table 2: OD 1.33–2.67 mm: ≥5 N; 3.00–4.67 mm: ≥15 N; >5 mm: ≥20 N. | WI p.5 §5.1.29 / p.10 includes Table 2 values, but method, OD category and component scope unclear. | GAP-017 ¦ P2 | Conditional supplier / component verification OPEN — TECHNICAL INPUT |
| R042 ¦ 8836:2014 Clause 7 ISO8836 PDF pp.19–22 / printed pp.8–11 | IF an actual respiratory suction catheter is part of the defined configuration, confirm further relevant ISO 8836 design/connection requirements against its complete component classification. | WI p.5 §5.1.28 lists suction catheter visual only. | GAP-017 ¦ P2 | Component qualification OPEN — TECHNICAL INPUT |

# 5. Detailed finding cards — Before / Problem / Action / After / Evidence

## GAP-001 | External general requirements not verified [P1]

ISO / requirements: R001, R003, R005

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Referenced ISO 18190:2016 Clauses 4, 5 and 6 require detailed source extraction; a full check is not possible from the uploaded files.

Required action: Acquire controlled ISO 18190:2016 and create clause-by-clause compliance assessment.

Evidence: Objective technical file / controlled results; owner: Engineering / QA

## GAP-002 | Alternative test controls not defined [P2]

ISO / requirements: R002

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: No documented method equivalence process found for substituted ISO type methods.

Required action: Define validation/approval gate before any substitution.

Evidence: Objective technical file / controlled results; owner: QA / Engineering

## GAP-003 | Gas pathway biocompatibility evidence missing [P1]

ISO / requirements: R004

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: ISO 16628 expressly requires ISO 18562-1 evaluation; original requirements and evidence not supplied.

Required action: Obtain applicable ISO 18562-1 and approved evaluation/test report.

Evidence: Objective technical file / controlled results; owner: QA / Regulatory

## GAP-004 | Size and axis evidence not fully measurable [P2]

ISO / requirements: R006, R007, R008

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI gives size/tolerance text but not approved nominal axes, raw measured values or defined conversion governance.

Required action: Introduce dimensional worksheet and require approved size/design drawing.

Evidence: Objective technical file / controlled results; owner: Engineering / QC

## GAP-005 | Connector conformity route incomplete [P1]

ISO / requirements: R009, R010, R021

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Generic connector visual/fit checks do not establish ISO 5356-1 and ISO 80369-7 conformance; referenced-source criteria are unavailable.

Required action: Add technical-file verification and control compatibility evidence.

Evidence: Objective technical file / controlled results; owner: Engineering / QA

## GAP-006 | Cuff leak record inadequately demonstrated [P1]

ISO / requirements: R011, R012

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI method substantially matches ISO 16628 but corresponding Endobronchial report/actual tests were not supplied.

Required action: Maintain existing test; require cuff-specific conditions and measurements in controlled record.

Evidence: Objective technical file / controlled results; owner: QC

## GAP-007 | Cuff Annex citations incorrect [P1]

ISO / requirements: R013, R014, R015

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI directs cuff diameter to ISO 5361 Annex C and cuff herniation to Annex E instead of B and D.

Required action: Correct citations, restore exact method prerequisites and dedicated worksheets.

Evidence: Objective technical file / controlled results; owner: QC / QA

## GAP-008 | Sharp-edge risk traceability not explicit [P1]

ISO / requirements: R016

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Cuff/shaft joint visual requirement lacks linkage to risk-management file.

Required action: Link risk analysis, documented design check and inspection.

Evidence: Objective technical file / controlled results; owner: Engineering / QA

## GAP-009 | Referenced cuff inflation system requirements incomplete [P1]

ISO / requirements: R017, R018, R019, R020, R022, R023, R024

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Tube/valves judged qualitatively without full ISO 5361:2016 5.6 dimension/functional clauses.

Required action: Add configuration-gated dimensional and functional verification.

Evidence: Objective technical file / controlled results; owner: Engineering / QC

## GAP-010 | ISO 18190 links for sterile supply, packaging and labelling unverified [P1]

ISO / requirements: R025, R026, R027

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: The WI notes sterilization and sealed pouches but not entire ISO 18190 referenced clauses.

Required action: Acquire ISO 18190; map to sterilization, packaging and label/IFU validation.

Evidence: Objective technical file / controlled results; owner: QA / Document Control

## GAP-011 | Marking exposure validation omitted [P1]

ISO / requirements: R028, R029

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI includes final readability/rubbing but not exposure media and duration before rub.

Required action: Establish exposure selection from IFU and risk file; execute marking durability qualification.

Evidence: Objective technical file / controlled results; owner: QA / Engineering

## GAP-012 | Required shaft markings not explicitly enumerated [P1]

ISO / requirements: R030, R031, R033, R034

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Br outer diameter, two effective lumen IDs, correct side and 2-cm graduations are not controlled as a complete checklist.

Required action: Add per-attribute ISO 16628 §9.2.2 fields and numeric verifications.

Evidence: Objective technical file / controlled results; owner: QC / Production

## GAP-013 | Effective inside diameter method missing [P1]

ISO / requirements: R032

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Generic tube-ID measurement cannot substitute for the specified three-bead functional effective ID test.

Required action: Add Annex C method/worksheet for both lumina with verified beads and mass.

Evidence: Objective technical file / controlled results; owner: Engineering / QC

## GAP-014 | Correct colour control needs explicit retained traceability [P2]

ISO / requirements: R035

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Existing blue bronchial color instruction is present; preserve it as a distinct check.

Required action: KEEP existing color rule; map to ISO §9.3 and specific report row.

Evidence: Objective technical file / controlled results; owner: QC

## GAP-015 | Package marking list incomplete [P1]

ISO / requirements: R036, R037, R038

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI p.10 lists generic label items but not major/minor axes or both cuff diameters.

Required action: Add approved label artwork verification and matching fields.

Evidence: Objective technical file / controlled results; owner: Document Control / QC

## GAP-016 | Advisory design recommendations need controlled rationale [P2]

ISO / requirements: R039, R040

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Bronchial bevel and design margin-of-safety are “should” recommendations; current WI only covers tip appearance.

Required action: Record applicability/design rationale in risk file, not as unsupported ISO “shall” limit.

Evidence: Objective technical file / controlled results; owner: Engineering

## GAP-017 | Suction catheter component applicability not established [P2]

ISO / requirements: R041, R042

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI refers to ISO 8836 Table 2 yet available records do not show if suction catheter is a separate qualifying component.

Required action: Confirm BOM/configuration and apply ISO 8836 selectively to the actual suction catheter.

Evidence: Objective technical file / controlled results; owner: Engineering / Supplier Quality

## GAP-018 | Mismatched report and unavailable product-specific form [P1]

ISO / requirements: Supporting document-control / record issue, not additional ISO shall

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: End.pdf p.1 shows an Endotracheal Tube report; Endobronchial report referenced by WI p.11 not supplied.

Required action: Retrieve current controlled Endobronchial report, verify revision and rebuild product-specific report.

Evidence: Objective technical file / controlled results; owner: Document Control / QC

## GAP-019 | WI section numbering ambiguous [P2]

ISO / requirements: Supporting document-control / record issue, not additional ISO shall

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: Existing WI p.4 repeats identifiers 5.1.14 and 5.1.15 for different controls; potential traceability issue.

Required action: Renumber proposed WI uniquely; keep original p/section provenance in register.

Evidence: Objective technical file / controlled results; owner: Document Control

## GAP-020 | Sampling and release approval basis unverified [P1]

ISO / requirements: Supporting document-control / record issue, not additional ISO shall

Existing / source basis: Current WI references in linked requirement matrix; correct product report NOT AVAILABLE FOR VERIFICATION

Exact problem: WI p.10 references AQL general/special inspection and two-hour production checks, but applicable SOP, AQL/level, controlled sampling scheme and release evidence absent.

Required action: Preserve source company checks; require approved SOP/plan and batch-specific Ac/Re entries; no invented sampling.

Evidence: Objective technical file / controlled results; owner: QA / QC

# 6. WI change register — summary

| WI CHANGE | GAP / REQUIREMENT | SOURCE / BEFORE | PROPOSED AFTER / CONTROL |
| --- | --- | --- | --- |
| W-C001 | GAP-001 R001, R003, R005 | WI p.1 §5.1 and pp.2–10 Receiving COA and visual inspection listed; no ISO 18190 applicability table. | ADD ¦ Before release of this WI revision, QA shall obtain ISO 18190:2016 Clauses 4, 5, 6, 7, 8 and 9; record applicability, approved control location, evidence and owner for each applicable requirement. Any unresolved item shall remain OPEN and may not be declared compliant. |
| W-C002 | GAP-002 R002 | WI pp.1–11 No alternative type-test validation clause identified. | ADD ¦ Where a specified ISO type-test method is replaced by an alternative, Engineering and QA shall document equivalent safety, validate the alternative against the referenced ISO method, approve the validation before use, and retain records. |
| W-C003 | GAP-003 R004 | WI pp.1–11 No ISO 18562-1 pathway reference identified. | ADD ¦ QA shall control an evaluation and test strategy for breathing-gas-pathway biocompatibility in accordance with ISO 18562-1. Evidence and acceptance shall be governed by the approved assessment; ISO 18562-1 requirements remain pending until the standard and assessment are reviewed. |
| W-C004 | GAP-004 R006, R007, R008 | WI p.2 §§5.1.2–5.1.3 Designated Fr based on averaged axes; major/minor axis tolerance ±0.5 mm. | MODIFY ¦ Measure and record tracheal-segment major and minor outer axes separately against the approved size-specific drawing; each measurement shall meet its approved nominal ±0.5 mm. Calculate/verify the declared whole-number French size from the controlled design rule. Treat the standard’s “should” size statement as guidance until Engineering documents the interpretation. |
| W-C005 | GAP-005 R009, R010, R021 | WI p.5 §5.1.25 and p.4 valves Connector declared compliant with ISO 5356-1; valves visually inspected. | MODIFY / ADD ¦ Verify machine-end 15-mm conical male connector conformity against ISO 5356-1 and cuff-inlet compatibility with ISO 80369-7 L1 male Luer-slip using approved design evidence. Check and reconcile referenced ISO 5361:2016 §5.6.5 requirements. Rejection limits shall come only from the approved referenced standards. |
| W-C006 | GAP-006 R011, R012 | WI p.3 §§5.1.5–5.1.6 and pp.7–8 9.0 kPa or 1.5× diameter, first reached; whole system immersion ≥10 s with no bubbles. | KEEP / MODIFY ¦ For each applicable tracheal and bronchial cuff, record the selected inflation endpoint (9.0 kPa/90 cmH2O or 1.5× measured cuff diameter, whichever occurs first), actual reached pressure/diameter, inflation-source disconnect, entire-system immersion and observation for at least 10 s. Accept only with zero bubbles; log cuff identity separately. |
| W-C007 | GAP-007 R013, R014, R015 | WI p.3 §§5.1.7–5.1.8; p.8 §§7.7–7.8 Cuff diameter incorrectly refers ISO 5361 Annex C; herniation incorrectly Annex E. | MODIFY ¦ Measure both cuff diameters using ISO 5361:2016 Annex B: stabilize at (2.0±0.1) kPa at (23±2)°C for 5 min; record four diameters 45° apart and calculate their mean (mm, two decimals). Compare each to its marked nominal within ±15%, disclosing the widest diameter for a noncircular bronchial cuff. Test herniation per ISO 5361:2016 Annex D, individually inflating each cuff; do not allow extension beyond the nearest edge of the appropriate patient-end segment. |
| W-C008 | GAP-008 R016 | WI p.3 cuffs Cuffs described as smooth, with no burst or leak. | LINK / ADD ¦ Inspect the cuff attachment and transition for sharp edges; link the product-specific risk analysis and approved visual acceptance standards. Record outcome for each applicable cuff. |
| W-C009 | GAP-009 R017, R018, R019, R020, R022, R023, R024 | WI pp.3–4 §§5.1.9–5.1.13 Visual tube/pilot/valve instructions, without complete dimensions or deflation checks. | MODIFY / ADD ¦ Control cuff inflation-tube dimensions, separation angle, pilot/deflation indicator, intentional deflation, connector interface, S3 length conditional cases and any separation-distance mark per ISO 5361:2016 §§5.6.1–5.6.8 as invoked by ISO 16628:2022 §6.6. Use the component-specific approved drawing and record quantitative actuals. |
| W-C010 | GAP-010 R025, R026, R027 | WI p.10 §§8–9 Seal, sterilization carton stamp and labels checked. | LINK / ADD ¦ Packaging, sterile supply and manufacturer information shall be released only against approved specifications and the documented applicability of ISO 18190:2016 Clauses 7–9. Record packaging integrity, sterilization-release link where applicable, and controlled artwork/IFU revision; do not equate carton chamber stamp with sterile-release evidence. |
| W-C011 | GAP-011 R028, R029 | WI p.5 §5.1.26 / p.9 §7.25 Legible at 1m ±10mm, 215 ±5lx, rub 15s wet cloth. | MODIFY ¦ For marking qualification, first expose markings for the intended-use cumulative exposure duration to applicable IFU-listed drugs/chemicals and simulated contact substances identified in the risk file. Then verify legibility with vision acuity 1 at (1m±10mm) and (215±5lx); rub 15s without undue pressure with distilled-water-wet cloth. Marking shall remain legible. |
| W-C012 | GAP-012 R030, R031, R033, R034 | WI p.2 graduation; pp.5,9 marking Graduation and marking described generally. | ADD ¦ Verify on each applicable tube: whole-number Fr; bronchial OD in mm prefixed “Br” (verified using ISO 16628 Annex B); effective IDs in mm for bronchial AND tracheal lumina (verified per Annex C); length marks at 2-cm intervals from the bronchial tip at the specified location; and the correct “right” or “left” identification. Record each result individually. |
| W-C013 | GAP-013 R032 | WI p.2 §5.1.3 Generic tube IDs and ODs required as specification. | ADD ¦ Determine maximum effective inside diameter of the bronchial and tracheal lumina using ISO 16628:2022 Annex C: calibrated three-bead set (diameter accuracy ±0.01 mm, length accuracy ±0.1 mm), spaced by 1–2 mm knots and weighted (100±5 g). Record maximum passing diameter for each lumen, expressed downward to the nearest 0.2 mm. Verify marked values match approved claims. |
| W-C014 | GAP-014 R035 | WI p.5 §5.1.27 Entire bronchial machine end, bronchial cuff and pilot balloon blue. | KEEP / LINK ¦ The machine end of the bronchial segment, the bronchial cuff, and its associated pilot balloon shall each be entirely blue. Retain the distinct company control that tracheal-side elements are clear, subject to approved configuration specification. |
| W-C015 | GAP-015 R036, R037, R038 | WI p.10 §8 Generic package label content specified. | ADD ¦ In addition to approved company packaging data, verify that the individual pack or insert gives content description, controlled major and minor tracheal outside axes, and both tracheal and bronchial cuff marked diameters. Compare these with design specification and applicable device markings. |
| W-C016 | GAP-016 R039, R040 | WI p.5 §5.1.24 Bronchial tip curve/cutting smoothness inspected. | LINK ¦ Engineering shall document consideration of bronchial segment length, bronchial cuff length/position, bronchial tip bevel, and relevant right-side anatomy in the approved risk file. The “should” statements in ISO 16628 Clauses 6.4.3/6.7 are advisory; no fabricated numeric acceptance limit shall be added. |
| W-C017 | GAP-017 R041, R042 | WI p.5 §5.1.29; p.10 §7.28 Suction catheter visual and ISO8836 Table2 pull force included. | MODIFY / LINK ¦ If product configuration includes a separately classifiable respiratory suction catheter, Supplier Quality/Engineering shall document ISO 8836 applicability and verify attached components using applicable §8.1.1 Table 2 and Annex B controls. The test shall not be applied to every endobronchial assembly connector without technical justification. |
| W-C018 | GAP-018 | WI p.11 §10; provided End.pdf p.1 WI lists Endobronchial report reference; provided report photograph is Endotracheal Tube. | ADD / LINK ¦ QC shall use only the verified current Endobronchial Inspection & Test Report revision linked to this WI. Until obtained and approved, the accompanying proposed form shall remain DRAFT and no claim regarding completeness of the current controlled Endobronchial report is permitted. |
| W-C019 | GAP-019 | WI p.4 §§5.1.14–5.1.15 duplicate 5.1.14 and 5.1.15 appear twice for separate components. | MODIFY ¦ Assign unique section identifiers in the proposed WI and preserve the original page/section mapping in the change-control register. Original WI remains authoritative until approved supersession. |
| W-C020 | GAP-020 | WI pp.1,6–7,10 Visual general AQL, physical special AQL, two-hour packaging QC and SOP-8-2-6 cited. | KEEP / LINK ¦ Maintain existing company requirements for visual/general and physical/special sampling and two-hour packaging checks, subject to the applicable approved sampling SOP/plan. Each lot/interval shall record lot quantity, inspection level, AQL, sample size, Ac, Re, randomization, evidence and disposition; no nominal level or AQL shall be assumed. |

# 7. Inspection report change register — summary

All proposed rows are ADD/LINK because no matching Endobronchial current form was supplied. This is not a claim that the current form lacks these fields; a form-to-form redline is blocked pending retrieval.

| CHANGE / PROPOSED ROW | GAP | BEFORE: VERIFIED REPORT | AFTER / REQUIRED RESULT |
| --- | --- | --- | --- |
| I-C001 ¦ Product/form identification | GAP-018 | NOT AVAILABLE FOR VERIFICATION | Correct Endobronchial form ID, rev, code, left/right, WI rev, lot/batch and controlled spec |
| I-C002 ¦ Sampling & disposition | GAP-020 | NOT AVAILABLE FOR VERIFICATION | Lot quantity, plan ID/rev, inspection level, AQL, code, n, Ac, Re, drawn IDs, disposition |
| I-C003 ¦ Major tracheal OD axis | GAP-004 | NOT AVAILABLE FOR VERIFICATION | Nominal from drawing ±0.5 mm; actual in mm; equipment ID |
| I-C004 ¦ Minor tracheal OD axis | GAP-004 | NOT AVAILABLE FOR VERIFICATION | Nominal from drawing ±0.5 mm; actual in mm; equipment ID |
| I-C005 ¦ French designated size | GAP-004 | NOT AVAILABLE FOR VERIFICATION | Whole number; measured axes / controlled size rule / printed Fr |
| I-C006 ¦ 15 mm conical connectors | GAP-005 | NOT AVAILABLE FOR VERIFICATION | ISO5356-1 conformance evidence ID and approved result |
| I-C007 ¦ Inflation inlet | GAP-005 | NOT AVAILABLE FOR VERIFICATION | ISO80369-7 L1 male Luer-slip compatibility record |
| I-C008 ¦ Tracheal cuff leakage | GAP-006 | NOT AVAILABLE FOR VERIFICATION | 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 |
| I-C009 ¦ Bronchial cuff leakage | GAP-006 | NOT AVAILABLE FOR VERIFICATION | 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 |
| I-C010 ¦ Tracheal cuff diameter | GAP-007 | NOT AVAILABLE FOR VERIFICATION | ISO5361 Annex B; claimed mm; measured mean mm; ±15% |
| I-C011 ¦ Bronchial cuff diameter | GAP-007 | NOT AVAILABLE FOR VERIFICATION | ISO5361 Annex B; declared/widest mm; mean mm; ±15% |
| I-C012 ¦ Cuff herniation per cuff | GAP-007 | NOT AVAILABLE FOR VERIFICATION | ISO5361 Annex D; tracheal/bronchial separate observations |
| I-C013 ¦ Cuff sharp edges | GAP-008 | NOT AVAILABLE FOR VERIFICATION | Visual and risk file ID; none at attachments/transitions |
| I-C014 ¦ Inflation tube / geometry | GAP-009 | NOT AVAILABLE FOR VERIFICATION | ≤3.0mm OD, ≤10% intrusion where applicable, separation angle ≤45°, S3 conditional |
| I-C015 ¦ Pilot function & deflation | GAP-009 | NOT AVAILABLE FOR VERIFICATION | Indication of inflation/deflation, unobstructed intentional deflation |
| I-C016 ¦ Marking exposure durability | GAP-011 | NOT AVAILABLE FOR VERIFICATION | Exposures/duration, 1m±10mm, 215±5lx, acuity1, 15s wet-rub result |
| I-C017 ¦ Fr / Br markings | GAP-012 | NOT AVAILABLE FOR VERIFICATION | Whole Fr; Br marked OD; Annex B actual mm |
| I-C018 ¦ Effective IDs / both lumina | GAP-013 | NOT AVAILABLE FOR VERIFICATION | ISO16628 Annex C; tracheal and bronchial actual and claimed mm |
| I-C019 ¦ Tip-origin length marks | GAP-012 | NOT AVAILABLE FOR VERIFICATION | 2cm gradations measured from bronchial tip |
| I-C020 ¦ Side marking | GAP-012 | NOT AVAILABLE FOR VERIFICATION | Right or left matches configuration |
| I-C021 ¦ Bronchial color | GAP-014 | NOT AVAILABLE FOR VERIFICATION | Bronchial machine end + cuff + pilot each blue |
| I-C022 ¦ Pack/insert content | GAP-015 | NOT AVAILABLE FOR VERIFICATION | Description, major and minor axes, tracheal and bronchial cuff diameters |
| I-C023 ¦ Pack sterile link | GAP-010 | NOT AVAILABLE FOR VERIFICATION | Pack sealing and sterile-release references where applicable; ISO18190 review pending |
| I-C024 ¦ External general/material/design technical file | GAP-001 | NOT AVAILABLE FOR VERIFICATION | ISO18190 Clauses 4–6 applicability and evidence ID; no implicit PASS |
| I-C025 ¦ Breathing-gas pathway biocompatibility | GAP-003 | NOT AVAILABLE FOR VERIFICATION | ISO18562-1 evaluation, approved results and reviewer |
| I-C026 ¦ Manufacturer information / IFU approval | GAP-010 | NOT AVAILABLE FOR VERIFICATION | ISO18190 Clause9 assessment, approved IFU/label evidence ID |
| I-C027 ¦ Optional suction catheter | GAP-017 | NOT AVAILABLE FOR VERIFICATION | BOM applicability; ISO8836 Table2 force by actual OD if needed |
| I-C028 ¦ Existing component appearance | GAP-020 | NOT AVAILABLE FOR VERIFICATION | Smooth tube, blue/clear subassemblies, bonds, connectors, lumen clear, x-ray line, label, stylet, swivel, seal |

# 8. Test and other control change registers

| ID / R | CURRENT | PROBLEM / ACTION | NEW METHOD / ACCEPTANCE | EVIDENCE |
| --- | --- | --- | --- | --- |
| T-C001 ¦ R012 | WI pp.3,7–8 cuff immersion method present | No cuff-specific recorded endpoints / KEEP / MODIFY | Record each cuff separately, inflation endpoint, pressure or diameter, ≥10s fully submerged bubble-free result No bubbles ≥10s | Cuff leakage worksheet |
| T-C002 ¦ R013, R014 | WI cites Annex C for diameter | Wrong Annex and no conditioning/readings / MODIFY | ISO5361 Annex B: 2.0±0.1kPa, 23±2°C, 5min, four 45° readings, mean mm 2dp ±15% of declared tracheal and bronchial cuff diameter | Cuff diameter worksheet |
| T-C003 ¦ R015 | WI cites Annex E for herniation | Wrong Annex / MODIFY | ISO5361 Annex D; 24h (40±1°C), weight 100g ≥60s, monitor deflation ≥10s; validate product fixture applicability Neither cuff reaches corresponding distal edge | Cuff herniation protocol |
| T-C004 ¦ R031 | Generic OD visual/dimensional inspection | No ISO Annex B five-wrap method / ADD | ISO16628 Annex B five closely spaced spiral wraps, OD = length/(5π), round up to next 0.5mm Br marking based on prescribed rounded value | Bronchial OD worksheet |
| T-C005 ¦ R032 | Generic inside diameter dimensional mention | No effective functional lumen ID / ADD | ISO16628 Annex C three-bead string, 100±5g mass, D calibration and 0.2mm increments; two lumina Actual max passing ID, expressed downward to nearest 0.2mm | Bead method worksheet |
| T-C006 ¦ R028, R029 | Wet-cloth rub described | Exposure conditioning not identified / MODIFY | ISO16628 §9.2.1 IFU/risk-selected exposure then acuity 1, 1m±10mm, 215±5lx, 15s wet rub Legible at end | Marking validation record |
| T-C007 ¦ R009, R010, R021 | Connector visual fit only | Normative ISO connector methods not verified / LINK | Obtain ISO5356-1/ISO80369-7 criteria and approved compatibility fixtures; technical-file verification Pass according to externally approved normative criteria | Connector qualification report |
| T-C008 ¦ R017, R018, R019, R020, R022, R023, R024 | Qualitative inflating tube/valve observations | No 5361 §5.6 coverage / ADD | Measure/verify all applicable 5361 §5.6 dimensions/angles, pilot indication, and deflation with conditional S3 cases Each specified limit met; pending design inputs | Inflation functional worksheet |
| T-C009 ¦ R041 | WI displays ISO8836 Table2 pull forces | Applicability and Annex B method missing / LINK / ADD | If actual qualifying suction catheter: Annex B 23±2°C/50±20%RH 1h, pull 200±20mm/min, Table2 force by OD ≥5/15/20N by OD category when applicable | Optional component bonding record |

| CONTROL | CURRENT / PROBLEM | ACTION / TARGET DOCUMENT | REQUIRED EVIDENCE |
| --- | --- | --- | --- |
| O-C001 ¦ Approved technical specifications | Nominal tracheal axes, declared cuff sizes and ID claims not supplied | ADD/LINK — controlled Product Specification / engineering drawings | Reviewed nominal values and revision |
| O-C002 ¦ Referenced standards | ISO18190 / ISO18562 and connector normative content not supplied | LINK — master ISO applicability assessment | Copies of controlled standards and applicability signoff |
| O-C003 ¦ Sampling plan | WI refers SOP-8-2-6 and AQL without controlled plan attachment | LINK — controlled Sampling Plan; preserve two-hour in-process rule subject to risk review | Approved level/AQL/n/Ac/Re by actual lot |
| O-C004 ¦ Device/label artwork | Approved left/right labels and axis/cuff artwork not supplied | LINK — Packaging / Device Marking Master | Approved artwork and proofs |
| O-C005 ¦ Sterile product release | WI mentions post-sterilization chamber stamp only | LINK — sterilization validation and authorized disposition | Sterilization lot and release record |
| O-C006 ¦ Component suction catheter | Actual component BOM and source applicability unverified | LINK — product BOM and supplier verification | Verified suction component classification |

# 9. ISO Closure register — ISO Says

| ID / REF | ISO Says | Existing WI Says | Revised WI Says | INSPECTION / EVIDENCE | CLOSURE STATUS |
| --- | --- | --- | --- | --- | --- |
| R001 4.1 | Applicable general requirements of ISO 18190:2016 Clause 4 apply; exact subrequirements cannot be substantiated without that source. | WI pp.1–11: no traceable ISO 18190 Clause 4 matrix. | W-C001 Section(s) of proposed revised WI | Proposed I-C024 / Controlled ISO18190 assessment and approved risk-management file | OPEN — TECHNICAL INPUT |
| R002 4.2 | If an alternative type-test method is selected, equivalent safety must be established and the alternative validated against the specified method. | WI pp.1–11: no procedure for alternative method approval. | W-C002 Section(s) of proposed revised WI | Proposed — / Method-equivalence rationale; approved validation report | DOCUMENTARY COVERED |
| R003 5.1 | Applicable materials requirements of ISO 18190:2016 Clause 5 apply; detailed obligations require access to the referenced standard. | WI p.1 receiving / CoA check; material-standard compliance not defined. | W-C001 Section(s) of proposed revised WI | Proposed I-C024 / Approved material specification, traceability, ISO18190 gap review | OPEN — TECHNICAL INPUT |
| R004 5.2 | Tracheobronchial tubes must be evaluated/tested according to ISO 18562-1 for breathing-gas-pathway biocompatibility. | WI pp.1–11: no ISO 18562-1 technical-file reference identified. | W-C003 Section(s) of proposed revised WI | Proposed I-C025 / ISO 18562-1 assessment and approved test/evaluation records | OPEN — TECHNICAL INPUT |
| R005 6.1 | Applicable design requirements of ISO 18190:2016 Clause 6 apply; review the referenced source for actual limits/methods. | WI pp.2–10 design characteristics; no linked ISO18190 review. | W-C001 Section(s) of proposed revised WI | Proposed I-C024 / Approved design verification matrix | OPEN — TECHNICAL INPUT |
| R006 6.2 | Designated French/Charrière size is expressed as a whole number and derived from average tracheal-segment major/minor outside axis dimensions; its stated ±1 mm phrase uses “should” and needs documented interpretation before use as a mandatory lot-rejection rule. | WI p.2 §5.1.2 includes whole Fr and averages major/minor axis; no actual values recorded. | W-C004 Section(s) of proposed revised WI | Proposed I-C005 / Major/minor actual readings, calculation, drawing, labelled Fr size | DOCUMENTARY COVERED |
| R007 6.3 | Major-axis dimension of the tracheal segment outside cross-section is subject to ±0.5 mm tolerance relative to its controlled nominal dimension. | WI p.2 §5.1.3 mentions ±0.5 mm; nominal/drawing not supplied. | W-C004 Section(s) of proposed revised WI | Proposed I-C003 / Drawing nominal; actual major-axis reading and disposition | OPEN — TECHNICAL INPUT |
| R008 6.3 | Minor-axis dimension of the tracheal segment outside cross-section is subject to ±0.5 mm tolerance relative to its controlled nominal dimension. | WI p.2 §5.1.3 mentions ±0.5 mm; nominal/drawing not supplied. | W-C004 Section(s) of proposed revised WI | Proposed I-C004 / Drawing nominal; actual minor-axis reading and disposition | OPEN — TECHNICAL INPUT |
| R009 6.4.1 | Machine-end tracheobronchial connectors must be 15 mm conical male connectors conforming to ISO 5356-1; technical-file conformance is required. | WI p.5 §5.1.25 references ISO 5356-1 without a test/technical-file link. | W-C005 Section(s) of proposed revised WI | Proposed I-C006 / ISO 5356-1 dimensional verification + approved technical file | OPEN — TECHNICAL INPUT |
| R010 6.4.2 | Cuff inflation inlet connectors must be compatible with the L1 male Luer-slip small-bore connector specified in ISO 80369-7; verify in the technical file. | WI pp.3–4 mentions valves and pilot balloons, not verified ISO 80369-7 compatibility. | W-C005 Section(s) of proposed revised WI | Proposed I-C007 / Approved ISO 80369-7 interface compatibility test | OPEN — TECHNICAL INPUT |
| R011 6.5.1 | If cuffs are fitted, each cuff must be permanently integrated with the tube and capable of inflation without leakage. | WI pp.3–4 cuff and pilot balloon attachment discussed. | W-C006 Section(s) of proposed revised WI | Proposed I-C008, I-C009 / Cuff attachment/leakage record for each required cuff | DOCUMENTARY COVERED |
| R012 6.5.1 | Cuff leak challenge: inflate up to 9.0 kPa (90 cmH2O) OR to 1.5 times measured cuff diameter, whichever is reached first; seal and disconnect the inflation source, immerse entire system, inspect ≥10 s, and accept only if no bubbles. | WI p.3 §§5.1.5–5.1.6 includes challenge and 10-s bubble check; actual results fields unverified. | W-C006 Section(s) of proposed revised WI | Proposed I-C008, I-C009 / Separate tracheal/bronchial cuff challenge pressure or diameter, ≥10-s observation | DOCUMENTARY COVERED |
| R013 6.5.2 | Tracheal cuff resting diameter must fall within ±15% of the declared value when measured per ISO 5361:2016 Annex B. | WI p.3 §5.1.7 cites incorrect ISO 5361 Annex C. | W-C007 Section(s) of proposed revised WI | Proposed I-C010 / Tracheal cuff marked diameter, four readings, measured mean, calculation | DOCUMENTARY COVERED |
| R014 6.5.2 | Bronchial cuff resting diameter must fall within ±15% of declared value when measured per ISO 5361:2016 Annex B; a non-round bronchial cuff must disclose its widest diameter. | WI p.3 §5.1.7 cites Annex C, although wide diameter language is present. | W-C007 Section(s) of proposed revised WI | Proposed I-C011 / Bronchial cuff four readings and widest/declared diameter evidence | DOCUMENTARY COVERED |
| R015 6.5.3 | Under ISO 5361:2016 Annex D cuff-herniation test, only the cuff under test is inflated and neither cuff may extend beyond nearest patient-end edge of its respective segment. | WI p.3 §5.1.8 cites incorrect ISO 5361 Annex E. | W-C007 Section(s) of proposed revised WI | Proposed I-C012 / Herniation test record per cuff with conditioning, fixture, force and observation | DOCUMENTARY COVERED |
| R016 6.5.4 | Cuff attachment transitions and outer surfaces must be free from sharp edges; demonstrate through review of risk-management documentation. | WI p.3 cuff surface mentions general appearance; risk-file link absent. | W-C008 Section(s) of proposed revised WI | Proposed I-C013 / Approved sharp-edge assessment and inspection records | DOCUMENTARY COVERED |
| R017 6.6 → 5361 5.6.1 | Where fitted, inflation-tube OD ≤3.0 mm; point of separation on concave side if curved; inflation lumen intrusion ≤10% of internal diameter; dimensions per referenced table/figures. | WI pp.3–4 describes the inflation tube qualitatively only. | W-C009 Section(s) of proposed revised WI | Proposed I-C014 / Inflation tube OD, drawing geometry, lumen intrusion assessment | OPEN — TECHNICAL INPUT |
| R018 6.6 → 5361 5.6.2 | Angle between inflation tube and tracheal tube at separation must not exceed 45°. | WI pp.3–4 contains no angular limit. | W-C009 Section(s) of proposed revised WI | Proposed I-C014 / Measured angle/drawing and approval | DOCUMENTARY COVERED |
| R019 6.6 → 5361 5.6.3 | Inflation tube must have a pilot balloon and/or another inflation/deflation indicator. | WI p.4 §§5.1.11–5.1.13 includes pilot balloons. | W-C009 Section(s) of proposed revised WI | Proposed I-C015 / Pilot balloon/indicator function record | DOCUMENTARY COVERED |
| R020 6.6 → 5361 5.6.4 | Cuff intentional deflation must not be obstructed by inflation tube, valve or other closure device. | WI p.4: valve integrity but no deflation test. | W-C009 Section(s) of proposed revised WI | Proposed I-C015 / Intentional-deflation time/observation and setup | DOCUMENTARY COVERED |
| R021 6.6 → 5361 5.6.5 | Inflation inlet must accept the applicable male 6% taper fitting under referenced 5361 Clause 5.6.5; reconcile with ISO 16628 Clause 6.4.2 requirement for ISO 80369-7 L1 male compatibility. | WI pp.3–5 mentions valves, with no interface verification. | W-C005 Section(s) of proposed revised WI | Proposed I-C007 / Documented reconciled connector-interface assessment | OPEN — TECHNICAL INPUT |
| R022 6.6 → 5361 5.6.6 | If inflation valve/closure is absent, inflation tube free S3 length must be ≥40 mm. | WI pp.3–4: no S3 dimension. | W-C009 Section(s) of proposed revised WI | Proposed I-C014 / S3 length measured on relevant configuration | DOCUMENTARY COVERED |
| R023 6.6 → 5361 5.6.7 | If closure device exists, S3 ≥10 mm unless pilot balloon and inflation valve are integral. | WI pp.3–4: no S3 dimension or exception decision. | W-C009 Section(s) of proposed revised WI | Proposed I-C014 / S3 measurement and actual valve configuration | DOCUMENTARY COVERED |
| R024 6.6 → 5361 5.6.8 | If inflation-tube separation distance from patient end is marked, actual dimension must equal mark ±10 mm. | WI pp.3–5 does not establish if that distance is marked. | W-C009 Section(s) of proposed revised WI | Proposed I-C014 / Marking applicability and measured dimension | DOCUMENTARY COVERED |
| R025 7 | Sterile-supplied tracheobronchial tubes must satisfy ISO 18190:2016 Clause 7, details not independently verifiable without ISO 18190. | WI p.10 §9 refers to post-sterilization carton chamber stamps only. | W-C010 Section(s) of proposed revised WI | Proposed I-C023 / Sterilization process validation, release records, ISO18190 Clause7 review | OPEN — TECHNICAL INPUT |
| R026 8 | Packaging for tracheobronchial tubes must meet ISO 18190:2016 Clause 8 requirements, which must be retrieved and verified. | WI p.10 §8 checks sealed pouch and label but lacks referenced technical requirements. | W-C010 Section(s) of proposed revised WI | Proposed I-C023 / Approved packaging specification and qualification / integrity evidence | OPEN — TECHNICAL INPUT |
| R027 9.1 | Applicable manufacturer-information requirements of ISO 18190:2016 Clause 9 apply; source must be obtained for full gap analysis. | WI pp.5,10 names marking/pack labels without controlled IFU review. | W-C010 Section(s) of proposed revised WI | Proposed I-C026 / Approved IFU and label artwork, ISO18190 Clause9 assessment | OPEN — TECHNICAL INPUT |
| R028 9.2.1 | Markings must remain durable/legible after representative exposure to materials they contact in intended use for the expected exposure duration, including listed drugs/chemicals and applicable simulated fluids. | WI p.5 §5.1.26 includes reading/rubbing step but omits representative-use chemical conditioning. | W-C011 Section(s) of proposed revised WI | Proposed I-C016 / Exposure media/rationale/duration, post-exposure result | OPEN — TECHNICAL INPUT |
| R029 9.2.1 | After exposure, check marking legibility at 1 m ±10 mm with visual acuity 1 and 215 ±5 lx; hand-rub without undue pressure for 15 s using distilled-water-wet cloth; remain legible. | WI p.5 §5.1.26 partly includes 1 m, 215 lx, 15 s and water cloth, but not acuity/pre-exposure. | W-C011 Section(s) of proposed revised WI | Proposed I-C016 / Controlled lux/distance/time recording and legibility result | DOCUMENTARY COVERED |
| R030 9.2.2(a) | Device shaft must display whole-number Charrière/French size designation. | WI p.2 §5.1.2 explains Fr but no label-specific inspection field. | W-C012 Section(s) of proposed revised WI | Proposed I-C017 / Product photo or completed marking record | DOCUMENTARY COVERED |
| R031 9.2.2(b) | Device shaft must state bronchial segment outside diameter in mm with prefix Br; verify using ISO 16628 Annex B. | WI pp.5,9 marks generally; not the Br numeric measurement/mark linkage. | W-C012 Section(s) of proposed revised WI | Proposed I-C017 / 5-turn thread length / (5π), rounded upward 0.5 mm; photo | DOCUMENTARY COVERED |
| R032 9.2.2(c) | Device shaft must show effective inside diameters (mm) for BOTH tracheal and bronchial lumen; determine using ISO 16628 Annex C bead method. | WI p.2 dimensional table only refers generic ID/OD, no effective ID bead method. | W-C013 Section(s) of proposed revised WI | Proposed I-C018 / Both lumen bead measurements & declared markings | DOCUMENTARY COVERED |
| R033 9.2.2(d) | Length graduations appear at 2 cm intervals, measured from bronchial tip, on the machine end of the tracheal segment. | WI p.2 notes accurate graduation only; origin/interval unspecified. | W-C012 Section(s) of proposed revised WI | Proposed I-C019 / Graduation interval and reference origin measurements | DOCUMENTARY COVERED |
| R034 9.2.2(e) | Device shaft must be labelled right or left, matching the manufactured configuration. | WI p.5 §5.1.27 discusses colors; right/left identification missing. | W-C012 Section(s) of proposed revised WI | Proposed I-C020 / Right/left label vs approved BOM/variant | DOCUMENTARY COVERED |
| R035 9.3 | Entire machine end of bronchial segment, bronchial cuff and corresponding pilot balloon must be blue. | WI p.5 §5.1.27 explicitly describes bronchial blue. | W-C014 Section(s) of proposed revised WI | Proposed I-C021 / Bronchial color verification and variant ID | DOCUMENTARY COVERED |
| R036 9.4(a) | Individual package or insert must describe package contents. | WI p.10 §8 requires label product name/code/lot etc, but does not explicitly check contents description. | W-C015 Section(s) of proposed revised WI | Proposed I-C022 / Approved artwork and package check | DOCUMENTARY COVERED |
| R037 9.4(b) | Individual package or insert must include the major-axis and minor-axis dimensions of the tracheal segment. | WI p.10 §8 does not state major/minor axis on packaging. | W-C015 Section(s) of proposed revised WI | Proposed I-C022 / Packaging nominal axes and proof of printing | DOCUMENTARY COVERED |
| R038 9.4(c) | Individual package or insert must state tracheal and bronchial cuff diameters as declared for Clause 6.5.2. | WI p.10 §8 does not state both cuff diameters on packaging. | W-C015 Section(s) of proposed revised WI | Proposed I-C022 / Approved pack print and marked-cuff diameter concordance | DOCUMENTARY COVERED |
| R039 6.4.3 (should) | Bronchial patient end is recommended to have a bevel; this is guidance wording, not a standalone “shall” obligation. | WI p.5 §5.1.24 includes bronchial tip curve visual inspection. | W-C016 Section(s) of proposed revised WI | Proposed — / Approved tip-geometry rationale | DOCUMENTARY COVERED |
| R040 6.7 (should) | Bronchial segment length, cuff location/length, bevel and right upper-lobe considerations are recommended inputs to design margin-of-safety assessment. | WI p.5 mentions tip curve but no margin-of-safety risk reference. | W-C016 Section(s) of proposed revised WI | Proposed — / Approved bronchial margin-of-safety design rationale | DOCUMENTARY COVERED |
| R041 8836:2014 7.3.2, 8.1.1 | IF a separate suction catheter component qualifies as a suction catheter under ISO 8836, verify securely attached connector per Annex B and Table 2: OD 1.33–2.67 mm: ≥5 N; 3.00–4.67 mm: ≥15 N; >5 mm: ≥20 N. | WI p.5 §5.1.29 / p.10 includes Table 2 values, but method, OD category and component scope unclear. | W-C017 Section(s) of proposed revised WI | Proposed I-C027 / Approved component applicability and axial pull force record | OPEN — TECHNICAL INPUT |
| R042 8836:2014 Clause 7 | IF an actual respiratory suction catheter is part of the defined configuration, confirm further relevant ISO 8836 design/connection requirements against its complete component classification. | WI p.5 §5.1.28 lists suction catheter visual only. | W-C017 Section(s) of proposed revised WI | Proposed I-C027 / Approved component design and supplier certificate | OPEN — TECHNICAL INPUT |

# 10. Cross-document consistency check

| CHECK | SUBJECT | RESULT | CONCLUSION |
| --- | --- | --- | --- |
| C001 | Cuff diameter | WI p.3 Annex C vs ISO16628 Annex B; proposed WI/IR corrected | CORRECTED IN DRAFT — NOT APPROVED |
| C002 | Herniation | WI p.3 Annex E vs ISO16628 Annex D; proposed WI/IR corrected | CORRECTED IN DRAFT — NOT APPROVED |
| C003 | Product report identity | Uploaded Endotracheal report mismatches Endobronchial WI | OPEN — matching form required |
| C004 | 15mm / Luer interfaces | ISO5356-1 and ISO80369-7 references; externally verified criteria unavailable | OPEN — independent source review |
| C005 | Packaging / cuff nominal values | Label, drawing and actual nominal technical inputs absent | OPEN — APPROVED PRODUCT SPECIFICATION |
| C006 | Sampling | WI sampling SOP referenced; n/Ac/Re cannot be determined from uploads | OPEN — PLAN / LOT INFO |
| C007 | Colour coding | Bronchial blue in WI and proposed records | DOCUMENTARY COVERED — no operational evidence |
| C008 | Tube marking, effective IDs | Proposed Annex B/C method vs old generic visual WI | CORRECTED IN DRAFT — evidence pending |
| C009 | Sterile and packaging release | ISO18190 referenced without source | OPEN — external standard |
| C010 | Per-row evidence traceability | Proposed report rows link IDs/workbooks, no actual sample results supplied | OPEN — IMPLEMENTATION |

# 11. QC Manager decision register

| DECISION | REQUEST / RECOMMENDATION | REFERENCE | FUNCTION | PRIORITY |
| --- | --- | --- | --- | --- |
| D001 | Authorize retrieval and controlled review of ISO18190:2016 and ISO18562-1, plus connector reference standards | GAP-001, GAP-003, GAP-005, GAP-010 | QA / Engineering | P1 |
| D002 | Approve product-specific configuration matrix/drawings and size/cuff declared nominal values | GAP-004, GAP-007, GAP-012, GAP-015 | Engineering / QC | P1 |
| D003 | Obtain actual controlled Endobronchial inspection form and verify status; do not use Endotracheal form | GAP-018 | Document Control / QA | P1 |
| D004 | Approve risk-based inspection/qualification split and sampling plan, including manufacturing checks | GAP-006, GAP-009, GAP-013, GAP-020 | QC / QA | P1 |
| D005 | Confirm suction catheter BOM classification and whether ISO8836 controls apply to its separately supplied component | GAP-017 | Engineering / Supplier Quality | P2 |
| D006 | Route proposed revised WI/form/worksheets through formal change-control and authority matrix | GAP-007, GAP-011, GAP-019 | QA / Document Control | P1 |
| D007 | Review outsourced test/validation needs for Annex C, cuff herniation and marking durability | GAP-007, GAP-011, GAP-013 | Engineering / QC / External Lab | P1 |

# 12. Corrective action plan

| ACTION | GAP | PR | ACTION | OWNER | EVIDENCE | STATUS |
| --- | --- | --- | --- | --- | --- | --- |
| CAP-001 | GAP-001 | P1 | Acquire controlled ISO 18190:2016 and create clause-by-clause compliance assessment. | Engineering / QA | Approvals / objective evidence named in finding | OPEN |
| CAP-002 | GAP-002 | P2 | Define validation/approval gate before any substitution. | QA / Engineering | Approvals / objective evidence named in finding | OPEN |
| CAP-003 | GAP-003 | P1 | Obtain applicable ISO 18562-1 and approved evaluation/test report. | QA / Regulatory | Approvals / objective evidence named in finding | OPEN |
| CAP-004 | GAP-004 | P2 | Introduce dimensional worksheet and require approved size/design drawing. | Engineering / QC | Approvals / objective evidence named in finding | OPEN |
| CAP-005 | GAP-005 | P1 | Add technical-file verification and control compatibility evidence. | Engineering / QA | Approvals / objective evidence named in finding | OPEN |
| CAP-006 | GAP-006 | P1 | Maintain existing test; require cuff-specific conditions and measurements in controlled record. | QC | Approvals / objective evidence named in finding | OPEN |
| CAP-007 | GAP-007 | P1 | Correct citations, restore exact method prerequisites and dedicated worksheets. | QC / QA | Approvals / objective evidence named in finding | OPEN |
| CAP-008 | GAP-008 | P1 | Link risk analysis, documented design check and inspection. | Engineering / QA | Approvals / objective evidence named in finding | OPEN |
| CAP-009 | GAP-009 | P1 | Add configuration-gated dimensional and functional verification. | Engineering / QC | Approvals / objective evidence named in finding | OPEN |
| CAP-010 | GAP-010 | P1 | Acquire ISO 18190; map to sterilization, packaging and label/IFU validation. | QA / Document Control | Approvals / objective evidence named in finding | OPEN |
| CAP-011 | GAP-011 | P1 | Establish exposure selection from IFU and risk file; execute marking durability qualification. | QA / Engineering | Approvals / objective evidence named in finding | OPEN |
| CAP-012 | GAP-012 | P1 | Add per-attribute ISO 16628 §9.2.2 fields and numeric verifications. | QC / Production | Approvals / objective evidence named in finding | OPEN |
| CAP-013 | GAP-013 | P1 | Add Annex C method/worksheet for both lumina with verified beads and mass. | Engineering / QC | Approvals / objective evidence named in finding | OPEN |
| CAP-014 | GAP-014 | P2 | KEEP existing color rule; map to ISO §9.3 and specific report row. | QC | Approvals / objective evidence named in finding | OPEN |
| CAP-015 | GAP-015 | P1 | Add approved label artwork verification and matching fields. | Document Control / QC | Approvals / objective evidence named in finding | OPEN |
| CAP-016 | GAP-016 | P2 | Record applicability/design rationale in risk file, not as unsupported ISO “shall” limit. | Engineering | Approvals / objective evidence named in finding | OPEN |
| CAP-017 | GAP-017 | P2 | Confirm BOM/configuration and apply ISO 8836 selectively to the actual suction catheter. | Engineering / Supplier Quality | Approvals / objective evidence named in finding | OPEN |
| CAP-018 | GAP-018 | P1 | Retrieve current controlled Endobronchial report, verify revision and rebuild product-specific report. | Document Control / QC | Approvals / objective evidence named in finding | OPEN |
| CAP-019 | GAP-019 | P2 | Renumber proposed WI uniquely; keep original p/section provenance in register. | Document Control | Approvals / objective evidence named in finding | OPEN |
| CAP-020 | GAP-020 | P1 | Preserve source company checks; require approved SOP/plan and batch-specific Ac/Re entries; no invented sampling. | QA / QC | Approvals / objective evidence named in finding | OPEN |

# 13. Final management conclusion

Existing WI coverage is PARTIAL. It includes substantive leak-test, size, marking and visual controls, but two incorrect cuff annex references, incomplete invoked-standard controls, insufficient product-specific measurement records, and incomplete traceability create material gaps. The photographed inspection report cannot be used to assess the current Endobronchial form. A complete proposed WI and report structure has been drafted, retaining valid company controls and making numerical/mechanical checks measurable. Some proposed controls remain blocked by missing drawings, referenced external standards, approved sampling plan, configuration and objective test records.

The proposed document revisions establish a traceable compliance control framework. Actual conformity remains subject to controlled approval, implementation, verification and review of objective evidence.

# 14. Source traceability register

| SOURCE | SCOPE | RELATIONSHIP |
| --- | --- | --- |
| S01 — ISO 16628:2022(E) | PDF pp.7–18 / printed pp.1–12; primary requirements | Source of §4–§9 and normative Annex B/C |
| S02 — ISO 5361:2016(E) | PDF p.7 / printed p.1 (scope), PDF p.20 / printed p.14 (§5.6); PDF p.32/printed 26 Annex B; PDF p.36/printed 30 Annex D | ONLY specifically invoked requirements; not wholesale primary product scope |
| S03 — BS EN ISO 8836:2014 | Scope/ISO §7.3, §8.1 Table2, Annex B | Conditional qualified suction catheter components only |
| S04 — End.pdf WI pp.2–12 | Printed WI pp.1–11 with document no./Rev/date | Current scanned Endobronchial WI baseline |
| S05 — End.pdf p.1 | Endotracheal product title and generic report lines | Non-matching other-product form; not product-specific report baseline |
| Missing records | ISO18190, ISO18562, ISO5356-1, ISO80369-7, product drawings, specifications, controlled report, lot results | NOT AVAILABLE FOR VERIFICATION |