QC SYSTEMS / MEDICAL DEVICES

# Endobronchial Tube — Inspection & Test Report

EDITABLE FORM (PROPOSED) | Route-based monitoring plus evidence references

| Product | Endobronchial Tube (double-lumen Tracheobronchial Tube) |
| --- | --- |
| Primary standard | ISO 16628:2022 (second edition) |
| Existing work instruction | WI-8-2-2-06-1 / Rev 00 / 05 Dec 2023 |
| Matching inspection report | NOT VERIFIED — supplied form is Endotracheal Tube |
| Document identifier | FORM ID / REV: OPEN — QA TO ASSIGN |
| Status | CONTROLLED DRAFT FOR TECHNICAL REVIEW — NOT APPROVED |
| Issue date | 09 October 2026 |

Important limitation: source files are photographed scans and standards; no approved lot test data, drawings, ISO 18190 copy, ISO 18562-1 report or matching Endobronchial report was supplied. No operational requirement is claimed closed.

NOT AN APPROVED RECORD. This document is a suggested blank controlled form. A separate form titled Endotracheal Tube cannot be assumed its predecessor. All measurement values, sizes, sample frequencies and form codes require approved source input.

| Receiving / batch identifiers | Entry | Product definition | Entry |
| --- | --- | --- | --- |
| Supplier / Production source | ________________ | Lot / batch | ________________ |
| PO / manufacturing order | ________________ | Lot size | ________________ |
| Product SKU / code | ________________ | Size (whole-number Fr) | ________________ |
| Config / RIGHT or LEFT | ________________ | Sterile / bulk | ________________ |
| Inspection date / shift | ________________ | Inspection type | Incoming / In-process / Final |
| Drawing / spec code and rev | ________________ | WI code and rev | ________________ |
| Inspection plan / SOP ID | ________________ | Inspection level / AQL | ________________ |
| Sampling n / Ac / Re | ________________ | Individual sample IDs | ________________ |
| Equipment / calibration ref | ________________ | COA / supplier trace | ________________ |

# Part A — Routine/lot/in-process checks (complete per approved control plan)

| No. | Inspection characteristic | Acceptance criteria (controlled source) | Requirement / WI | Method / route | Actual result / unit / specimen | Pass / Fail / Pending | Evidence ID |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 01 | Pouch/pack integrity, receipt/visual | Intact, correctly sealed/undamaged; correct PO, COA and label as approved | WI §6,§13 | Visual / recorded lot | ________________ | ________ | ________ |
| 02 | Main tube surface / main limbs | Transparent/smooth with no specified defects, cracks, holes, black spots or burns | WI §8 / Table A | Visual / lot | ________________ | ________ | ________ |
| 03 | Tracheal major axis (mm) | Approved nominal major axis ±0.5 mm (actual target OPEN) | 16628 §6.3 / GAP-004 | Caliper / lot | ________________ | ________ | ________ |
| 04 | Tracheal minor axis (mm) | Approved nominal minor axis ±0.5 mm (actual target OPEN) | 16628 §6.3 / GAP-004 | Caliper / lot | ________________ | ________ | ________ |
| 05 | Whole-number Fr designation | Correct approved size derived from major/minor mean (approved map OPEN) | 16628 §6.2 / GAP-004 | Label and worksheet | ________________ | ________ | ________ |
| 06 | Tracheal cuff immersion leak | 9.0 kPa OR 1.5× cuff resting diameter first; ≥10 s; no bubbles | 16628 §6.5.1 / GAP-009 | Manometer/immersion | ________________ | ________ | ________ |
| 07 | Bronchial cuff immersion leak | 9.0 kPa OR 1.5× cuff resting diameter first; ≥10 s; no bubbles | 16628 §6.5.1 / GAP-009 | Manometer/immersion | ________________ | ________ | ________ |
| 08 | Both inflation circuits and pilot balloons | No leaks/blockage/holes; inflation & intentional deflation functional | 16628 §6.6 / GAP-013 | Functional | ________________ | ________ | ________ |
| 09 | Bronchial end / cuff / pilot balloon colour | Each entirely blue; company tracheal identifiers per artwork | 16628 §9.3 / GAP-016 | Visual | ________________ | ________ | ________ |
| 10 | Right/left and whole-number Fr markings | Correct side and readable whole-number Fr for SKU | 16628 §9.2.2 / GAP-015 | Artwork visual | ________________ | ________ | ________ |
| 11 | Br-prefixed bronchial OD marking (mm) | Actual “Br [mm]” agrees with qualified Annex B determination and artwork | 16628 §9.2.2b / GAP-005 | Artwork / linked worksheet | ________________ | ________ | ________ |
| 12 | Two-lumen effective ID markings (mm) | Both values shown, agree with Annex C qualified results | 16628 §9.2.2c / GAP-006 | Artwork / linked worksheet | ________________ | ________ | ________ |
| 13 | Length marks | Every 2 cm from bronchial tip, tracheal limb machine-end position | 16628 §9.2.2d / GAP-015 | Ruler / visual | ________________ | ________ | ________ |
| 14 | Connector / swivel / endoscopic fit | Approved BOM part, secure assembly, damage-free and working | WI §8 / GAP-007 | Visual / functional | ________________ | ________ | ________ |
| 15 | Cuff transitions and both tips | No sharp edges or dangerous protrusions / no blockage | 16628 §6.5.4 / WI §8 | Visual | ________________ | ________ | ________ |
| 16 | Suction accessory / stylet, if supplied | Correct item identity, undamaged, linked supplier qualification | WI §8 / GAP-019/020 | Visual / supplier | ________________ | ________ | ________ |
| 17 | Individual pack/IFU content | Description, major/minor axes, both cuff diameters, side/size, approved artwork | 16628 §9.4 / GAP-017 | Label master | ________________ | ________ | ________ |
| 18 | Packaging and sterile/lot release | Pouch intact, labels match, sterilization & approved release links if sterile | 16628 §§7-8 / GAP-022 | Production/QA records | ________________ | ________ | ________ |

# Part B — Separate qualification/supplier/periodic evidence links

| Test / document | Standard | Gap | Verification route | Approved report ID/rev | Validity / scope / lot link | QA verified (Y/N) | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Annex B bronchial OD (five turns; raw length / 5π; round UP 0.5 mm) | 16628 Annex B | GAP-005 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Annex C each effective ID, beads 0.2 mm; 100 g ±5 g; round DOWN 0.2 mm | 16628 Annex C | GAP-006 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Cuff diameters separately; ±15% of marked values | 16628 §6.5.2 / 5361 Annex B | GAP-010 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Both cuff herniation/no protrusion; only tested cuff inflated | 16628 §6.5.3 / 5361 Annex D | GAP-011 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| 15 mm connector conformance and L1 male Luer-slip inlet compatibility | 16628 §§6.4.1-6.4.2 | GAP-007/008 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Inflation tube dimensional & functional design verification | 16628 §6.6 / 5361 §5.6 | GAP-013 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Marking chemical exposure, 15 s rub, 1 m ±10 mm, 215 ±5 lx | 16628 §9.2.1 | GAP-014 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Gas-pathway biological safety evidence | 16628 §5.2 / 18562-1 | GAP-003 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| ISO 18190 Clauses 4-9 requirement/control crosswalk | 16628 §§4-9 | GAP-002 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Conditional ISO 8836 suction catheter / detachment force | 8836 §8.1 Annex B, conditional | GAP-019/020 | Qualification / supplier | ________________ | ________________ | ______ | Pending |
| Risk-file design assessments and sterile barrier validation | 16628 §§6.7,7-8 | GAP-018/022 | Qualification / supplier | ________________ | ________________ | ______ | Pending |

# Part C — Detailed cuff leak test observations

| Specimen / lot | Cuff under test | Resting Ø (mm) | 1.5 × Ø (mm) | Pressure (kPa) | Endpoint reached first | Immersion duration (s) | Bubbles YES / NO | Result and equipment ID |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ____________ | Tracheal | ______ | ______ | ______ | ______ | ______ | ______ | ____________ |
| ____________ | Bronchial | ______ | ______ | ______ | ______ | ______ | ______ | ____________ |

Cuff test instruction: verify actual pressure is 9.0 kPa maximum OR inflation reaches 1.5 times the cuff resting diameter first; detach inflator, seal and immerse complete system for ≥10 s. Acceptance is NO bubbles. Attach additional sample sheets when n > 1.

# Part D — Sample plan, test results and approvals

Sample plan: SOP ID ______  n ______  Ac _____  Re _____  Lot ______  Time / interval ______  Defects ____  Result ACCEPT / REJECT / HOLD (circle). Each actual numerical result must be retained against specimen ID, calibrated apparatus and WI clause; use linked worksheets when the box is too small.

Discrepancies / NCR ID: __________________  Quarantine / disposition: __________________  Evidence missing and HOLD decision: __________________

Prepared/tested by QC: __________________  Date/time: __________  QC reviewer: __________________  QA release authorization (where required): __________________  Date: __________