# PROPOSED INSPECTION & TEST REPORT

Endobronchial (Double-Lumen / Tracheobronchial) Tube — fillable by editing Word cells

STATUS: PROPOSED CONTROLLED DRAFT  |  NO SIGNATURES OR RELEASE AUTHORIZATION

Source basis: ISO 16628:2022; ISO 5361:2016 invoked clauses/annexes; conditional ISO 8836:2014; scanned WI Rev 00 and non-matching inspection report photo.

CONTROL NOTICE: FORM NUMBER / REVISION NOT ALLOCATED; NOT FOR ROUTINE RELEASE UNTIL QMS APPROVED. The uploaded photographed form is for Endotracheal Tube, not Endobronchial Tube. The existing Endobronchial Inspection & Test Report is NOT AVAILABLE FOR VERIFICATION.

| PRODUCT / TRACEABILITY | TO COMPLETE | PRODUCT / TRACEABILITY | TO COMPLETE |
| --- | --- | --- | --- |
| Product / code | ________________ | Lot / batch | ________________ |
| Right / left; French size | ________________ | Material supplier/lot | ________________ |
| PO / receipt date | ________________ | Received / produced quantity | ________________ |
| Inspection date / shift | ________________ | Inspection stage | Incoming / In-process / Final / Qualification |
| WI ID / revision | ________________ | Current report ID / revision | ________________ |
| ISO reference | ISO 16628:2022 | Approved specification/revision | ________________ |
| Inspector ID | ________________ | Reviewer ID | ________________ |
| Sterilization run (if applicable) | ________________ | Approved drawing / BOM | ________________ |

# A. Sampling and equipment — mandatory actuals

| SAMPLING FIELD | ACTUAL / REFERENCE | SAMPLING FIELD | ACTUAL / REFERENCE |
| --- | --- | --- | --- |
| Lot size | ____________ | Approved sampling plan ID/Rev | ____________ |
| Inspection level | ____________ | AQL | ____________ |
| Sample code letter | ____________ | Sample size (n) | ____________ |
| Acceptance number (Ac) | ____________ | Rejection number (Re) | ____________ |
| Drawn sample IDs | ____________ | Sampling method | ____________ |

| EQUIPMENT NAME / TEST | EQUIPMENT ID | CALIBRATION / QUALIFICATION STATUS | ENVIRONMENT / SETUP |
| --- | --- | --- | --- |
| Pressure source / gauge | ____________ | ____________ | ____________ |
| Dimensional tools | ____________ | ____________ | ____________ |
| Temperature / immersion vessel | ____________ | ____________ | ____________ |
| Light meter / distance rule | ____________ | ____________ | ____________ |
| Bead set / calibrated mass | ____________ | ____________ | ____________ |
| Force tester / fixture | ____________ | ____________ | ____________ |

# B. Product-specific inspection & test results

Fill Required Limit from verified ISO / approved controlled specification. Actual numeric value is mandatory wherever acceptance is quantitative. For qualification-only controls, enter “Controlled qualification record ID” and “PENDING” if not reviewed; never invent lot data.

| ROW / REQ | CHARACTERISTIC | REQUIRED / METHOD | ACTUAL VALUE / UNIT | SAMPLE & EVIDENCE ID | PASS / FAIL / PENDING |
| --- | --- | --- | --- | --- | --- |
| 01 ¦ | Product/form identification | Correct Endobronchial form ID, rev, code, left/right, WI rev, lot/batch and controlled spec | ___________________ | ___________________ | __________ |
| 02 ¦ | Sampling & disposition | Lot quantity, plan ID/rev, inspection level, AQL, code, n, Ac, Re, drawn IDs, disposition | ___________________ | ___________________ | __________ |
| 03 ¦ R007 | Major tracheal OD axis | Nominal from drawing ±0.5 mm; actual in mm; equipment ID | ___________________ | ___________________ | __________ |
| 04 ¦ R008 | Minor tracheal OD axis | Nominal from drawing ±0.5 mm; actual in mm; equipment ID | ___________________ | ___________________ | __________ |
| 05 ¦ R006 | French designated size | Whole number; measured axes / controlled size rule / printed Fr | ___________________ | ___________________ | __________ |
| 06 ¦ R009 | 15 mm conical connectors | ISO5356-1 conformance evidence ID and approved result | ___________________ | ___________________ | __________ |
| 07 ¦ R010,R021 | Inflation inlet | ISO80369-7 L1 male Luer-slip compatibility record | ___________________ | ___________________ | __________ |
| 08 ¦ R011,R012 | Tracheal cuff leakage | 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 | ___________________ | ___________________ | __________ |
| 09 ¦ R011,R012 | Bronchial cuff leakage | 9.0 kPa or 1.5×; measured pressure/diameter; submerge ≥10s; bubbles=0 | ___________________ | ___________________ | __________ |
| 10 ¦ R013 | Tracheal cuff diameter | ISO5361 Annex B; claimed mm; measured mean mm; ±15% | ___________________ | ___________________ | __________ |
| 11 ¦ R014 | Bronchial cuff diameter | ISO5361 Annex B; declared/widest mm; mean mm; ±15% | ___________________ | ___________________ | __________ |
| 12 ¦ R015 | Cuff herniation per cuff | ISO5361 Annex D; tracheal/bronchial separate observations | ___________________ | ___________________ | __________ |
| 13 ¦ R016 | Cuff sharp edges | Visual and risk file ID; none at attachments/transitions | ___________________ | ___________________ | __________ |
| 14 ¦ R017,R018,R022,R023,R024 | Inflation tube / geometry | ≤3.0mm OD, ≤10% intrusion where applicable, separation angle ≤45°, S3 conditional | ___________________ | ___________________ | __________ |
| 15 ¦ R019,R020 | Pilot function & deflation | Indication of inflation/deflation, unobstructed intentional deflation | ___________________ | ___________________ | __________ |
| 16 ¦ R028,R029 | Marking exposure durability | Exposures/duration, 1m±10mm, 215±5lx, acuity1, 15s wet-rub result | ___________________ | ___________________ | __________ |
| 17 ¦ R030,R031 | Fr / Br markings | Whole Fr; Br marked OD; Annex B actual mm | ___________________ | ___________________ | __________ |
| 18 ¦ R032 | Effective IDs / both lumina | ISO16628 Annex C; tracheal and bronchial actual and claimed mm | ___________________ | ___________________ | __________ |
| 19 ¦ R033 | Tip-origin length marks | 2cm gradations measured from bronchial tip | ___________________ | ___________________ | __________ |
| 20 ¦ R034 | Side marking | Right or left matches configuration | ___________________ | ___________________ | __________ |
| 21 ¦ R035 | Bronchial color | Bronchial machine end + cuff + pilot each blue | ___________________ | ___________________ | __________ |
| 22 ¦ R036,R037,R038 | Pack/insert content | Description, major and minor axes, tracheal and bronchial cuff diameters | ___________________ | ___________________ | __________ |
| 23 ¦ R025,R026 | Pack sterile link | Pack sealing and sterile-release references where applicable; ISO18190 review pending | ___________________ | ___________________ | __________ |
| 24 ¦ R001,R003,R005 | External general/material/design technical file | ISO18190 Clauses 4–6 applicability and evidence ID; no implicit PASS | ___________________ | ___________________ | __________ |
| 25 ¦ R004 | Breathing-gas pathway biocompatibility | ISO18562-1 evaluation, approved results and reviewer | ___________________ | ___________________ | __________ |
| 26 ¦ R027 | Manufacturer information / IFU approval | ISO18190 Clause9 assessment, approved IFU/label evidence ID | ___________________ | ___________________ | __________ |
| 27 ¦ R041,R042 | Optional suction catheter | BOM applicability; ISO8836 Table2 force by actual OD if needed | ___________________ | ___________________ | __________ |
| 28 ¦ | Existing component appearance | Smooth tube, blue/clear subassemblies, bonds, connectors, lumen clear, x-ray line, label, stylet, swivel, seal | ___________________ | ___________________ | __________ |

# C. Additional test worksheets — controlled data collection

## WS-01 | Tracheal and bronchial cuff leakage

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Lot & sample ID | _________________________________________ |
| Cuff: tracheal / bronchial | _________________________________________ |
| Marked cuff diameter (mm) | _________________________________________ |
| Measured cuff diameter (mm) | _________________________________________ |
| 1.5 × limit (mm) | _________________________________________ |
| Pressure challenge (kPa / cmH2O) | _________________________________________ |
| Actual endpoint first reached | _________________________________________ |
| Inflation source disconnected? | _________________________________________ |
| Whole assembly submerged? | _________________________________________ |
| Immersion time (s, ≥10) | _________________________________________ |
| Observed bubble count | _________________________________________ |
| Result / inspector / reviewer | _________________________________________ |

## WS-02 | Cuff diameter - ISO 5361 Annex B

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Cuff identity | _________________________________________ |
| Marked nominal diameter (mm) | _________________________________________ |
| Temperature (°C; 23±2) | _________________________________________ |
| Inflation pressure (kPa; 2.0±0.1) | _________________________________________ |
| Stabilization time (min; 5) | _________________________________________ |
| Read 1 (0°) mm | _________________________________________ |
| Read 2 (45°) mm | _________________________________________ |
| Read 3 (90°) mm | _________________________________________ |
| Read 4 (135°) mm | _________________________________________ |
| Arithmetic mean (mm to 2dp) | _________________________________________ |
| Deviation % from nominal | _________________________________________ |
| Non-round bronchial: widest measurement | _________________________________________ |
| ±15% Pass/Fail / reference | _________________________________________ |

## WS-03 | Cuff herniation - ISO 5361 Annex D

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Cuff tested (inflate one only) | _________________________________________ |
| Fixture configuration verification ID | _________________________________________ |
| Conditioning temperature (°C; 40±1) | _________________________________________ |
| Conditioning time (h; 24) | _________________________________________ |
| Inflation pressure (kPa; specified method) | _________________________________________ |
| Suspended mass (g; 100) | _________________________________________ |
| Mass suspension time (s; ≥60) | _________________________________________ |
| Deflation observation (s; ≥10) | _________________________________________ |
| Patient-edge extension observed? | _________________________________________ |
| Result / evidence | _________________________________________ |

## WS-04 | Bronchial OD - ISO 16628 Annex B

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Bronchial cuff deflated / midpoint confirmed | _________________________________________ |
| Thread length across 5 turns (mm) | _________________________________________ |
| OD = length/(5×π) mm | _________________________________________ |
| Rounded UP next 0.5 mm | _________________________________________ |
| Device “Br” claim (mm) | _________________________________________ |
| Difference / confirmation | _________________________________________ |
| Device sample and instrument ID | _________________________________________ |

## WS-05 | Effective inside diameters - ISO 16628 Annex C

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Lumen: tracheal or bronchial | _________________________________________ |
| 3-bead set ID / D accuracy ±0.01 mm | _________________________________________ |
| Bead length accuracy ±0.1 mm | _________________________________________ |
| Knots 1–2mm / 100±5g mass ID | _________________________________________ |
| Largest passing bead (mm) | _________________________________________ |
| Rounded DOWN next 0.2 mm | _________________________________________ |
| Printed effective ID (mm) | _________________________________________ |
| Water-soluble lubricant if used | _________________________________________ |
| Result / signature | _________________________________________ |

## WS-06 | Marking durability - ISO 16628 Clause 9.2.1

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| IFU-listed contact media and rationale | _________________________________________ |
| Exposure period and equivalent in-use cumulative duration | _________________________________________ |
| Artificial saliva/mucus/skin oil applicability | _________________________________________ |
| Visual acuity 1 | _________________________________________ |
| Distance 1m±10mm | _________________________________________ |
| Illuminance 215±5lx | _________________________________________ |
| Rub time 15s with wet cloth | _________________________________________ |
| Final legibility / reviewer | _________________________________________ |

## WS-07 | Inflation system - ISO 5361 Clause 5.6

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Inflation OD (mm ≤3) | _________________________________________ |
| Lumen intrusion ≤10% verified? | _________________________________________ |
| Angle at split (° ≤45) | _________________________________________ |
| Pilot indicator function | _________________________________________ |
| Intentional deflation functional | _________________________________________ |
| Inlet compatibility standard / evidence | _________________________________________ |
| S3 length and exception applicability | _________________________________________ |
| Marked separation distance ±10 mm if used | _________________________________________ |

## WS-08 | Optional ISO 8836 Annex B pull strength

| PARAMETER | ACTUAL / RECORD |
| --- | --- |
| Component confirmed suction catheter in BOM? | _________________________________________ |
| Shaft OD category (mm) | _________________________________________ |
| Table2 minimum (N) | _________________________________________ |
| Conditioning 23±2°C / 50±20%RH for 1h | _________________________________________ |
| Separation speed 200±20mm/min | _________________________________________ |
| Actual detachment force (N) | _________________________________________ |
| Pass/Fail / test report ID | _________________________________________ |

# D. Review and disposition

| RESULT GATE | VERIFICATION / PERSON / DATE |
| --- | --- |
| All required test rows completed or linked to reviewed controlled qualification evidence | _______________________________ |
| NCR / deviation / quarantine references | _______________________________ |
| Documentary open issues confirmed resolved? | Yes / No — details: __________________ |
| QC inspector / date | _______________________________ |
| QC Supervisor reviewed / date | _______________________________ |
| QA or QMS authorized release / date | _______________________________ |
| Final disposition (select only after review) | Accepted / Rejected / Quarantine / Pending Further Verification |

No signatures, measured results, instruments, pass states, or source document approvals have been prefilled. Template acceptance values are ISO-sourced only where verifiable; all company-specific nominal dimensions, sampling and referenced-standard limits require actual approved technical inputs.