ENDOTRACHEAL TUBE

ISO 16628:2022 / ISO 5361:2016

WI & Inspection Report Gap Analysis

| Department | Quality Control (QC) |
| --- | --- |
| Product | Endotracheal Tube |
| Report basis | Uploaded ISO standards + current Inspection Report + retrievable prior project WI evidence |
| Assessment type | Evidence-based compliance and document-control gap analysis |
| Report date | 06 October 2026 |

CONTROLLED REVIEW NOTE: The current report is evaluated as a blank controlled template. Empty execution fields are not treated as a failed lot record; only missing or inadequate form content is treated as a documentation gap.

# 1. Executive Summary

The review identified a fundamental classification and document-linkage issue before the technical gaps are considered. The current form is titled “Inspection & Test Report for Endotracheal Tube,” while the retrievable project WI is WI-8-2-2-06-1 Rev.00, “Inspection Procedure for Endobronchial Tube.” These are not interchangeable standard scopes. ISO 5361:2016 applies to tracheal/endotracheal tubes and explicitly excludes endobronchial/tracheobronchial tubes; ISO 16628:2022 applies to double-lumen tracheobronchial tubes. The project had previously mapped the two documents together for gap-review purposes, but that operational mapping must not replace controlled product classification.

The most important technical defects are concentrated in cuff verification and the WI’s normative references. The WI criteria previously extracted for cuff leakage, cuff diameter and cuff herniation are not adequately recordable in the current inspection template. In addition, the WI extraction references ISO 5361 Annex C for cuff diameter and Annex E for cuff herniation; ISO 5361:2016 places cuff diameter in Annex B and cuff herniation in Annex D. That is a P1 technical/document-control error because it directs the inspector to the wrong test method.

ISO 8836:2014 is not a governing standard for the Endotracheal Tube itself; its scope is suction catheters for respiratory-tract suction. It is therefore excluded from the product compliance matrix except as a scope determination.

| Finding | Classification | Priority | Outcome |
| --- | --- | --- | --- |
| Product / standard / WI scope mismatch | Document Control Gap | P1 / High | Must be closed before controlled release of the WI/report set. |
| WI uses wrong ISO 5361 annex references for cuff diameter / herniation | Document Control Gap | P1 / High | Correct Annex B and Annex D references. |
| Cuff leakage method cannot be fully recorded | Evidence Gap | P1 / High | Add pressure/condition/time/cuff identity/actual result. |
| Cuff diameter ±15% not recordable | Evidence Gap | P1 / High | Add marked value, allowed range and actual diameter for each cuff. |
| Cuff herniation not present as a report item | Evidence Gap | P1 / High | Add independent herniation result field. |
| Standards delegated to technical file are not linked | Evidence Gap | P1 / High | Link risk, biocompatibility, sterility, packaging and type-test evidence rather than forcing lot testing. |

# 2. Documents Reviewed

| Document | Title | Extent | Use in this review |
| --- | --- | --- | --- |
| ISO 16628:2022 | Anaesthetic and respiratory equipment — Tracheobronchial tubes | 20 PDF pages | Primary only if device is tracheobronchial/double-lumen. |
| ISO 5361:2016 | Anaesthetic and respiratory equipment — Tracheal tubes and connectors | 48 PDF pages | Direct for endotracheal/tracheal tube; excluded for endobronchial except where ISO 16628 invokes it. |
| BS EN ISO 8836:2014 | Suction catheters for use in the respiratory tract | 40 PDF pages | Not applicable to the Endotracheal Tube itself. |
| WI-8-2-2-06-1 Rev.00 | Inspection Procedure for Endobronchial Tube | 11 internal pages (per prior project extraction) | Original WI PDF was not retrievable in the current file layer; retained project analysis was used for known clauses. |
| Current Inspection Report | Inspection & Test Report for Endotracheal Tube | 1 page | Uploaded image; Rev. 4 visible; Doc.# and date are partially cropped. |

# 3. Document Identification

| Product | Endotracheal Tube (current inspection report title). |
| --- | --- |
| Mapped WI | WI-8-2-2-06-1 — Inspection Procedure for Endobronchial Tube. |
| WI Revision | Rev. 00. |
| WI Date | Not verifiable from the currently retrievable project evidence. |
| Uploaded standards | ISO 16628:2022; ISO 5361:2016; BS EN ISO 8836:2014. |
| Inspection Report | Inspection & Test Report for Endotracheal Tube. |
| Report Revision | Rev. 4 (visible on uploaded image). |
| Report Document No. | Partially visible as “F-823-T…”; full number cannot be verified from the image. |
| Report Date | Partially visible as “23.04.202…”; full year cannot be verified from the image. |
| Report Page | 1/1. |
| Report WI linkage field | “According to WI:” is blank on the template. |
| Other linked forms | Prior project extraction indicated an Endobronchial inspection/test form reference, but the source WI PDF is not currently available to re-verify the number. No number is asserted as verified in this report. |

# 4. Scope

This review assesses only the Endotracheal Tube document set and the uploaded standards. It distinguishes direct product-standard requirements from requirements that are expected to be controlled in the technical file, risk-management file, supplier/CoA system, validation, qualification or periodic/type testing. No requirement is assumed merely because a standard title appears in a reference list.

# 5. Methodology

Determine standard applicability from each standard’s scope before comparing technical clauses.

Extract requirement, test method and acceptance criterion from the uploaded standard.

Compare the requirement with retained WI evidence and the current inspection-report row structure.

Classify each item as Covered, Partial Gap, Gap, Evidence Gap, Document Control Gap or Not Applicable.

Use NOT EXPECTED IN THIS REPORT when the standard expects technical-file/type-test evidence rather than routine lot inspection.

Do not convert a type/validation requirement into a routine production-lot test unless the source requires that frequency.

Where the source WI itself is not currently retrievable, mark unrecovered WI coverage as not verifiable rather than inventing content.

# 6. Standard Applicability Determination

| Standard | Scope determination | Applies directly? | How used here | Control conclusion |
| --- | --- | --- | --- | --- |
| ISO 5361:2016 | Tracheal tubes/connectors; explicitly excludes endobronchial (including tracheobronchial) tubes. | YES if product is true Endotracheal Tube. | Conditional direct-compliance track. | Report title points here, but mapped WI does not. |
| ISO 16628:2022 | Tracheobronchial tubes; double-lumen devices used to isolate one lung from the other. | YES if product is Endobronchial/Tracheobronchial double-lumen. | Conditional direct-compliance track matching the WI family. | WI title points here, but report design is largely single-lumen ETT style. |
| BS EN ISO 8836:2014 | Open/closed suction catheters used to suction the respiratory tract. | NO for the Endotracheal Tube itself. | Scope check only. | Do not import its tests into this product report unless a separate suction catheter is itself being evaluated. |

Decision required: The controlled product specification/intended use must state whether this is a single-lumen tracheal/endotracheal tube or a double-lumen tracheobronchial/endobronchial tube. Until that is fixed, neither the WI nor the report can be considered a reliable compliance record set.

# 7. ISO Requirements Review

## 7.1 ISO 16628:2022 — requirements relevant if the device is tracheobronchial

| Clause | Requirement | Page | Control type |
| --- | --- | --- | --- |
| 4.1 / 5.1 / 6.1 / 7 / 8 / 9.1 | ISO 18190:2016 requirements are incorporated for general, material, design, sterile, packaging and manufacturer-information controls. | PDF pp.8–13 (ISO pp.2–7) | Mandatory by normative incorporation; source ISO 18190 was not supplied. |
| 5.2 | Evaluate/test biological safety of the breathing-gas pathway in conformity with ISO 18562-1; verify in technical file. | PDF pp.8–9 (ISO pp.2–3) | Technical-file evidence. |
| 6.2 | Designated size in Charrière/French gauge; whole-number designation based on average tracheal-segment major/minor OD. | PDF p.9 (ISO p.3) | Design requirement. |
| 6.3 | Major and minor tracheal-segment OD axes tolerance: ±0.5 mm. | PDF p.9 (ISO p.3) | Numeric design/inspection criterion. |
| 6.4.1 | Machine-end connectors: 15 mm conical connectors conforming to ISO 5356-1. | PDF p.9 (ISO p.3) | Connector compliance. |
| 6.4.2 | Cuff-inflation inlets compatible with L1 male Luer-slip connector of ISO 80369-7. | PDF p.9 (ISO p.3) | Connector compliance. |
| 6.5.1 | Cuffs integral and leak-free; inflate to 9.0 kPa (90 cmH2O) or 1.5× cuff diameter, immerse, observe ≥10 s; no bubbles. | PDF pp.9–10 (ISO pp.3–4) | Functional test. |
| 6.5.2 | Tracheal and bronchial cuff diameters within ±15% of marked value using ISO 5361:2016 Annex B. | PDF p.10 (ISO p.4) | Numeric test. |
| 6.5.3 | Cuff herniation per ISO 5361:2016 Annex D; cuff must not extend beyond nearest patient-end edge. | PDF p.10 (ISO p.4) | Functional test. |
| 6.6 | Cuff inflation systems shall meet ISO 5361:2016 clause 5.6; verify by functional testing. | PDF p.10 (ISO p.4) | Functional/dimensional control. |
| 9.2.1 | Marking durability/legibility after intended exposure; specified visual/rub verification. | PDF p.12 (ISO p.6) | Marking verification. |
| 9.2.2 / 9.3 / 9.4 | Tube and package marking includes Fr size, bronchial OD, effective IDs, length marks, right/left designation; bronchial components blue; package includes major/minor axes and cuff diameters. | PDF pp.12–13 (ISO pp.6–7) | Label/marking controls. |
| Annex B / C | Normative methods for bronchial OD and effective inside diameters. | PDF pp.16–18 (ISO pp.10–12) | Type/specification verification or controlled inspection as applicable. |

## 7.2 ISO 5361:2016 — requirements relevant if the device is a true Endotracheal Tube

| Clause | Requirement | Page | Control type |
| --- | --- | --- | --- |
| 1 | Scope includes oro-/naso-tracheal tubes and specialized tracheal tubes; endobronchial/tracheobronchial tubes are excluded. | PDF p.7 (ISO p.1) | Applicability gate. |
| 5.1 / 5.2.1 | Size designation and tube dimensions, including marked ID/OD tolerances and size-dependent minimum lengths. | PDF pp.12–14 (ISO pp.6–8) | Numeric design/inspection control. |
| 5.2.2 | Connector dimensional requirements; machine end is a 15 mm male conical connector to ISO 5356-1. | PDF pp.16–18 (ISO pp.10–12) | Connector control. |
| 5.3 | Biological safety; durable/legible marking; reuse-material controls if applicable. | PDF p.19 (ISO p.13) | Technical-file/material control. |
| 5.4 | Bevel angle and patient-end sharp-edge requirements. | PDF p.19 (ISO p.13) | Visual/dimensional control. |
| 5.5.1–5.5.7 | Cuff leak, cuff position/diameter, tube collapse, herniation, tracheal seal and sharp-edge requirements. | PDF pp.19–20 (ISO pp.13–14) | Functional/type-test mix. |
| 5.6 | Inflation-system dimensions, angle, pilot-balloon function, deflation, connector and separation-distance controls. | PDF p.20 (ISO p.14) | Functional/dimensional control. |
| 5.7 | Curvature and smooth internal/external surface requirements. | PDF pp.20–21 (ISO pp.14–15) | Design/inspection control. |
| 5.8 | Radiopacity performance if the tube is labelled radiopaque. | PDF p.22 (ISO p.16) | Conditional test. |
| 5.9 / Annex H | Kink resistance; ball must pass freely after prescribed bend/conditioning test. | PDF pp.22, 45–46 (ISO pp.16, 39–40) | Functional/type or periodic verification. |
| 6 | Murphy eye requirements if present: area, location, collapse risk and sharp-edge control. | PDF p.23 (ISO p.17) | Conditional design/inspection. |
| 7 / 8 | Sterility/sterile packaging and device/package/insert marking. | PDF pp.23–27 (ISO pp.17–21) | Validation + labeling/inspection. |

# 8. ISO ↔ WI ↔ Inspection Report Unified Gap Matrix

Legend: “Not verifiable” means the original WI PDF was not retrievable in the current file layer and the requirement was not present in the retained project extract. It is not automatically treated as an implementation failure.

| ID | ISO Requirement | ISO Clause | ISO Page | WI Coverage | WI Page | Inspection Report Coverage | Report Row/Page | Status | Priority |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| DC-01 | Product / standard classification | ISO 5361 §1; ISO 16628 §1 | 5361 PDF p.7 / 16628 PDF p.7 | WI is Endobronchial; project previously mapped it to this report. | WI identity retained from prior analysis | Report is titled Endotracheal Tube and uses mostly single-lumen ETT characteristics. | Header + Rows 3–23 / p.1 | Document Control Gap | P1 / High |
| DC-02 | WI linkage on report | Controlled-document linkage | Report p.1 | Mapped WI known from project evidence. | N/A | “According to WI:” field is blank; Doc.# is cropped in image, so exact form identity cannot be confirmed. | Header / p.1 | Document Control Gap | P2 / Medium |
| DC-03 | Correct normative test-method references for cuff diameter and herniation | ISO 16628 §§6.5.2–6.5.3; ISO 5361 Annex B / Annex D | 16628 PDF p.10; 5361 PDF pp.32,36 | Retained WI extract states cuff diameter with ISO 5361 Annex C and herniation with Annex E. | WI p.3/11, §§5.1.7–5.1.8 | Report has no cuff-diameter or herniation rows. | No rows / p.1 | Document Control Gap | P1 / High |
| TBT-01 | [If tracheobronchial] General/risk requirements incorporated from ISO 18190 | ISO 16628 §4.1 | PDF p.8 (ISO p.2) | Not verifiable from retained WI evidence; prior project review identified ISO 18190 linkage as insufficient. | Not revalidated | Not expected in this report; should link to risk/technical file evidence. | NOT EXPECTED IN THIS REPORT | Evidence Gap | P1 / High |
| TBT-02 | [If tracheobronchial] Biological safety of breathing-gas pathway | ISO 16628 §5.2 | PDF pp.8–9 (ISO pp.2–3) | No retrievable WI evidence linking ISO 18562-1/technical-file verification. | Not revalidated | Not expected as lot inspection; no linked evidence field exists. | NOT EXPECTED IN THIS REPORT | Evidence Gap | P1 / High |
| TBT-03 | [If tracheobronchial] Designated size in Fr / average tracheal OD | ISO 16628 §6.2 | PDF p.9 (ISO p.3) | Prior project review indicated size controls existed, but exact WI page is not currently revalidated. | Not revalidated | Row 9 says “As per specification O,D, I,D” without Fr rule, numeric limit or actual value. | Row 9 / p.1 | Partial Gap | P1 / High |
| TBT-04 | [If tracheobronchial] Major/minor tracheal OD axes tolerance ±0.5 mm | ISO 16628 §6.3 | PDF p.9 (ISO p.3) | Prior project review indicated this was addressed; exact current WI evidence unavailable. | Not revalidated | No major/minor-axis fields or ±0.5 mm acceptance field. | Row 9 generic / p.1 | Evidence Gap | P1 / High |
| TBT-05 | [If tracheobronchial] 15 mm machine-end connector to ISO 5356-1 | ISO 16628 §6.4.1 | PDF p.9 (ISO p.3) | Prior review found connector coverage partial; 15 mm requirement was not explicit. | Not revalidated | Rows 4–5 use “Good fitting / According to size” and visual status only. | Rows 4–5 / p.1 | Partial Gap | P1 / High |
| TBT-06 | [If tracheobronchial] Cuff-inflation inlet compatible with ISO 80369-7 L1 male Luer slip | ISO 16628 §6.4.2 | PDF p.9 (ISO p.3) | Prior project review identified this as missing/insufficient. | Not revalidated | No connector-compatibility requirement/record field. | Rows 11,18–20 do not verify it | Gap | P1 / High |
| TBT-07 | [If tracheobronchial] Cuff leak-free test: 9.0 kPa or 1.5× diameter; water; ≥10 s; no bubbles | ISO 16628 §6.5.1 | PDF pp.9–10 (ISO pp.3–4) | Covered by WI §§5.1.5–5.1.6. | WI p.3/11 | Rows 19–20 only state “Should not be leak”; no pressure, immersion, duration, cuff identity or actual result. | Rows 19–20 / p.1 | Evidence Gap | P1 / High |
| TBT-08 | [If tracheobronchial] Tracheal/bronchial cuff diameter within ±15% of marked value | ISO 16628 §6.5.2; ISO 5361 Annex B | 16628 PDF p.10; 5361 PDF p.32 | Criterion present, but retained WI reference points to wrong Annex C. | WI p.3/11, §5.1.7 | No cuff-diameter measurement/marked-value field. | No row / p.1 | Evidence Gap | P1 / High |
| TBT-09 | [If tracheobronchial] Cuff herniation test | ISO 16628 §6.5.3; ISO 5361 Annex D | 16628 PDF p.10; 5361 PDF pp.36–37 | Requirement present, but retained WI reference points to wrong Annex E. | WI p.3/11, §5.1.8 | No cuff-herniation item/result field. | No row / p.1 | Evidence Gap | P1 / High |
| TBT-10 | [If tracheobronchial] Cuff/tube transition free of sharp edges | ISO 16628 §6.5.4 | PDF p.10 (ISO p.4) | Not verifiable from retained WI evidence. | Not revalidated | Cuff row says HVLP; no cuff attachment/transition edge check. | Row 13 / p.1 | Partial Gap | P2 / Medium |
| TBT-11 | [If tracheobronchial] Cuff inflation system complies with ISO 5361 §5.6 | ISO 16628 §6.6 | PDF p.10 (ISO p.4) | Prior review found partial inflation-system controls; complete §5.6 coverage not demonstrated. | Not revalidated | Rows 11,18–20 cover kink/printing/leak only; no dimensional limits, deflation function or connector compliance. | Rows 11,18–20 / p.1 | Partial Gap | P1 / High |
| TBT-12 | [If tracheobronchial] Bronchial segment / cuff position / margin-of-safety rationale | ISO 16628 §6.7 | PDF p.10 (ISO p.4) | No retrievable risk-file linkage. | Not revalidated | Not expected in routine report; no linked design/risk evidence. | NOT EXPECTED IN THIS REPORT | Evidence Gap | P1 / High |
| TBT-13 | [If tracheobronchial] Sterile-device requirements incorporated from ISO 18190 | ISO 16628 §7 | PDF p.11 (ISO p.5) | Not verifiable from retained WI evidence. | Not revalidated | Row 23 “Sterile Or Bulk” is only a status statement, not sterility-assurance evidence. | Row 23 / p.1 | Evidence Gap | P1 / High |
| TBT-14 | [If tracheobronchial] Packaging requirements incorporated from ISO 18190 | ISO 16628 §8 | PDF p.11 (ISO p.5) | Prior project review indicated packaging/labeling was handled elsewhere/partially; source not revalidated. | Not revalidated | Row 1 “Good without any damage” is insufficient for sterile-barrier/package compliance. | Row 1 / p.1 | Evidence Gap | P2 / Medium |
| TBT-15 | [If tracheobronchial] Marking durability/legibility verification | ISO 16628 §9.2.1 | PDF p.12 (ISO p.6) | Covered by WI §5.1.26 with illumination/rub conditions. | WI p.5/11 | Rows 15–18 verify only printing clarity/straightness; no durability/rub test field. | Rows 15–18 / p.1 | Evidence Gap | P2 / Medium |
| TBT-16 | [If tracheobronchial] Tube marking: Fr size, bronchial OD “Br”, effective IDs | ISO 16628 §9.2.2(a–c) | PDF p.12 (ISO p.6) | Not verifiable in retained WI evidence. | Not revalidated | Rows 9/16 are generic; no Fr, Br-OD, two-lumen effective-ID evidence. | Rows 9,16 / p.1 | Gap | P1 / High |
| TBT-17 | [If tracheobronchial] Length marks at 2 cm intervals and right/left designation | ISO 16628 §9.2.2(d–e) | PDF pp.12–13 (ISO pp.6–7) | Graduation-scale criterion retained; right/left designation not demonstrated in retrievable WI evidence. | WI p.2/11 (graduation only) | Row 16 “Clear Printed” only; no 2 cm interval basis or right/left designation. | Row 16 / p.1 | Partial Gap | P2 / Medium |
| TBT-18 | [If tracheobronchial] Bronchial segment/cuff/pilot balloon entirely blue | ISO 16628 §9.3 | PDF p.13 (ISO p.7) | Covered by WI §5.1.27; WI also defines tracheal components as transparent. | WI p.5/11 | No color-coding row. | No row / p.1 | Evidence Gap | P2 / Medium |
| TBT-19 | [If tracheobronchial] Package/insert: description, major/minor axes, cuff diameters | ISO 16628 §9.4 | PDF p.13 (ISO p.7) | Not fully verifiable; prior review indicated package/label controls existed outside this report. | Not revalidated | Not expected in this inspection table unless linked to packaging/label inspection record. | NOT EXPECTED IN THIS REPORT | Evidence Gap | P2 / Medium |
| TBT-20 | [If tracheobronchial] Bronchial OD test method for marking | ISO 16628 Annex B | PDF p.16 (ISO p.10) | No retrievable WI method evidence. | Not revalidated | No method/actual result field for bronchial OD. | Row 9 generic / p.1 | Gap | P1 / High |
| TBT-21 | [If tracheobronchial] Effective ID test for both lumens | ISO 16628 Annex C | PDF pp.17–18 (ISO pp.11–12) | No retrievable WI method evidence. | Not revalidated | No bead-test or two-lumen effective-ID result field. | Row 9 generic / p.1 | Gap | P1 / High |
| ETT-01 | [If true Endotracheal] Size designation + ID/OD tolerances + length | ISO 5361 §§5.1, 5.2.1; Table 1a | PDF pp.12–13 (ISO pp.6–7) | Mapped WI is for Endobronchial Tube; therefore it is not a valid direct WI basis for an ordinary ETT. | WI mismatch | Rows 9–10 say “As per specification / According to Size” but do not show controlled limits or actuals. | Rows 9–10 / p.1 | Gap | P1 / High |
| ETT-02 | [If true Endotracheal] Connector dimensions / 15 mm machine end | ISO 5361 §5.2.2 | PDF pp.16–17 (ISO pp.10–11) | Mapped WI is the wrong product family for direct ETT compliance. | WI mismatch | Rows 4–5 are qualitative only; no 15 mm / ISO 5356-1 evidence. | Rows 4–5 / p.1 | Gap | P1 / High |
| ETT-03 | [If true Endotracheal] Biological safety | ISO 5361 §5.3.1 | PDF p.19 (ISO p.13) | Wrong WI family; direct ETT material-control linkage not demonstrated. | WI mismatch | Not expected in routine lot report; needs technical-file/biocompatibility evidence. | NOT EXPECTED IN THIS REPORT | Evidence Gap | P1 / High |
| ETT-04 | [If true Endotracheal] Bevel angle and sharp-edge control | ISO 5361 §5.4 | PDF p.19 (ISO p.13) | Endobronchial WI tip criteria cannot establish full ETT clause coverage. | WI mismatch | Row 3 checks smooth beveled tip but no angle acceptance or separate sharp-edge criterion. | Row 3 / p.1 | Partial Gap | P1 / High |
| ETT-05 | [If true Endotracheal] Cuff leak-free test | ISO 5361 §5.5.1 | PDF p.19 (ISO p.13) | Retained WI has equivalent leak method but belongs to Endobronchial WI. | WI p.3/11 (mapped only) | Rows 19–20 lack test pressure, immersion, ≥10 s and actual result. | Rows 19–20 / p.1 | Evidence Gap | P1 / High |
| ETT-06 | [If true Endotracheal] Cuff dimension / position controls | ISO 5361 §§5.5.2–5.5.3; Annex B | PDF p.19; Annex B PDF p.32 | Cuff diameter criterion retained; product-family mismatch remains. | WI p.3/11, §5.1.7 (mapped only) | No cuff position or measured cuff-diameter fields. | No row / p.1 | Evidence Gap | P1 / High |
| ETT-07 | [If true Endotracheal] Cuffed-tube collapse + cuff herniation | ISO 5361 §§5.5.4–5.5.5; Annex C/D | PDF pp.19–20; Annexes PDF pp.33–37 | Herniation requirement retained but wrong Annex reference; tube-collapse evidence not demonstrated. | WI p.3/11 + not verifiable | No tube-collapse or herniation rows. | No row / p.1 | Gap | P1 / High |
| ETT-08 | [If true Endotracheal] Tracheal seal performance | ISO 5361 §5.5.6; Annex G | PDF p.20; Annex G PDF pp.42–44 | No retrievable direct ETT WI evidence. | WI mismatch | NOT EXPECTED as routine lot test; no linked type-test evidence shown. | NOT EXPECTED IN THIS REPORT | Evidence Gap | P1 / High |
| ETT-09 | [If true Endotracheal] Cuff inflation system controls | ISO 5361 §5.6 | PDF p.20 (ISO p.14) | Mapped WI is not a valid direct ETT WI; prior project review found partial inflation-system controls. | WI mismatch | Rows 11,18–20 do not cover complete §5.6 functions/dimensions. | Rows 11,18–20 / p.1 | Partial Gap | P1 / High |
| ETT-10 | [If true Endotracheal] Curvature and smooth inside/outside surfaces | ISO 5361 §5.7 | PDF pp.20–21 (ISO pp.14–15) | Direct ETT WI coverage not demonstrated. | WI mismatch | Row 8 says comply with drawing/specification; no controlled reference or actual curvature verification. | Row 8 / p.1 | Partial Gap | P2 / Medium |
| ETT-11 | [If labelled radiopaque] Radiopaque marker performance | ISO 5361 §5.8 | PDF p.22 (ISO p.16) | Direct ETT WI coverage not demonstrated. | WI mismatch | Row 15 checks printed straight X-ray line but not radiographic distinguishability/test evidence. | Row 15 / p.1 | Partial Gap | P2 / Medium |
| ETT-12 | [If true Endotracheal] Kink resistance | ISO 5361 §5.9; Annex H | PDF p.22; Annex H PDF pp.45–46 | Direct ETT WI coverage not demonstrated. | WI mismatch | Row 2 states “Kink resistance / Termo sensitive tube” but no Annex H method or controlled evidence route. | Row 2 / p.1 | Partial Gap | P1 / High |
| ETT-13 | [If Murphy eye present] Murphy eye size/location/collapse/sharp-edge requirements | ISO 5361 §6 | PDF p.23 (ISO p.17) | Direct ETT WI coverage not demonstrated. | WI mismatch | Row 14 only checks smooth edge; does not verify area ≥80%, opposite-bevel location or collapse risk. | Row 14 / p.1 | Partial Gap | P2 / Medium |
| ETT-14 | [If supplied sterile] Sterility assurance and sterile packaging | ISO 5361 §7 | PDF pp.23–24 (ISO pp.17–18) | Direct ETT WI/validation linkage not demonstrated. | WI mismatch | Row 23 “Sterile Or Bulk” does not link validation/sterile-barrier evidence. | Row 23 / p.1 | Evidence Gap | P1 / High |
| ETT-15 | [If true Endotracheal] Tube/pack/insert marking requirements | ISO 5361 §8 | PDF pp.24–27 (ISO pp.18–21) | Direct ETT WI basis not demonstrated. | WI mismatch | Rows 15–18 cover selected printing only; no complete label/pack checklist or controlled reference. | Rows 15–18 / p.1 | Partial Gap | P2 / Medium |
| NA-01 | BS EN ISO 8836:2014 suction-catheter requirements | ISO 8836 §1 | PDF p.12 (ISO p.1) | Not applicable to the Endotracheal Tube itself. | N/A | Do not add suction-catheter tests to this report unless a separate suction catheter is the inspected device/component. | N/A | Not Applicable | P3 / Low |

# 9. Critical Findings

1. Product classification / document set is not controlled
The report is Endotracheal while the mapped WI is Endobronchial. ISO 5361 and ISO 16628 have different scopes. This is not only a naming issue because the WI/report content also differs (double-lumen components versus standard ETT attributes).

2. WI points to the wrong ISO 5361 annexes for two cuff tests
The retained WI extract links cuff diameter to Annex C and cuff herniation to Annex E. ISO 5361:2016 uses Annex B for cuff diameter, Annex C for cuffed-tube collapse, Annex D for cuff herniation and Annex E only as informative connector guidance.

3. Cuff leakage is implemented in the WI but not recordable with adequate evidence
The report’s “Should not be leak” entries omit the actual pressure/condition, water immersion, minimum 10 s observation, cuff identity and actual result.

4. Cuff diameter and herniation have no dedicated report evidence
The WI contains these checks, but the report has no fields to record measured cuff diameter, marked value, tolerance result or herniation result.

5. Dimensional and connector requirements are too generic
“As per specification” / “According to size” / “Good fitting” are not sufficient by themselves. Either the numeric limit must appear on the controlled report or the report must identify the exact approved specification/drawing and revision.

6. Technical-file requirements are not linked
Risk-management, biological-safety, sterility/packaging and type-test requirements should not be forced into each lot inspection report, but the WI/report system must identify the controlled evidence that demonstrates them.

# 10. Document Control Findings

| ID | Finding | Evidence | Priority | Required control |
| --- | --- | --- | --- | --- |
| DC-01 | Product naming/scope mismatch | Report: Endotracheal Tube; WI: Endobronchial Tube. | P1 / High | Establish controlled product classification and issue a matching WI/report set. |
| DC-02 | Blank WI traceability field | “According to WI:” is blank on the report template. | P2 / Medium | Preprint or control the WI number/revision in the form master. |
| DC-03 | Form number not verifiable from source image | Doc.# is cropped as “F-823-T…”. | P2 / Medium | Confirm current controlled form number and compare with the WI’s referenced form before release. |
| DC-04 | Wrong normative annex references in WI | Cuff diameter → Annex C and herniation → Annex E in retained WI extract. | P1 / High | Correct to ISO 5361 Annex B and Annex D respectively; assess change impact on training/records. |
| DC-05 | Ambiguous qualitative criteria | Examples: “Good fitting”, “Strong Gluing”, “Should not be leak”, “Clear Printed”. | P2 / Medium | Replace with measurable criteria or exact controlled-specification reference and result field. |
| DC-06 | Conditional configuration not structured | Cuffed/uncuffed, oral/nasal/reinforced, Murphy eye, radiopaque and sterile/bulk conditions share one generic form. | P2 / Medium | Add N/A logic and configuration-controlled rows so only applicable requirements are executed. |

# 11. Recommended Actions

| ID | Requirement / Area | Current State | Required Action | Priority |
| --- | --- | --- | --- | --- |
| A-01 | Product classification | Endotracheal report is paired with Endobronchial WI; applicable standard is unresolved. | Freeze the document pair until the approved product specification confirms single-lumen ETT versus double-lumen tracheobronchial/endobronchial. Then assign the correct primary standard and matching WI/report. | P1 / High |
| A-02 | WI normative references | WI directs cuff diameter/herniation to wrong ISO 5361 annexes. | Revise cuff-diameter method reference to Annex B and cuff-herniation reference to Annex D. Add a document-change impact assessment for prior records/training. | P1 / High |
| A-03 | Cuff leakage evidence | Report rows 19–20 only state “Should not be leak”. | Add Cuff ID (tracheal/bronchial where applicable), test pressure/1.5× condition, water immersion, minimum 10 s observation, actual observation and Pass/Fail. | P1 / High |
| A-04 | Cuff diameter evidence | No row records marked/actual cuff diameter. | Add Marked Diameter, ±15% Allowed Range, Actual Diameter, Result, test method reference and cuff identity. | P1 / High |
| A-05 | Cuff herniation evidence | No report row. | Add a dedicated herniation result line linked to ISO 5361 Annex D; record Pass/Fail and applicable cuff identity. | P1 / High |
| A-06 | Dimensions | Rows 9–10 use generic “specification/size” wording. | For routine inspection, either print size-specific ID/OD/length limits on the form or require entry of controlled specification/drawing number and revision plus actual values. | P1 / High |
| A-07 | Connector compliance | Report checks fitting/appearance only. | If tracheobronchial: document 15 mm machine connector and ISO 80369-7 cuff-inflation connector evidence. If ETT: document ISO 5361/ISO 5356-1 connector requirements. Use supplier/spec/type-test evidence where appropriate. | P1 / High |
| A-08 | Inflation system | Current rows do not demonstrate complete functional requirements. | Add or link controlled checks for pilot-balloon function, deflation capability and required connector/dimensional criteria. Do not force non-lot design verification into every lot record if controlled by qualification/type testing. | P1 / High |
| A-09 | Marking / coding | Printing clarity is present; durability/coding and product-specific mandatory markings are incomplete. | Add marking-rub/durability evidence, color-coding checks for tracheobronchial devices, and the correct size/OD/ID/right-left/oral-nasal fields according to the confirmed product family. | P2 / Medium |
| A-10 | Risk / biological safety / sterility / packaging | Standards require technical-file evidence that is not a normal dimensional lot check. | Create explicit WI references to the Risk Management File, ISO 18562/biocompatibility evidence, sterilization validation and packaging validation/records. Inspection report may show a reference/status rather than repeating validation tests. | P1 / High |
| A-11 | Configuration logic | The form mixes cuffed/uncuffed, reinforced, Murphy eye, radiopaque and sterile/bulk options. | Add applicability/N/A controls and require the inspected configuration to be selected before sampling/testing. | P2 / Medium |
| A-12 | Document traceability | WI number field is blank and form number is not confirmable from the source image. | Control the report master with fixed Doc.#, Rev., effective date and linked WI number/revision. Confirm whether the WI-referenced form and current Rev.4 form are the same controlled record. | P2 / Medium |

# 12. Priority Action Plan

| Action | Deliverable | Owner | Timing | Closure Evidence |
| --- | --- | --- | --- | --- |
| P1-01 | Confirm actual device classification and intended use; lock the governing standard. | QCM / RA / Product Owner | Before revising WI/report | Closed classification decision + approved spec. |
| P1-02 | Correct WI Annex B / Annex D references and issue controlled revision. | Document Control / QC | Immediate | Approved WI revision. |
| P1-03 | Revise report for cuff leak, cuff diameter and herniation evidence fields. | QC | Immediate | Controlled report revision. |
| P1-04 | Add controlled numeric/spec references for ID, OD, length and connector requirements. | QC / Engineering | Before next controlled release | Approved spec/drawing references in form. |
| P1-05 | Establish links to technical-file evidence: risk, biological safety, type tests, sterilization and packaging. | QA / RA / QC | Before compliance closure | Traceable evidence register. |
| P2-01 | Add marking durability, product-family-specific marking/color coding and configuration logic. | QC | Next revision cycle | Updated controlled form. |
| P2-02 | Confirm current report Doc.# against WI-linked form and complete header linkage. | Document Control | Next revision cycle | Verified form/WI cross-reference. |

# 13. Final Conclusion

The current controlled set cannot be rated fully compliant because the governing product classification is unresolved and the WI/report pair belongs to different standard families. The strongest evidence gaps are not cosmetic: the report cannot fully record the required cuff leakage conditions, cuff diameter verification or cuff herniation result, and the retained WI references the wrong ISO 5361 annexes for two of those tests.

Once the product family is confirmed, the corrective path is straightforward. For a true Endotracheal Tube, use ISO 5361:2016 as the direct uploaded standard and create/use a matching ETT WI. For a double-lumen Endobronchial/Tracheobronchial Tube, use ISO 16628:2022 as the direct standard and apply ISO 5361 only where ISO 16628 invokes it. BS EN ISO 8836:2014 should not be used as a governing ETT standard.

Several ISO requirements are correctly better controlled through the technical file, risk-management file, qualification/type testing, supplier evidence, sterilization validation or packaging validation rather than being repeated on every production-lot inspection record. The corrective action is therefore to create traceable evidence links, not to add every ISO test to every lot inspection.

# 14. Source Traceability

| Source | Traceability used |
| --- | --- |
| ISO 16628:2022 | Scope/normative references: PDF p.7; general/material requirements: pp.8–9; design/cuff: pp.9–10; sterile/packaging: p.11; marking: pp.12–13; Annex B: p.16; Annex C: pp.17–18. |
| ISO 5361:2016 | Scope: PDF p.7; general/size/dimensions: pp.11–18; materials/bevel/cuff: pp.19–20; curvature/radiopacity/kink: pp.20–22; Murphy/sterility/marking: pp.23–27; Annex B: p.32; Annex C: pp.33–35; Annex D: pp.36–37; Annex G: pp.42–44; Annex H: pp.45–46. |
| BS EN ISO 8836:2014 | Scope: PDF p.12 (suction catheters for respiratory-tract suction). |
| Prior QC project analysis | QC_Inspection_Report_WI_Content_Gap_Analysis.docx, pages 11–13 and 17: WI identity, retained WI page/section criteria, and mapping to the Endotracheal Tube report. |
| Current Inspection Report image | Inspection & Test Report for Endotracheal Tube, Rev.4, page 1/1; rows 1–23 transcribed from the uploaded image. Full Doc.# and date are cropped in the provided source image. |

WI evidence limitation: The original Wi.pdf used in the earlier project analysis was searched in the current conversation/library and was not directly retrievable. Therefore, this report uses only WI content that is preserved in the prior project analysis and marks other WI coverage as “not verifiable” instead of inferring it.

# Appendix A — Current Inspection Report Row Map

| Row | Description | Current requirement text |
| --- | --- | --- |
| 1 | Packing Condition | Good without any damage |
| 2 | Tube Status | Kink resistance / Termo sensitive tube |
| 3 | Tip Forming status | Smooth beveled tip |
| 4 | Fitting with Connector | Good Fitting / According to size |
| 5 | Connector Status | No flashing, black spot / smooth edge |
| 6 | Type | Cuffed / Uncuffed |
| 7 | Tube formation | Oral / Nasal / Oral nasal / Reinforced |
| 8 | Angle formation | comply with drawing / specification |
| 9 | Tube Dimension | As per specification O,D, I,D |
| 10 | Tube Length | According to Size |
| 11 | Inflation tube Status | Kink Resistance tube |
| 12 | Inflation tube Gluing | Strong Gluing, no leakage |
| 13 | Cuff Status | High volume low pressure cuff |
| 14 | Murphy Eye status | Smooth edged Murphy eye |
| 15 | X-Ray Line Marking | Clear Printed Straight |
| 16 | Graduation Scale | Clear Printed |
| 17 | Guide Mark | With dept Guide mark |
| 18 | Printing in pilot balloon | Printed clearly |
| 19 | Leak Test for Pilot balloon | Should not be leak |
| 20 | Leak Test for balloon | Should not be leak |
| 21 | Reinforced spring status | Good smooth fitting, circle shape |
| 22 | Logo | Mais Logo |
| 23 | Sterility Condition | Sterile Or Bulk |

# Appendix B — Source Report Snapshot

Uploaded Inspection & Test Report for Endotracheal Tube — source snapshot used for row-level mapping.