# EPIDURAL SET | ISO CLOSURE REGISTER

CLAUSE-LEVEL REQUIREMENT, APPLICABILITY & OBJECTIVE EVIDENCE REGISTER  |  DD-05-ISO-REG-DRAFT-01

Epidural Set — DD-05  |  Source WI-DD-05 Rev.01  |  Evaluation date: 22 Oct 2026

DOCUMENT STATUS: PROPOSED CONTROLLED DRAFT — NOT IMPLEMENTED / NOT RELEASED

Scope: Epidural Set in-process assembly and manual packaging controls. This document does not constitute final product inspection, final release, or evidence that product tests have passed.

Coverage register includes 52 individual clause and normative/informative entries from the two uploaded standards. It is scoped to identified syringe or connector components, never automatic complete-kit coverage. No required executed evidence was supplied; all operational closures remain unverified.

# 1. Applicability and document limitations

Current product variant, syringe configurations and connector routes are NOT confirmed. Requirements in ISO 7886-1 and ISO 80369-7 are therefore assessed as Conditional / Alternative / Configuration-dependent / Luer-lock-only / Informative rather than incorrectly marked compliant. Separate neuraxial ISO 80369-6 must be obtained before claiming neuraxial connection conformity.

# 2. Clause-to-Gap Coverage and Closure

| ISO requirement | ISO Says — requirement synopsis | PDF / ISO printed page | Applicability | Gap / WI change IDs | Required Evidence / Route | Verified available evidence | Documentary / Operational closure status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 7886 1 | Scope and non-lot-release limitation | PDF 9 / printed 1 | Conditional | GAP-001,029 | Syringe model / intended-use applicability decision | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 5 | General appearance, flange and packaging design inputs | PDF 13 / printed 5 | Conditional | GAP-014,015,023 | Design validation and visual qualification | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 6.1 | Fluid-contact surface foreign matter | PDF 13 / printed 5 | Conditional | GAP-006 | Visual incoming/QC report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 6.2 | Extract pH shift not above 1.0 pH unit | PDF 13 / printed 5 | Conditional | GAP-012 | Annex A test report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 6.3 | Extractable metals Pb/Sn/Zn/Fe ≤5 mg/kg, Cd <0.1 mg/kg | PDF 13 / printed 5 | Conditional | GAP-012 | Analytical qualification report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 7 | Lubricant type/amount/location | PDF 14 / printed 6 | Conditional | GAP-013 | Lubricant qualification/process record | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 8 | Graduated-volume tolerance Table 1 | PDF 14–15 / printed 6–7 | Conditional | GAP-007 | Volume accuracy design data | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 9.1 | Scale graduation characteristics | PDF 15–16 / printed 7–8 | Conditional | GAP-008 | Scale dimensional drawings | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 9.2 | Scale numbering | PDF 16 / printed 8 | Conditional | GAP-008 | Printed scale artwork/inspection | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 9.3 | Scale length per Table 1 | PDF 16 / printed 8 | Conditional | GAP-008 | Dimensional qualification | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 9.4 | Zero graduation alignment | PDF 17 / printed 9 | Conditional | GAP-008 | Design and visual records | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 10.1 | Maximum capacity risk assessment | PDF 17 / printed 9 | Conditional | GAP-014 | Design risk controls | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 10.2 | Barrel flanges / no sharp flash / roll | PDF 17 / printed 9 | Conditional | GAP-014 | Flange qualification & visual | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 11.1 | Piston design and stopper retention | PDF 17 / printed 9 | Conditional | GAP-014 | Retention and drawings | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 12.1 | Conical nozzle design connector conformity | PDF 18 / printed 10 | Conditional | GAP-009,016 | Connector design qualification | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 12.2 | Nozzle center/eccentric placement | PDF 18 / printed 10 | Conditional | GAP-009 | Approved nozzle drawing | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 12.3 | Nozzle lumen ≥1.2 mm | PDF 18 / printed 10 | Conditional | GAP-009 | Nozzle dimensional report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 13.1 | Dead space Table 1 | PDF 18 / printed 10 | Conditional | GAP-010 | Annex C dead space report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 13.2 | Air & water tightness stopper | PDF 18 / printed 10 | Conditional | GAP-011 | Annex B and D reports | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 13.3 | Piston force measurement recommended (not mandatory) | PDF 18 / printed 10 | Recommended | GAP-029 | Approved optional procedure if adopted | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 13.4 | No gravity-induced piston motion | PDF 18 / printed 10 | Conditional | GAP-014 | Design test report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 14.1 | Unit packaging / self-contained syringe | PDF 19 / printed 11 | Conditional | GAP-015 | Applicable pack configuration and validation | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 14.2 | Multiple unit packs | PDF 19 / printed 11 | Configuration-dependent | GAP-015 | Packaging evidence if configuration used | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 14.3 | User packaging | PDF 19 / printed 11 | Conditional | GAP-015 | Distribution packaging validation | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.1 | Manufacturer information | PDF 20 / printed 12 | Conditional | GAP-015,024 | Approved label/artwork | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.2 | Syringe marking | PDF 20 / printed 12 | Conditional | GAP-008,015 | Barrel marking qualification | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.3 | Unit packaging label | PDF 20–21 / printed 12–13 | Conditional | GAP-015 | Artwork validation | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.4 | Multiple pack label | PDF 21 / printed 13 | Configuration-dependent | GAP-015 | Pack label applicability and approval | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.5 | User packaging label | PDF 21–22 / printed 13–14 | Conditional | GAP-015 | Labeling control | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.6 | Storage container marks | PDF 22 / printed 14 | Configuration-dependent | GAP-015 | Packaging spec and mark inspection | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 15.7 | Transport wrapping | PDF 22 / printed 14 | Configuration-dependent | GAP-015 | Packaging configuration control | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 Annex A (normative) | Extract prep for 6.2–6.3 | PDF 23 / printed 15 | Conditional | GAP-012 | Executed extraction worksheets | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 Annex B (normative) | Plunger aspiration air leakage/retention | PDF 24–25 / printed 16–17 | Conditional | GAP-011,014 | Executed vacuum test | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 Annex C (normative) | Dead space determination | PDF 26 / printed 18 | Conditional | GAP-010 | Mass/volume dead space test | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 Annex D (normative) | Plunger liquid leakage compression | PDF 27–28 / printed 19–20 | Conditional | GAP-011 | Executed pressure test | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 Annex E (informative) | Suggested piston force test | PDF 29–33 / printed 21–25 | Informative | GAP-029 | Optional if adopted in spec | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 7886 Annex F (informative) | Suggested silicone oil method | PDF 33–36 / printed 25–28 | Informative | GAP-013 | Optional measurement route | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 1 | Intravascular/hypodermic connector scope | PDF 8 / printed 1 | Conditional | GAP-001,016 | Approved route matrix | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 4.1 | General and non-interconnectable characteristics | PDF 10 / printed 3 | Conditional | GAP-001,016 | Design/risk and CAD/test evidence | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 4.2 | Material modulus >700 MPa | PDF 11 / printed 4 | Conditional | GAP-017 | ASTM test record | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 4.3 | Type test evidence | PDF 11 / printed 4 | Conditional | GAP-017,028 | Approved type-test package | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 5 | Luer variant geometry Annex B | PDF 11 / printed 4 | Conditional | GAP-016 | Specific B.1–B.8 report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.1.1 | Fluid leakage: approved test method selection | PDF 11 / printed 4 | Conditional | GAP-018 | Qualification plan | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.1.2 | Pressure-decay option | PDF 11 / printed 4 | Alternative | GAP-018 | Pressure-decay test, if selected | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.1.3 | Positive-water pressure option | PDF 12 / printed 5 | Alternative | GAP-018 | Water-drop test, if selected | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.2 | Subatmospheric air leak | PDF 12 / printed 5 | Conditional | GAP-019 | Air-leak test report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.3 | Stress cracking | PDF 12 / printed 5 | Conditional | GAP-020 | Stress/post-test leakage report | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.4 | Axial separation | PDF 12 / printed 5 | Conditional | GAP-021 | Axial load test | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.5 | Unscrewing (L2 only) | PDF 12 / printed 5 | Luer lock only | GAP-021 | Torque test | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 6.6 | Overriding (L2 only) | PDF 12–13 / printed 5–6 | Luer lock only | GAP-021 | Overriding test | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 Annex B (normative) | Luer drawings/tolerances | PDF 18–28 / printed 11–21 | Conditional | GAP-016 | Type-variant-specific dimensions | Not provided / Pending | Open — Evidence & applicability; no verified closure |
| 80369-7 Annex C (normative) | Reference connectors | PDF 29–35 / printed 22–28 | Conditional | GAP-018,019,020,021 | Test fixture reference ID | Not provided / Pending | Open — Evidence & applicability; no verified closure |

# 3. Cross-Gap Action Closure

| Gap ID / priority | Closure responsibility | Current state | Exact evidence/acceptance basis still needed |
| --- | --- | --- | --- |
| GAP-001 P1 | Engineering / QA | Open — Technical Input | Approved component-route matrix, product risk assessment, approved BOM/drawings, correct standard(s), test reports and change-control decision. Route: Design qualification + incoming ID verification + assembly traceability. No test result provided. |
| GAP-002 P2 | Document Control / QA | Open — Approval | Master list, WI revision history, DD-05-PS, DD-05-PDBM and approved change record confirming cross-references. Route: Document Control. |
| GAP-003 P1 | Engineering / Production / QC | Open — Technical Input | Approved SKU/BOM matrix, component drawings, purchase specifications, first-article reconciliation, deviation/rework approval if any. Route: Engineering + incoming + routine assembly. |
| GAP-004 P2 | QC / Supplier Quality | Open — Technical Input | Approved AQL procedure, item specifications, completed F-823-T017/T90/T37/T61 (as applicable), supplier lots and disposition. Route: Incoming; in-process links verified accepted lots. |
| GAP-005 P1 | Production / QC | Open — Evidence | Completed production order, controlled issue slips, BOM reconciliation, inventory and QC release statuses; completed in-process assembly evidence. Route: routine assembly. |
| GAP-006 P2 | QC | Open — Evidence | Approved method, recorded incoming visual results, sample details, defect log and linkage to assembled kit lot. Route: incoming + routine assembly spot check. |
| GAP-007 P1 | Engineering / Supplier Quality / QC | Open — Evidence | Verified calibrated test or supplier qualification report with Table 1 calculation for each applicable model and valid scope, QA approval. Route: design/type/supplier + incoming trace. |
| GAP-008 P2 | Engineering / QC | Open — Evidence | Approved drawing with 3/5 ml configuration, dimensional report, incoming/assembly records. Route: design/supplier + routine visual. |
| GAP-009 P1 | Engineering / QA | Open — Technical Input | Supplier design drawings, nozzle dimensional qualification, applicable connection standard evidence and QA-approved clinical use matrix. Route: design qualification + component identity. |
| GAP-010 P2 | Engineering / Supplier Quality | Open — Evidence | Approved and traceable Annex C test results for the applicable model; supplier qualification record. Route: design/type/supplier. |
| GAP-011 P1 | QC Laboratory / Supplier Quality | Open — Evidence | ISO 7886 Annex B and D completed reports, equipment calibration, supplier traceability, QA disposition. Route: design/type/supplier with risk-based periodic verification. |
| GAP-012 P1 | Engineering / QC Laboratory | Open — Evidence | Signed Annex A analytical reports and controlled material/supplier equivalence documents; test method and qualification scope. Route: supplier/design qualification + change control. |
| GAP-013 P2 | Engineering / Supplier Quality | Open — Evidence | Supplier formulation controls, lubricant quantitative design or process evidence, applicable engineering approval and change-history. Route: supplier/process qualification. |
| GAP-014 P2 | Engineering / QC | Open — Evidence | Verified design test report, approved drawing, incoming defect checks and component lot traceability. Route: qualification + routine visual. |
| GAP-015 P2 | QA / Regulatory / Packaging | Open — Technical Input | Approved packaging validation, labeled master artwork, kit labeling specification, traceable packaging records. Route: packaging design validation + routine label verification. |
| GAP-016 P1 | Engineering / Supplier Quality | Open — Technical Input | Signed relevant Annex B drawing/dimensional qualification and first-article/periodic data linked to each design and lot scope. Route: type/qualification, supplier incoming. |
| GAP-017 P2 | Engineering / Supplier Quality | Open — Evidence | Approved material certificate, ASTM report, design/type-test plan and QA approval. Route: design/supplier qualification. |
| GAP-018 P1 | Engineering / QC Laboratory | Open — Evidence | Completed ISO 80369-7/20 test report with instrument calibration, geometry, sample/design applicability and QA approval. Route: type/supplier qualification. |
| GAP-019 P1 | Engineering / QC Laboratory | Open — Evidence | Validated test report and approved configuration link; route type/supplier qualification. |
| GAP-020 P1 | Engineering / Supplier Quality | Open — Evidence | Approved stress-crack type-test report, post-exposure leak results, QA configuration approval. Route: type/supplier qualification. |
| GAP-021 P1 | Engineering / QC Laboratory | Open — Evidence | Approved design/type mechanical tests and calibration reports, supplier qualification and approved risk-based control plan. Route: type/supplier qualification. |
| GAP-022 P2 | Engineering / QC | Open — Technical Input | Approved component-specific criteria, appearance master / drawings, inspection work method, completed line records, traceable QC review. Route: routine in-process and incoming. |
| GAP-023 P1 | Packaging Engineering / Production / QA | Open — Technical Input | Packaging validation IQ/OQ/PQ or equivalent approved record, machine calibration, pouch specification, sampled seal width/peel strength/continuity records, investigation when out of limit. Route: process validation + routine checks. |
| GAP-024 P2 | QA / Regulatory / Production | Open — Approval | Approved kit artwork, labeling specifications, regulatory authorization and completed pouch/carton verification. Route: routine line + regulatory document control. |
| GAP-025 P2 | Supplier Quality / QC | Open — Technical Input | Supplier CoA/label, expiry/stability evidence, approved material specification and in-process lot/date records. Route: incoming + routine kit check. |
| GAP-026 P1 | Production / QC / QA | Open — Evidence | Completed hold tag, NCR, suspect quantity reconciliation, authorized disposition/recheck/release signatures. Route: real-time production control. |
| GAP-027 P1 | QA / QCM / Sterilization | Open — Evidence | Executed approved sterilization cycle, laboratory results, F-823-F20, final release authorization and traceable stock transfer. Route: sterilization validation + final release (separate). |
| GAP-028 P2 | QC / Document Control | Open — Evidence | Current approved and completed F-823-T017/T90/T37/T61/F20, DD-05-PS, DD-05-PDBM and related controlled testing forms. Route: evidence collection and QC review. |
| GAP-029 P1 | Engineering / QA | Open — Technical Input | Approved LOR syringe technical specification, intended-use risk assessment, compatible interface/functional validation and supplier report. Route: design/supplier qualification + in-process component verification. |
| GAP-030 P2 | Production / QC | Open — Technical Input | Approved risk-based control plan, workcenter sampling plan, completed actual timed checks, shift/changeover records and line clearance. Route: routine assembly and packaging. |

# 4. Evidence classification and closure gate

• Source verified = actual supplied document and its identity/clause/row have been checked. It does not mean any operational test is complete.

• Documentary Covered — Operational Evidence Pending = approved controlled documentation actually covers a confirmed clause but no valid execution evidence exists. For this proposed DRAFT no such approval is established.

• Verified Closed = only after verified approved control, applicability to specific component/lot/design scope and objective executed evidence. None assigned in this review.

• Open — Technical Input = missing approved product route, part variant, performance specification, validated setting or sampling decision.

• Open — Approval = revision/reference conflict or document changes remain pending controlled approval.

• N/A with Justification = only after QA and Engineering approve a route/configuration-specific rationale and cite actual drawing/standard.

# 5. Source Index

S-01 ISO 7886-1:2017(E) corrected Aug 2019 uploaded PDF 38 pages; S-02 ISO 80369-7:2016(E) uploaded PDF 53 pages; S-03 WI-DD-05 Rev.01, original six pages; S-04 F-823-i22 Rev.5 Excel original rows 2–39 (28 component lines rows 11–38); S-05 F-823-i4-2 Rev.6 Excel original rows 2–30 (16 characteristic lines rows 14–29). Referenced but NOT available: DD-05-PS, DD-05-PDBM, ISO 80369-6/20, approved company sampling/packaging validation, F-823-T017/T90/T37/T61/F20, completed QC checks and final release.