# EPIDURAL SET | IN-PROCESS CONTROL RECORD

EDITABLE PROPOSED FORM • ASSEMBLY + MANUAL PACKAGING  |  F-823-i22 / F-823-i4-2 • DRAFT

Epidural Set — DD-05  |  Source WI-DD-05 Rev.01  |  Evaluation date: 22 Oct 2026

DOCUMENT STATUS: PROPOSED CONTROLLED DRAFT — NOT IMPLEMENTED / NOT RELEASED

Scope: Epidural Set in-process assembly and manual packaging controls. This document does not constitute final product inspection, final release, or evidence that product tests have passed.

Purpose: Reconstructed execution form based on the 28-row assembly record and 16-row manual packaging record. It is NOT an independent final inspection report, type-test certificate, or product release. Values are not approved without verified current process specifications.

# A. Kit / Production Order Identification

| Product code / kit SKU | Production order | Kit lot / batch | BOM ID / revision | Shift / date | QC inspector / team |
| --- | --- | --- | --- | --- | --- |
| ________________________ | ________________________ | ________________________ | ________________________ | ________________________ | ________________________ |

| Clinical route matrix / risk ref | WI ID/rev validated | F-823-i22 original row source | Component issue slip | Line clearance / last good check | Approved sampling plan / interval |
| --- | --- | --- | --- | --- | --- |
| ___________________ | ___________________ | F-823-i22 Rev.5 | ___________________ | ___________________ | ___________________ |

# B. Epidural Set Assembly — 28 Source-Row Characteristics (F-823-i22 Rev.5)

At each planned check, record time/shift, actual observation (or measured value), result P/F/Pending/N/A with rationale, source component lot and applicable approved evidence reference. Criteria below are faithful historical form wording and require approved drawing/BOM confirmation. Do NOT assume pre-approval.

| Row | Description of Components | Original historical requirement | Approved spec / part / revision | Actual observation / value | Component lot | Check time / sample ID | Result | QC initials / Evidence / NCR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Touhy needle status | Sharp tip, without tip bend. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 2 | Epidural catheter status | Accurate in length, graduation clearly printed. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 3 | 5 ml Syringe (Loss Resistance) | Smooth movement, clear graduations. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 4 | Luer lock flat filter (0.2µ) | Proper fitting. No leakage. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 5 | Gauze swab white 50×50 mm | Clean; size 50×50 mm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 6 | Gauze swab white 75×75 mm | Clean; size 75×75 mm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 7 | Hypodermic needle (white hub) | 16×40 mm, sharp tip, without bend. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 8 | Hypodermic needle (green hub) | 21×40 mm, sharp tip, without bend. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 9 | Hypodermic needle (orange hub) | 25×16 mm, sharp tip, without bend. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 10 | Straw filter 5µ (Introducer status) | Proper fitting, sharp tip, without bend. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 11 | Applicator sponge | Clean sponge with handle. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 12 | 3 ml Syringe Luer Lock | Smooth movement, clear graduations. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 13 | 5 ml Syringe Luer Lock | Smooth movement, clear graduations. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 14 | Epidural catheter clamp | Strong clamping, no loose fitting. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 15 | Catheter connector | Good fitting, accurate assembly. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 16 | Catheter guide | Good quality, secured catheter. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 17 | Non-woven fenestrated drape | Per specification, dimensions. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 18 | Paper fenestrated drape | Per specification, dimensions. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 19 | Paper drape | Per specification, dimensions. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 20 | Povidone iodine swab stick | Check expiration 3 years [historical wording; unverified]. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 21 | Self-adhesive bandage | Good quality; 2×7 cm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 22 | Self-adhesive bandage | Good quality; 6×7 cm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 23 | Non-woven adhesive bandage | Strong; 6×30 cm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 24 | Tray (BU-Hips Tray-1) | Clean without damage [WI cites Tray-2]. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 25 | Hospital wrapping sheet | Size 75×75 cm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 26 | Ready pouch | Size 50×30 cm. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 27 | Components arrangement | Arrange completely. | _________________ | ________ | ________ | ________ | ______ | _______ |
| 28 | Label Printing on Pouch | Clear printed. | _________________ | ________ | ________ | ________ | ______ | _______ |

# C. Assembly check event / deviations / review

| Time / shift | Trigger / sample size | Production lot / checked units | Failed row / NCR | Affected quantity held | Recheck / disposition / reviewer |
| --- | --- | --- | --- | --- | --- |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |

Checked by QC: _____________________  Date/time: ______________    QC Supervisor/QCM review: _____________________  Date/time: ______________

# D. Manual Sealing Machine — F-823-i4-2 Rev.6 Controls

HISTORICAL VALUES IN EXISTING FORM, NOT VERIFIED APPROVED DD-05 VALIDATION LIMITS: width 5.0–12.0 mm; temperature 200–370 °C; sealing time 2.0–10 s. Approval required before these are operational acceptance windows. Machine operating limits must be read from approved validation/SOP.

| Product/SKU | Kit lot | Sealer ID & calibration | Pouch/material lot | Validated recipe ID | Sealing WI / spec rev |
| --- | --- | --- | --- | --- | --- |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |

| Timestamp / shift | Trigger / sample ID | Required width (mm) from approved spec | Actual width (mm) | Required temp (°C) from approved spec | Actual temp (°C) | Required dwell (s) from approved spec | Actual dwell (s) | Seal/peel integrity result | Reviewer / NCR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |

# E. Packaging / Final In-Process Characteristics (16 Original Rows)

A generic source form contains some product-irrelevant checks; assess and justify N/A. Fixed CE code or pipe/tube flow tests shall not be deemed applicable without an approved kit product/label specification. All results must be real observations or linked records.

| Source row | Characteristic | Original requirement (historic; subject to spec) | Applicable? Y/N + reason | Actual observation / value | Sample/time | Pass/Fail/Pending | Approved evidence/NCR |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Physical appearance | Clear, free from marks and dust. | _________ | _________ | ________ | _________ | _________ |
| 2 | Cutting status | Perfect, without fibers or sharp edge. | _________ | _________ | ________ | _________ | _________ |
| 3 | Plastic forming status | Good forming; no irregularities. | _________ | _________ | ________ | _________ | _________ |
| 4 | Peel test / sealing status | Good sealing; no fibers or sealed gap. | _________ | _________ | ________ | _________ | _________ |
| 5 | Gluing status of products | Not pulled out/detached. | _________ | _________ | ________ | _________ | _________ |
| 6 | Label printing status | Name, ref, lot, production, expiry, symbols per label. | _________ | _________ | ________ | _________ | _________ |
| 7 | Free from irregularities | No black, flash, deformation, short shot. | _________ | _________ | ________ | _________ | _________ |
| 8 | Filter air vent status | Intact and undamaged. | _________ | _________ | ________ | _________ | _________ |
| 9 | Connectors MLL/FLL status | Good fitting, without leak/crack. | _________ | _________ | ________ | _________ | _________ |
| 10 | Adapter/connection status | Good connection, not loose/damaged. | _________ | _________ | ________ | _________ | _________ |
| 11 | Components status | Complete components. | _________ | _________ | ________ | _________ | _________ |
| 12 | Tube dimension ID×OD×TL | Required values blank; controlled drawing needed. | _________ | _________ | ________ | _________ | _________ |
| 13 | Flow rate | Flow rate blank, ml/min; applicability required. | _________ | _________ | ________ | _________ | _________ |
| 14 | CE certified | Historic “CE-1639 printed, only CE printed”; per approved artwork only. | _________ | _________ | ________ | _________ | _________ |
| 15 | Clamp status | Good clamping without leak. | _________ | _________ | ________ | _________ | _________ |
| 16 | Graduation scale on products | Legible marking. | _________ | _________ | ________ | _________ | _________ |

# F. In-Process Evidence Linkage (not mandatory routine design re-test)

| Component / verification group | Design / supplier qualified evidence ID | Approved incoming report ID / lot | Scope / status / verifier |
| --- | --- | --- | --- |
| 3 ml / 5 ml syringe ISO 7886-1 | ________________________ | ________________________ | _______________________ |
| 5 ml LOR syringe special qualification | ________________________ | ________________________ | _______________________ |
| Clinical connector route & relevant ISO | ________________________ | ________________________ | _______________________ |
| Relevant Luer type qualification ISO 80369-7 | ________________________ | ________________________ | _______________________ |
| Approved BOM and labeling artwork | ________________________ | ________________________ | _______________________ |
| Pouch sealing validated window + seal evidence | ________________________ | ________________________ | _______________________ |

# G. Failure, Containment and Follow-on Release Link

| NCR / hold ID | Last accepted check / first affected unit | Affected kit lot / quantity | Disposition approval / recheck evidence | QC reviewed by / time |
| --- | --- | --- | --- | --- |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |
| ___________________ | ___________________ | ___________________ | ___________________ | ___________________ |

Assembly QC Sign/Date: _______________________________    Packaging QC Sign/Date: _______________________________    Supervisor/QCM Review Sign/Date: _______________________________

Sterilization load / transfer record: __________________   F-823-F20 final control report: __________________   Final release authorization: NOT PROVIDED / OUTSIDE THIS FORM.

# H. Reviewer decision and readiness to transfer to sterilization (not batch release)

| Required in-process record | Record reference / reviewer | Disposition / comment |
| --- | --- | --- |
| Correct approved BOM/configuration | ________________________________ | Pass / Fail / Pending |
| Components traceable to accepted source lots | ________________________________ | Pass / Fail / Pending |
| Assembly inspection results reviewed | ________________________________ | Pass / Fail / Pending |
| Sealer and sterile barrier check reviewed | ________________________________ | Pass / Fail / Pending |
| Artwork, lot and date labeling reviewed | ________________________________ | Pass / Fail / Pending |
| Open NCRs contained and dispositioned | ________________________________ | Pass / Fail / Pending |
| Production quantity reconciled | ________________________________ | Pass / Fail / Pending |

# I. Gap ID linkage — in-process versus external qualification

This traceability index ensures this form does not silently omit requirements: design and supplier qualification is documented OUTSIDE routine in-process checks. Evidence IDs in section F shall point to those approved records.

| Gap ID | Priority | Corrective-control topic | Form section / separate approved evidence route |
| --- | --- | --- | --- |
| GAP-001 | P1 | Clinical connector-route applicability | F / clinical route or external connector qualification |
| GAP-002 | P2 | WI cross-reference and document revisions | A–C / BOM, source-lot and actual timed assembly records |
| GAP-003 | P1 | Conflicting component configuration / BOM | A–C / BOM, source-lot and actual timed assembly records |
| GAP-004 | P2 | Incoming component acceptance and sampling basis | A–C / BOM, source-lot and actual timed assembly records |
| GAP-005 | P1 | Assembly lot and component traceability | A–C / BOM, source-lot and actual timed assembly records |
| GAP-006 | P2 | Syringe internal surfaces / particulate checks | B + F / syringe incoming or design qualification |
| GAP-007 | P1 | Syringe delivered-volume / graduation accuracy | B + F / syringe incoming or design qualification |
| GAP-008 | P2 | Syringe scales, fiducial alignment and legibility | B + F / syringe incoming or design qualification |
| GAP-009 | P1 | Syringe nozzle configuration and correct connection | F / clinical route or external connector qualification |
| GAP-010 | P2 | Syringe dead-space design verification | B + F / syringe incoming or design qualification |
| GAP-011 | P1 | Syringe piston air and water tightness | B + F / syringe incoming or design qualification |
| GAP-012 | P1 | Syringe extractables and chemical performance | B + F / syringe incoming or design qualification |
| GAP-013 | P2 | Syringe lubricant limits and process evidence | B + F / syringe incoming or design qualification |
| GAP-014 | P2 | Piston retention, fit, flanges and dimensional safety | B + F / syringe incoming or design qualification |
| GAP-015 | P2 | Syringe-specific package / information applicability | D–E / sealer, packaging and approved artwork |
| GAP-016 | P1 | Intravascular/hypodermic connector geometry and variant | F / clinical route or external connector qualification |
| GAP-017 | P2 | Luer material modulus and type-test control | F / clinical route or external connector qualification |
| GAP-018 | P1 | Luer fluid leakage evidence | F / clinical route or external connector qualification |
| GAP-019 | P1 | Luer sub-atmospheric air leakage | F / clinical route or external connector qualification |
| GAP-020 | P1 | Luer stress cracking after loading | F / clinical route or external connector qualification |
| GAP-021 | P1 | Luer mechanical connection integrity | F / clinical route or external connector qualification |
| GAP-022 | P2 | Non-measurable in-process component criteria | A–C / BOM, source-lot and actual timed assembly records |
| GAP-023 | P1 | Manual pouch sealing parameter approval and evidence | D–E / sealer, packaging and approved artwork |
| GAP-024 | P2 | Packaging and label requirements ambiguity | D–E / sealer, packaging and approved artwork |
| GAP-025 | P2 | Povidone iodine shelf-life reference | A–C / BOM, source-lot and actual timed assembly records |
| GAP-026 | P1 | Nonconforming kit handling and result dispositions | G–H / containment or final-release evidence outside this form |
| GAP-027 | P1 | Sterilization and final release evidence linkage | G–H / containment or final-release evidence outside this form |
| GAP-028 | P2 | Missing referenced source inspection / test records | F / source record not supplied; link only after verification |
| GAP-029 | P1 | Loss-of-resistance syringe applicability and clinical function | B + F / syringe incoming or design qualification |
| GAP-030 | P2 | Sampling / timing and cross-shift control methodology | A–C / BOM, source-lot and actual timed assembly records |