# EPIDURAL SET | REVISED WORK INSTRUCTION

RECONSTRUCTED WI • PRESERVES ORIGINAL §§1–13  |  WI-DD-05 • PROPOSED Rev.02

Epidural Set — DD-05  |  Source WI-DD-05 Rev.01  |  Evaluation date: 22 Oct 2026

DOCUMENT STATUS: PROPOSED CONTROLLED DRAFT — NOT IMPLEMENTED / NOT RELEASED

Scope: Epidural Set in-process assembly and manual packaging controls. This document does not constitute final product inspection, final release, or evidence that product tests have passed.

# 1. Purpose

To define controlled receiving, component verification, issuance, assembly, packaging, in-process checking, sterilization hand-off and documentation for Epidural Set (DD-05). This draft builds on WI-DD-05 Rev.01 and requires approval of identified OPEN inputs before use. [GAP-002,003,022]

# 2. Scope

Applicable to DD-05 epidural kit configurations specifically approved in the controlled BOM and production order. Covers in-process controls; finished-product release remains governed by separately approved procedures and documented test records. ISO 7886-1 applies only to in-scope hypodermic syringes; ISO 80369-7 only to in-scope hypodermic/intravascular Luer connectors. [GAP-001,027]

# 3. Definitions

• In-process control: documented inspections during kit component assembly and manual pouch packaging; not final acceptance or batch release.

• Approved BOM: current authorized configuration of part numbers, versions, sizes and quantities for each saleable kit SKU.

• Qualification/type test: controlled design evidence that shall not be silently duplicated as a per-lot requirement.

• NCR: nonconforming product record including affected scope and authorized disposition.

• OPEN TECHNICAL INPUT REQUIRED: no acceptance or applicability basis was provided and approval is necessary.

# 4. Responsibilities

4.1 Production controller: issue approved production order and parts/lot traceability, reconcile quantities, enforce identification/line clearance.

4.2 Assembly and packaging operators: assemble to approved BOM, record actual checks, stop work/segregate failures and notify QC.

4.3 QC: verify incoming-approved status; execute documented in-process inspection and sampling; record data and NCR references.

4.4 QC Supervisor/QCM: review in-process records and deviations; final batch authorization only according to controlled release procedure.

4.5 QA/Engineering/Supplier Quality: approve product specifications, connector application matrix, design/type evidence, controlled machine settings and validations.

4.6 Document Control: verify WI/forms/approved artwork revisions and release effective versions.

# 5. Procedure: Receiving and component verification

## 5.1 Receiving inspection of bulk components from supplier

5.1.1 Confirm purchase order, authorized vendor, item description/part number/revision, supplier lot, quantity and pack integrity. 5.1.2 Verify product configuration against current DD-05-PS and DD-05-PDBM (NOT SUPPLIED). 5.1.3 Use an approved, item-specific incoming acceptance/sampling plan and record actual sample quantities and acceptance decisions. 5.1.4 Do not substitute a visually similar component or mix different unapproved grades. 5.1.5 A component may enter the line only with documented incoming approval and a traceable issue-slip. [GAP-003–005]

## 5.2 Epidural catheter set bulk components — approved configuration

The original WI Rev.01 includes Tuohy G19×80 mm, LOR syringe 5 ml, epidural catheter and related fittings, 0.2µ filter, gauze, syringes 3/5 ml, needles G21/G25, drapes, tray “Hips Tray-2”, etc. The current Rev.5 assembly form includes an additional 16×40 mm needle, straw filter 5µ and “BU-Hips Tray-1”. This is NOT resolved by this draft. The correct contents and quantities for each kit must be loaded exclusively from approved SKU-specific drawings and BOM, reviewed by QA/Engineering. [GAP-003]

## 5.2.1 Incoming component-specific visual and functional requirements

Tuohy needle/stylet: examine tip curvature/sharpness, no bending or obstruction, surface finish, visible graduation, controlled dimensional acceptance and attachment integrity. Catheter and guide: verify accurate type/gauge/length to approved drawing, visible legible graduations, movement, no cracks/deformation. LOR syringe: verify approved 5 ml catheter-tip type, clear barrel/graduations, gasket/fit, absence of visible foreign matter and approved LOR functional specification. 3/5 ml syringes: inspect packaging/status, clear graduations, smooth motion and visual damage; do not treat those as evidence of ISO syringe accuracy, leakage or chemical tests. Hypodermic needle: verify approved actual gauge/length, tip and hub integrity, no blockage. [GAP-003,006–014,029]

Filter, catheter connector and catheter clamp: inspect correct part/route, integrity and approved mating/assembly specification. Gauze (50×50 and 75×75 mm), sponge, swab sticks, dressings, drapes, hospital wrap and trays: verify against approved part-specific drawings/visual acceptance guides, dimensions, contamination status, quantity and remaining expiry. Historic “3-year swab expiry” is not an approved substitute for supplier/controlled shelf-life evidence. [GAP-001,022,025]

## 5.2.2 Connector route and clinical safety assessment

Engineering/QA shall approve and maintain a component-level intended clinical route matrix. Clearly separate intravascular/hypodermic Luer connections in ISO 80369-7 from any neuraxial epidural interface. Obtain the applicable neuraxial connector standard and related test evidence before declaring conformity. Tests and dimensional tables apply only after the actual connector type and route are verified. Use a change-control gate before approving substitutions. OPEN TECHNICAL INPUT REQUIRED: route, device mating configuration and detailed 80369-6 source. [GAP-001,009,016]

## 5.2.3 Design/type verification for syringes and in-scope Luer connectors

For an in-scope 3 ml/5 ml manual syringe, supplier or manufacturer approved records must address ISO 7886-1 5–13, with test-method/calibration and validity: particle cleanliness (6.1), extracted pH/metals (6.2–6.3, Annex A), lubricant controls (7), graduated-volume accuracy and dimensions (8/9 and Table 1), flanges/stopper (10/11), nozzle (12), dead space (13.1/Annex C), air and liquid stopper leakage (13.2/Annex B/D), gravity-induced piston movement (13.4). The force-test in Annex E is informative/recommended and must be included only if adopted in approved specifications. ISO 7886-1 does NOT by itself define kit lot release. [GAP-006–014]

For a clinically appropriate intravascular/hypodermic Luer connector, type verification shall document ISO 80369-7:2016 cl.4 material modulus/type testing, cl.5 applicable Annex B variant dimensions, cl.6.1 fluid leakage (one allowable method), cl.6.2 air-leak under vacuum, cl.6.3 post-stress crack performance, and cl.6.4–6.6 forces/torques as applicable to slip/lock geometry. The actual test method is ISO 80369-20:2015; obtain this standard and the applicable reference connectors. Do not use 80369-7 values for a different clinical route. [GAP-016–021]

## 5.2.4 In-process acceptance criteria and evidence linking

For every component characteristic in F-823-i22, the approved specification/drawing and revision shall be identified. The checker records actual observation/measurement, result, sample ID, time, kit and component lot, initials and deviation as applicable. If the specification is missing, identify OPEN TECHNICAL INPUT REQUIRED, prevent an unjustified Pass and request QA/Engineering disposition. In-process observations may use approved incoming qualification records when the characteristic is not suitable for routine retest. [GAP-004,022,030]

# 6. Issuance and receipt of bulk components for assembly

6.1 Issue: Production controller confirms production order, exact kit SKU/BOM revision and required quantities; record issue-slip, lot and QC-accepted status for each part. 6.2 Receipt: line personnel reconcile issued actual to approved kit BOM and check identification/physical integrity. 6.3 Traceability: link every issued component lot and qualified evidence source to the in-process kit lot. Any mismatch triggers hold/NCR. [GAP-003,005]

# 7. Assembly of Epidural Set

7.1 Confirm approved order, line clearance, cleanliness, location/equipment and kit components; segregate parts in labeled trays. 7.2 Assemble each component into its approved tray position and quantity; do not infer the required contents from conflicting WI and form lists. 7.3 QC executes checks using approved sampling/shift/control plan and revises frequency only through approval. 7.4 Check identity, lot approval, specific component condition, arrangement, completeness and accurate approved pouch label. 7.5 Enter actual results to F-823-i22, with sample/time/component lot and source evidence. 7.6 Hold failed or pending work, record affected output since last acceptable check and open NCR before rework/restart. [GAP-003,005,022,026,030]

# 8. Packaging (manual pouch and cartons)

8.1 Following authorized in-process checks, cover tray with specified hospital wrap and insert in the approved pouch material/size according to the current BOM and process instructions. 8.2 Before and during manual-machine sealing, confirm machine/calibration, approved process validation and recipe, record actual sealing width, temperature, dwell time, approved seal-integrity or peel result and sample/time. Existing F-823-i4-2 values (width 5.0–12.0 mm, temperature 200–370°C, dwell 2.0–10 s) are historical form values ONLY: they must not be made effective for DD-05 until engineering/QA verify the approved validation/specification. 8.3 Verify packing compatibility and intended sterile barrier. 8.4 Compare product, lot, manufacture/expiry dates, symbols, description and any CE/registration identification with current approved artwork; no universal hardcoded CE code. 8.5 Stop, contain and NCR nonconforming pouch, seal, label or carton. 8.6 Pack to order quantity and transfer segregated, in-process-approved units to sterilization with records. [GAP-015,023,024,026,030]

# 9. Sterilization

Finished kits shall be transferred under controlled lot/quantity identity and sterilized in accordance with WI-7-5-1-1 and applicable validated process (source not supplied). Link kit lot, pouch material, load, cycle record and lab/sterilization evidence. In-process check completion does not mean sterilization has passed. [GAP-027]

# 10. Final Control and Release

Retain the existing intent of QC check after sterilization, QA/QCM review of cycle parameters and inspection reports, laboratory results and signed final authorization. Finished product release must follow current controlled procedure and F-823-F20 (not supplied) with segregation until review completed. The in-process forms in this draft do not grant finished-good release. Production records actual quantities transferred to FG only following documented authorization. [GAP-027]

# 11. Documentation and Records

Original referenced records: F-823-T017 (Tuohy needle); F-823-T90 (general products); F-823-T37 (hypodermic needle); F-823-T61 (syringes); F-823-i22 (kit in-process); F-823-i4-2 (manual sealing in-process); F-823-F20 (final kit). Existing product form header additionally refers to WI-DD-05-1, DD-05-PS and DD-05-PDBM; reconcile approved cross-references before release. Preserve completed batch/lot source records, change control, deviations, test reports and approvals. [GAP-002,028]

# 12. References

Verified in this assessment: ISO 7886-1:2017(E), corrected 2019-08; ISO 80369-7:2016(E); WI-DD-05 Rev.01; F-823-i22 Rev.5; F-823-i4-2 Rev.6. Existing WI also lists BS 6196, ISO 7864, ISO 9626, ISO 80369-20, ISO 22610 and ISO 14079: these sources were NOT supplied or assessed. Neuraxial connector assessment requires an approved ISO 80369-6 source. Confirm all currently effective editions using the approved controlled standards register before WI approval.

# 13. Modification / Documentation Change History

| Revision | Date | Change | Release status |
| --- | --- | --- | --- |
| 00 | 18.05.2021 | Original issue per WI history | Historic |
| 01 | 27.02.2023 | Final control updated per WI history | Source |
| 02 (PROPOSED) | 22.10.2026 | Documentary reconstruction; component/ISO mapping; in-process traceability and validation link | DRAFT — PENDING APPROVAL |

# Appendix A — Original In-Process Component List Reconciliation

Historic F-823-i22 Rev.5 row numbering is preserved below; the listed criteria must be crosschecked with approved drawings and BOM. This table is NOT an approved product master specification.

| Row | Original form component | Original requirement (condensed) |
| --- | --- | --- |
| 1 | Touhy needle status | Sharp tip, without tip bend. |
| 2 | Epidural catheter status | Accurate in length, graduation clearly printed. |
| 3 | 5 ml Syringe (Loss Resistance) | Smooth movement, clear graduations. |
| 4 | Luer lock flat filter (0.2µ) | Proper fitting. No leakage. |
| 5 | Gauze swab white 50×50 mm | Clean; size 50×50 mm. |
| 6 | Gauze swab white 75×75 mm | Clean; size 75×75 mm. |
| 7 | Hypodermic needle (white hub) | 16×40 mm, sharp tip, without bend. |
| 8 | Hypodermic needle (green hub) | 21×40 mm, sharp tip, without bend. |
| 9 | Hypodermic needle (orange hub) | 25×16 mm, sharp tip, without bend. |
| 10 | Straw filter 5µ (Introducer status) | Proper fitting, sharp tip, without bend. |
| 11 | Applicator sponge | Clean sponge with handle. |
| 12 | 3 ml Syringe Luer Lock | Smooth movement, clear graduations. |
| 13 | 5 ml Syringe Luer Lock | Smooth movement, clear graduations. |
| 14 | Epidural catheter clamp | Strong clamping, no loose fitting. |
| 15 | Catheter connector | Good fitting, accurate assembly. |
| 16 | Catheter guide | Good quality, secured catheter. |
| 17 | Non-woven fenestrated drape | Per specification, dimensions. |
| 18 | Paper fenestrated drape | Per specification, dimensions. |
| 19 | Paper drape | Per specification, dimensions. |
| 20 | Povidone iodine swab stick | Check expiration 3 years [historical wording; unverified]. |
| 21 | Self-adhesive bandage | Good quality; 2×7 cm. |
| 22 | Self-adhesive bandage | Good quality; 6×7 cm. |
| 23 | Non-woven adhesive bandage | Strong; 6×30 cm. |
| 24 | Tray (BU-Hips Tray-1) | Clean without damage [WI cites Tray-2]. |
| 25 | Hospital wrapping sheet | Size 75×75 cm. |
| 26 | Ready pouch | Size 50×30 cm. |
| 27 | Components arrangement | Arrange completely. |
| 28 | Label Printing on Pouch | Clear printed. |

# Appendix B — Verification Routes and Holding Points

| Stage | Documented control | Action if missing/failed |
| --- | --- | --- |
| Receiving | Accepted lot + incoming F-823 source + specification | Hold at incoming |
| Kit assembly | Correct BOM, linked lots, actual timed checks and QC | Stop line / NCR / segregate |
| Pouch sealing | Validated recipe, actual width/temp/time, seal check | Stop/segregate affected interval |
| Type / Supplier | Syringe/connector qualification evidence | Block unqualified substitution; engineering/QA decision |
| Sterilization / final | Executed cycle, labs, F-823-F20, authorized release | No FG release |

# Appendix C — Verified Technical Criteria for Conditional Component Qualification

The source-derived criteria below are required only for the component/design for which Engineering/QA confirms the standard applies. They are not routine inspection or acceptance limits for the whole Epidural Set. Record the test with approved equipment, actual values, sample ID, report and scope; link those reports from the in-process form. Never mark a product or kit compliant solely because a test row appears below. [GAP-007–021]

| Source / clause | Applies to | Technical acceptance and verification conditions (source-derived) |
| --- | --- | --- |
| ISO 7886-1 §6.2–6.3; Annex A | In-scope sterile hypodermic syringe | Extract at least 3 syringes: distilled water at 37–40 °C for 8 h to 8 h 15 min; pH shift ≤1.0 unit vs control; combined rise in Pb/Sn/Zn/Fe ≤5 mg/kg and Cd <0.1 mg/kg. Record extraction and lab analysis. |
| ISO 7886-1 §7 | Syringe lubricant | Surface-applied lubricant ≤0.25 mg/cm²; incorporated lubricant ≤0.6% by mass; applied lubricant shall not enter Luer channel when plunger fully inserted. Source Annex F method informative. |
| ISO 7886-1 §8, Table 1 | 3 ml syringe: 2≤V<5 ml | Below half nominal capacity: ±(1.5% of V + 2% of expelled volume). At/above half: ±5% of expelled volume. Record tested graduation, target ml, measured ml and absolute error. |
| ISO 7886-1 §8, Table 1 | 5 ml syringe: 5≤V<10 ml | Below half nominal capacity: ±(1.5% of V + 1% of expelled volume). At/above half: ±4% of expelled volume. Record target, measurement and error. |
| ISO 7886-1 §9 Table 1 | 3 ml / 5 ml scale design | 3 ml: maximum interval 0.2 ml, scale numbering every 1 ml and overall scale length ≥27 mm; 5 ml: interval 0.5 ml, numbering every 1 ml and scale length ≥36 mm. Clauses 9.1–9.4 also govern scale marking, arrangement and zero alignment. |
| ISO 7886-1 §11–12 | Relevant nozzle/plunger | Minimum plunger projection ≥8 mm at zero. 3 ml nozzle central; 5 ml nozzle central or eccentrically positioned per §12.2; nozzle lumen ≥1.2 mm. Nozzle Luer shall comply with applicable ISO 80369-7 where clinically appropriate. |
| ISO 7886-1 §13.1, Table 1, Annex C | 3 ml / 5 ml syringe dead space | Maximum dead space: 3 ml ≤0.070 ml; 5 ml ≤0.075 ml. Annex C uses a balance with resolution 1 mg or better and residual water mass after expulsion. |
| ISO 7886-1 §13.2, Annex B | Syringe air leakage | Draw at least 25% nominal capacity of water at 18–28°C; reach 88 kPa below ambient, isolate and observe 60 s. No leakage past plunger/stopper and no fall in manometer reading; no stopper separation. |
| ISO 7886-1 §13.2, Annex D, Table 1 | 3 ml / 5 ml syringe liquid leakage | Water leakage under compression: Table 1 axial gauge pressure 300 kPa (±5%) for both sizes, sideways force 1.0 N (±5%) for 3 ml, 2.0 N (±5%) for 5 ml; apply pressure for 30–35 s; no leakage past stopper except allowable small droplets between seals. |
| ISO 80369-7 §4.2–4.3 | Clinically applicable IV/hypodermic Luer | Connector nominal tensile/flexural modulus >700 MPa by ASTM D638-14 or D790-15e2; test by TYPE TEST and identify component material revision. |
| ISO 80369-7 §5; Annex B | Approved L1/L2 Luer variant | Perform dimensions/tolerances for the specific chosen male/female slip/lock variant using matching Annex B Fig. B.1–B.8 and Tables B.1–B.8. OPEN TECHNICAL INPUT REQUIRED: actual connector variant and controlled drawing. |
| ISO 80369-7 §6.1 | Applicable IV/hypodermic Luer | Choose one design type-test: (A) pressure-decay leak ≤0.005 Pa·m³/s at 300–330 kPa for 15–20 s, OR (B) no falling drop at 300–330 kPa liquid pressure over 30–35 s. ISO 80369-20 test method and Annex C mating reference connector required. |
| ISO 80369-7 §6.2–6.3 | Applicable IV/hypodermic Luer | Vacuum air leak ≤0.005 Pa·m³/s under 80–88 kPa sub-atmospheric pressure, 15–20 s. Stress cracking: after ISO 80369-20 Annex E stress exposure still meet §6.1 fluid leakage; keep post-stress test report. |
| ISO 80369-7 §6.4–6.6 | Slip / lock interface according to variant | Axial nonseparation 10–15 s: Luer slip at 23–25 N; Luer lock at 32–35 N. Luer lock no unscrewing at 0.0198–0.0200 N·m over 10–15 s; no overriding at 0.15–0.17 N·m over 5–10 s. Confirm exact variant and appropriate ISO 80369-20 test/report. |
| ISO 7886-1 §13.3, Annex E | Piston force | Force measurement recommended/informative, not an ISO mandatory numeric force threshold. Apply only if controlled syringe functional specification adopts a value; OPEN TECHNICAL INPUT REQUIRED for LOR performance. |

# Appendix D — Change-Control Traceability for Every GAP ID

This index links every proposed GAP ID to the relevant reconstructed WI section. It does not indicate acceptance or approval; use the Master Gap table for full corrective wording.

| Gap ID | Priority | Related WI section(s) | Change / remaining evidence |
| --- | --- | --- | --- |
| GAP-001 | P1 | 5.2.2 / 5.2.3 | Clinical connector-route applicability; Open — Technical Input |
| GAP-002 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | WI cross-reference and document revisions; Open — Approval |
| GAP-003 | P1 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Conflicting component configuration / BOM; Open — Technical Input |
| GAP-004 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Incoming component acceptance and sampling basis; Open — Technical Input |
| GAP-005 | P1 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Assembly lot and component traceability; Open — Evidence |
| GAP-006 | P2 | 5.2.1 / 5.2.3 / Appendix C | Syringe internal surfaces / particulate checks; Open — Evidence |
| GAP-007 | P1 | 5.2.1 / 5.2.3 / Appendix C | Syringe delivered-volume / graduation accuracy; Open — Evidence |
| GAP-008 | P2 | 5.2.1 / 5.2.3 / Appendix C | Syringe scales, fiducial alignment and legibility; Open — Evidence |
| GAP-009 | P1 | 5.2.2 / 5.2.3 | Syringe nozzle configuration and correct connection; Open — Technical Input |
| GAP-010 | P2 | 5.2.1 / 5.2.3 / Appendix C | Syringe dead-space design verification; Open — Evidence |
| GAP-011 | P1 | 5.2.1 / 5.2.3 / Appendix C | Syringe piston air and water tightness; Open — Evidence |
| GAP-012 | P1 | 5.2.1 / 5.2.3 / Appendix C | Syringe extractables and chemical performance; Open — Evidence |
| GAP-013 | P2 | 5.2.1 / 5.2.3 / Appendix C | Syringe lubricant limits and process evidence; Open — Evidence |
| GAP-014 | P2 | 5.2.1 / 5.2.3 / Appendix C | Piston retention, fit, flanges and dimensional safety; Open — Evidence |
| GAP-015 | P2 | 5.2.1 / 5.2.3 / Appendix C | Syringe-specific package / information applicability; Open — Technical Input |
| GAP-016 | P1 | 5.2.2 / 5.2.3 | Intravascular/hypodermic connector geometry and variant; Open — Technical Input |
| GAP-017 | P2 | 5.2.2 / 5.2.3 | Luer material modulus and type-test control; Open — Evidence |
| GAP-018 | P1 | 5.2.2 / 5.2.3 | Luer fluid leakage evidence; Open — Evidence |
| GAP-019 | P1 | 5.2.2 / 5.2.3 | Luer sub-atmospheric air leakage; Open — Evidence |
| GAP-020 | P1 | 5.2.2 / 5.2.3 | Luer stress cracking after loading; Open — Evidence |
| GAP-021 | P1 | 5.2.2 / 5.2.3 | Luer mechanical connection integrity; Open — Evidence |
| GAP-022 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Non-measurable in-process component criteria; Open — Technical Input |
| GAP-023 | P1 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Manual pouch sealing parameter approval and evidence; Open — Technical Input |
| GAP-024 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Packaging and label requirements ambiguity; Open — Approval |
| GAP-025 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Povidone iodine shelf-life reference; Open — Technical Input |
| GAP-026 | P1 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Nonconforming kit handling and result dispositions; Open — Evidence |
| GAP-027 | P1 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Sterilization and final release evidence linkage; Open — Evidence |
| GAP-028 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Missing referenced source inspection / test records; Open — Evidence |
| GAP-029 | P1 | 5.2.1 / 5.2.3 / Appendix C | Loss-of-resistance syringe applicability and clinical function; Open — Technical Input |
| GAP-030 | P2 | 5.1 / 5.2 / 6 / 7 / 8 / 9 / 10 / 11 | Sampling / timing and cross-shift control methodology; Open — Technical Input |

# Approval block — must be completed on controlled release

| Prepared / function | Technically reviewed / function | QA / QCM approval |
| --- | --- | --- |
| Name: __________________ Date: ______ Signature: ______ | Name: __________________ Date: ______ Signature: ______ | Name: __________________ Date: ______ Signature: ______ |