QUALITY CONTROL  /  TECHNICAL ASSESSMENT

# ISO CLOSURE REGISTER

TF-11 Extension Tube — ISO-to-control mapping, evidence status and document reconstruction

| Product / programme | Extension Tube — TF-11 |
| --- | --- |
| Existing WI | WI-8-2-2-11, Rev.6 (15 Mar 2023) |
| In-process sources | F-823-i60 Rev.7; F-823-i15 Rev.16; F-823-i4 Rev.19 |
| Standards | ISO 8536-10:2015 (partial preview); 2004 full legacy copy; ISO 8536-9:2015 scope check |
| Evidence availability | BLANK controlled templates only; no executed batch/qualification evidence |
| Status | CONTROLLED DRAFT — NOT APPROVED / NOT RELEASED |
| Review date | 15 October 2026 |

Critical limitation: ISO 8536-10:2015 Annex A/B/C and the full ISO 8536-9:2015, approved TF-11-PS/PDBM, actual process-control plan and execution records were not provided. The 2004 annex methods are historical context, not current approved acceptance methods.

# 1. Rules for closure

Every entry is OPEN or conditional: no approved and executed test/production evidence was submitted. Proposed WI and form text are drafts and do not demonstrate closure. ISO 8536-10:2015 Annexes A/B/C and ISO 8536-9 full text remain unavailable, so documentary coverage percentages cannot be certified. “N/A” requires approved product/BOM justification.

# 2. ISO applicability and source-coverage register

| Requirement ID / source | Verified normative subject | Applicability / reviewed evidence | Gap linkage | Documentary coverage | Actual evidence / operational closure |
| --- | --- | --- | --- | --- | --- |
| 2015 cl.1 | Scope/accessory/line classification | Open — Applicability | GAP-001 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.2 | Normative references (full text availability) | Open — Source | GAP-001,005,017 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.3 | Materials | Partial — company incoming checks | GAP-002,012,015 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.1 | No trapped air bubbles in flow channels | Partial visual; technical method unverified | GAP-006 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.2 | Particles controlled / contamination index | Open — test evidence | GAP-006 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.3 | 15 N / 15 s tensile, conditional scope | Partial 15 N; hold/evidence absent | GAP-004,008 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.4 | No leakage under A.4 method | Partial IP screen; A.4 absent | GAP-003 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.5 | Conical adapter standard compliance | Open — connector evidence | GAP-005 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.6 | Protective caps via ISO 8536-4 | Partial appearance; method/evidence absent | GAP-005,017 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.7 | Stopcock manipulation | Conditional N/A (no integrated stopcock verified) | GAP-001 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.8 | Injection-site loss criteria | Conditional — variant unverified | GAP-007 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.4.9 | Check-valve test per ISO 8536-12 | Conditional N/A (not verified in BOM) | GAP-001 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.5 | Chemical requirements via ISO 8536-4 | Open — external method/evidence | GAP-017 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.6.1 | Validated sterilization | WI reference only; evidence absent | GAP-016 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.6.2 | Pyrogen freedom | Open — lab/qualification | GAP-017 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.6.3 | Haemolysis assessment | Open — lab/qualification | GAP-017 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.7 | Packaging via ISO 8536-4 | Partial IP packaging; references missing | GAP-013 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.8.1–8.3 | Labels and packaging-level content | Partial WI/form; artwork absent | GAP-014 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 cl.9 | Safe disposal guidance | Open — IFU applicability and evidence | GAP-014 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 Annex A–C | Normative detailed methods | Unable to verify — annexes unavailable | GAP-001,003,004,006,007,017 | Draft changes proposed; not approved | None submitted / OPEN |
| 2015 ISO 8536-9 | Possible fluid-line primary standard | Open — full edition & applicability | GAP-001 | Draft changes proposed; not approved | None submitted / OPEN |

# 3. Corrective gap closure register

| Gap ID / priority | Required documentary revision | Required verification evidence | Available evidence / status | Owner / next decision |
| --- | --- | --- | --- | --- |
| GAP-001 / P1 | Applicability and controlled ISO edition | approved configuration BOM / intended use, ISO 8536-9 and 8536-10 applicability matrix, controlled full ISO 2015. | None supplied. Open — Technical Input / Approval | QA / Engineering — approve source/criteria and obtain executed records. |
| GAP-002 / P2 | Controlled product specification not available | approved TF-11-PS, PDBM and drawing, current revision, verified reading records. | None supplied. Open — Technical Input | Engineering / Document Control — approve source/criteria and obtain executed records. |
| GAP-003 / P1 | Leakage screen is not established as ISO conformity test | full Annex A.4 (2015), approved in-process screen method, equipment calibration/qualification, completed screening and applicable normative evidence or approved equivalence. | None supplied. Open — Technical Input / Evidence | QC / Engineering / QA — approve source/criteria and obtain executed records. |
| GAP-004 / P1 | Tensile holding time omitted | approved tensile SOP / applicability, calibrated equipment, 15 N/15 s completed data, approved sampling route. | None supplied. Open — Evidence / Approval | QC / QA — approve source/criteria and obtain executed records. |
| GAP-005 / P2 | Luer adapter conformance evidence not linked | approved connector design specification, qualification report, mating master and component lot linkage. | None supplied. Open — Evidence / Technical Input | Engineering / Supplier Quality / QC — approve source/criteria and obtain executed records. |
| GAP-006 / P2 | Air-bubble and particulate control route unverified | approved contamination/air-bubble verification strategy, referenced test method & recorded results if applicable. | None supplied. Open — Evidence / Technical Input | QC / QA / Laboratory — approve source/criteria and obtain executed records. |
| GAP-007 / P2 | Optional injection-site controls not fully defined | BOM confirming injection site, approved Annex A.6 test record/method, actual results or documented N/A. | None supplied. Open — Technical Input / Evidence | Engineering / QC — approve source/criteria and obtain executed records. |
| GAP-008 / P1 | Cyclohexanone gluing process lacks approved operating limits | approved solvent spec, hazard controls, process qualification/validation where required, cure criteria, completed IP results, tensile/leak evidence. | None supplied. Open — Technical Input / Evidence | Production / Engineering / QA — approve source/criteria and obtain executed records. |
| GAP-009 / P2 | Dimension and tube-length measurements are not fully traceable | TF-11-PS with dimensional tolerances, calibrated instruments, completed readings. | None supplied. Open — Technical Input / Evidence | QC / Engineering — approve source/criteria and obtain executed records. |
| GAP-010 / P2 | Hourly WI frequency conflicts with two-hour grids | approved IP frequency/control plan and supporting hourly/approved-period records. | None supplied. Open — Approval / Evidence | QC / Production / QA — approve source/criteria and obtain executed records. |
| GAP-011 / P2 | Result, sample and instrument traceability weak | completed traceable IP worksheets, calibrated equipment and review signatures. | None supplied. Open — Evidence / Approval | QC / Document Control — approve source/criteria and obtain executed records. |
| GAP-012 / P2 | Molding/extrusion acceptance hierarchy is unclear | per-part approved specifications, justified limits and representative completed molding/extrusion records. | None supplied. Open — Technical Input / Evidence | Engineering / QC / Production — approve source/criteria and obtain executed records. |
| GAP-013 / P1 | Packaging seal process limits and integrity acceptance not reconciled | approved seal process qualification/validation, machine recipe matrix, completed monitoring/integrity test & calibration. | None supplied. Open — Technical Input / Evidence | Packaging Engineering / QA / QC — approve source/criteria and obtain executed records. |
| GAP-014 / P2 | Pack and label master / line clearance linkage incomplete | approved artwork, market and packaging BOM, signed line-clearance checks and label samples. | None supplied. Open — Evidence / Approval | QA / Packaging / Regulatory — approve source/criteria and obtain executed records. |
| GAP-015 / P2 | Incoming material identity and traceability need confirmed records | F-823-T1, material specifications, COAs where applicable, actual receiving releases and batch history. | None supplied. Open — Evidence | Supplier Quality / Warehouse / QC — approve source/criteria and obtain executed records. |
| GAP-016 / P1 | Sterilization validation and batch linkage not demonstrated | approved WI-7-5-1-1, validation, actual cycle printouts/load history, released lab tests. | None supplied. Open — Evidence | Sterilization / QA / QC Manager — approve source/criteria and obtain executed records. |
| GAP-017 / P2 | Chemical, pyrogen, and haemolysis evidence route unverified | approved ISO 8536-4 and test methods, relevant lab/qualification data, control-plan frequency, signed review. | None supplied. Open — Technical Input / Evidence | QA / Laboratory / Regulatory — approve source/criteria and obtain executed records. |
| GAP-018 / P1 | Release decision and cited final record not verifiable | approved F-823-F47, real signed batch-release records, lab and sterilization links, nonconformity closure. | None supplied. Open — Evidence | QC Manager / QA — approve source/criteria and obtain executed records. |
| GAP-019 / P3 | Sheet-tab revision mismatch on shared form | controlled master revision register and aligned issue after approval. | None supplied. Open — Approval | Document Control — approve source/criteria and obtain executed records. |
| GAP-020 / P2 | In-process nonconforming response and restart control underdefined | nonconformance/hold SOP, actual NCR/hold records, approved disposition and restart verification. | None supplied. Open — Approval / Evidence | QC / Production / QA — approve source/criteria and obtain executed records. |

# 4. Conditions before setting Verified Closed

Confirm each product configuration, included optional components, classification and current applicable standard including controlled edition.

Obtain and verify actual current ISO 8536-10:2015 Annex A/B/C, relevant ISO 8536-9 and ISO 8536-4, including requirements by printed ISO page and actual ISO test settings.

Approve WI and revised forms with controlled revision IDs, implemented training, sample plan, product specifications and calibrated instruments.

Review documented qualification/validation/supplier/laboratory controls wherever normative tests are outside routine IP.

For each applicable batch check objective inspection, sterilization, release, acceptance and nonconformance records; independent QC Manager/QA review per QMS.

Change status to Verified Closed only after approver, exact record ID/date and acceptance evidence have been checked; do not imply 100% product conformity from draft content.

# 5. Source register and traceability

WI Rev.6 pp.1–4 (uploaded legacy DOC); F-823-i15 Rev.16 rows 1–32; F-823-i60 Rev.7 rows 2–28; F-823-i4 Rev.19 rows 1–24.

ISO 8536-10:2004 supplied full PDF: withdrawn edition used solely for historical technical comparison; ISO printed pp.1–10.

ISO 8536-10:2015 official product page: https://www.iso.org/standard/63621.html

ISO 8536-10:2015 limited preview (clauses 1–9, Foreword; not Annex A/B/C): https://preview.sist.si/sist-preview/63621/dc8e407194264df5b138759b7d8876db/ISO-8536-10-2015.pdf

ISO 8536-9:2015 scope only: https://www.iso.org/cms/%20render/live/en/sites/isoorg/contents/data/standard/06/36/63620.html