QUALITY CONTROL  /  TECHNICAL ASSESSMENT

# REVISED IN-PROCESS CONTROL REPORT

Review-ready controlled form package — NOT an Inspection & Test / Final Release Report

| Product / programme | Extension Tube — TF-11 |
| --- | --- |
| Existing WI | WI-8-2-2-11, Rev.6 (15 Mar 2023) |
| In-process sources | F-823-i60 Rev.7; F-823-i15 Rev.16; F-823-i4 Rev.19 |
| Standards | ISO 8536-10:2015 (partial preview); 2004 full legacy copy; ISO 8536-9:2015 scope check |
| Evidence availability | BLANK controlled templates only; no executed batch/qualification evidence |
| Status | CONTROLLED DRAFT — NOT APPROVED / NOT RELEASED |
| Review date | 15 October 2026 |

Critical limitation: ISO 8536-10:2015 Annex A/B/C and the full ISO 8536-9:2015, approved TF-11-PS/PDBM, actual process-control plan and execution records were not provided. The 2004 annex methods are historical context, not current approved acceptance methods.

PROPOSED TEMPLATE ONLY | Replaces no currently issued F-823-i15, F-823-i60 or F-823-i4 until controlled approval | Supporting WI: WI-8-2-2-11 proposed review draft

# A. Lot / configuration and applicability header

| Mandatory information | Entry |
| --- | --- |
| Product / product code / TF-11 type | ____________________________________________ |
| Actual intended use and ISO scope matrix ID/rev | ____________________________________________ |
| Production lot / quantity / order | ____________________________________________ |
| Tube OD/ID/length / approved TF-11-PS revision | ____________________________________________ |
| PDBM revision / connector and injection-site configuration | ____________________________________________ |
| Raw resin/component lot IDs / incoming acceptance IDs | ____________________________________________ |
| Shift/date/machine/extruder/mould/cavity/operator | ____________________________________________ |
| Applicable approved IP sampling plan frequency / revision | ____________________________________________ |
| Instruments / calibration verification IDs | ____________________________________________ |
| QC checker / reviewer | ____________________________________________ |

Rule: use actual recorded times according to currently approved WI and control plan. The original WI requires hourly molding/extrusion checks; this draft offers expandable rows rather than defaulting to two-hour spacing. Blank means not performed, NOT Pass.

# B. Molding / extrusion stage (source F-823-i15 Rev.16)

| Actual date/time / shift | Part / machine / cavity / sample | Characteristic + approved spec limits | Actual value / defect code / unit | Result P/F/N/A/Pending | Corrective action / reference |
| --- | --- | --- | --- | --- | --- |
| ________ | ________ | Appearance / streaks / bubbles / black spots; per PS | ________ | ________ | ________ |
| ________ | ________ | OD / ID / wall thickness / length; PS limits | ________ | ________ | ________ |
| ________ | ________ | Flash / sink / gate / blockage / cracks; per PS | ________ | ________ | ________ |
| ________ | ________ | Weight/fit/mating only if specified by part PS | ________ | ________ | ________ |
| ________ | ________ | Restart / after machine setting change | ________ | ________ | ________ |

Approved component specification and tolerances (TF-____-PS Rev ___): ______  Material lot / colourant: ______  Instrument/tool ID: ______  QC stamp/transfer tag: ______

# C. Extension Tube assembly & functional IP (source F-823-i60 Rev.7)

| ID / source row | IP characteristic | Approved acceptance / conditional applicability | Actual recorded evidence / units | Sample / time / equipment | P/F/N/A/Pending |
| --- | --- | --- | --- | --- | --- |
| 1 / R10 | Appearance | Clear; no approved appearance defects; PS rev ___ | Observed defect/no defect ___ | ___ | ___ |
| 2 / R11 | End plug / cap | Fit, correct threads, no damage/looseness | Fit/defect ___; part lot ___ | ___ | ___ |
| 3 / R12 | Female luer lock | No flash/block/damage; approved mating check | Master/gauge or part ID ___; result ___ | ___ | ___ |
| 4 / R13 | Tube status | Clear, flexible, kink resistance; no dust/fish eye/holes per PS | Observed defect / kinking condition ___ | ___ | ___ |
| 5–6 / R14,17 | Tube ID / OD / length | PS ___ LSL/USL: ID ___ mm; OD ___ mm; length ___ cm | Actual ID ___ mm, OD ___ mm, length ___ cm | ___ | ___ |
| 7 / R18 | Bond / gluing | Solvent/bond qualified parameters rev ___; no separation/blockage | Solvent lot ___; cure rule ___; actual ___ | ___ | ___ |
| 8 / R19 | Injection site (if included) | Per approved BOM; no crack, leak/blockage; qualification link | Present Y/N ___; test report ID ___ | ___ | ___ |
| 9 / R20 | Male luer lock | Fit/threads good, no leak/damage; applicable qualification link | Mating master ___; record ___ | ___ | ___ |
| 10 / R21 | Male luer cap | Cap fit/no damage/looseness | Observed condition ___ | ___ | ___ |
| 11 / R22 | Clamp | Secure fit/closing; no crack/leak (when applicable) | Installed Y/N ___; result ___ | ___ | ___ |
| 12 / R23 | Routine IP leakage screen | Approved method ___; historical template 50 kPa / 15 s ONLY IF APPROVED | Actual kPa ___; time s ___; medium ___; leaks Y/N ___ | Sample ___ / fixture ___ | ___ |
| 13 / R24 | Tensile / bond integrity | 15 N for 15 s if ISO 8536-10 cl.4.3 applies & approved method | Actual N ___; seconds ___; separation Y/N ___ | Joint ___ / device ___ | ___ |
| 14 / R25 | Tube winding | Correct winding, no bent/occluded tube per PS | Observed condition ___ | ___ | ___ |

Separate qualification evidence reference (not routine lot test unless authorized): ISO method/rev ___; full current Annex A.4 ___; tensile method ___; connector/particles/bubble/injection report ___; configuration/lot scope ___; approved sampling plan ___; QA review ___.

# D. Packaging machine and final in-process check (source F-823-i4 Rev.19)

| Characteristic / source | Approved requirement (pending controlled recipe) | Actual / unit / sample | Pass/Fail/Pending |
| --- | --- | --- | --- |
| Sealing width / row 9 | Approved recipe ID ___; width LSL/USL ___ mm | ___ mm / sample ___ | ___ |
| Temperature / rows 10,12 | Machine/product recipe selected ___; limits ___ °C | ___ °C / machine ___ | ___ |
| Sealing time / row 11 | Recipe limits ___ s | ___ s | ___ |
| Appearance / forming / cutting rows 14–16 | No visible defects per approved packaging spec | Observation ___ | ___ |
| Seal/peel / row 17 | Controlled seal-integrity/peel method & acceptance ___ | Actual measurement / worksheet ___ | ___ |
| Artwork/printing / row 18 | Label master & market-approved artwork rev ___ | Label code/lot/expiry verified ___ | ___ |
| Product / components / rows 19–20 | Approved PS/PDBM; no damage/part omission | Count, defect codes ___ | ___ |
| Dimensions & length / rows 21–22 | Approved size & acceptance limits per PS | Actual ID/OD/length ___ | ___ |
| Graduation scale / row 23 | N/A unless product specification demands | N/A justification / actual ___ | ___ |

# E. Exceptions, process disposition and review

| Required trace | Entry |
| --- | --- |
| Deviation / NCR / hold ID; lot / quantity affected | ____________________________________________ |
| Corrective adjustment / machine / last valid checkpoint | ____________________________________________ |
| Post-adjustment restart sample and actual results | ____________________________________________ |
| QC in-process disposition (Transfer / Hold / Reject) | ____________________________________________ |
| QC checker / date-time and QC Supervisor review | ____________________________________________ |
| Independent normative verification report link (if applicable) | ____________________________________________ |
| Final-release record cross-reference only: F-823-F47 | ____________________________________________ |

Important: all missing accepted limits remain OPEN TECHNICAL INPUT REQUIRED. The in-process check is NOT the market-release decision; separate sterilization and final-control records govern that decision.

# F. Gap / in-process form traceability

The following controlled-draft section links ensure every master gap ID is traceable to a proposed WI or in-process record change; some verifications remain outside routine in-process control.

| Gap ID | Proposed location / evidence route |
| --- | --- |
| GAP-001 | Form A — configuration / ISO scope |
| GAP-002 | A,C — PS and BOM revision + limits |
| GAP-003 | C — screen and qualification evidence link |
| GAP-004 | C — actual 15 N / 15 s data |
| GAP-005 | C — mating fit and connector qualification |
| GAP-006 | B,C — visual / external verification references |
| GAP-007 | C — injection site conditional branch |
| GAP-008 | C — solvent / bond process trace |
| GAP-009 | B,C,D — actual dimensional data |
| GAP-010 | A,B,C — approved sampling times |
| GAP-011 | A–E — sample/tool/result traceability |
| GAP-012 | B — molding / extrusion checkpoints |
| GAP-013 | D — packaging setpoints / sealing |
| GAP-014 | D — artwork and labels |
| GAP-015 | A,B — incoming material ID |
| GAP-016 | E — sterilization separate link only |
| GAP-017 | C,E — qualification/lab separate link only |
| GAP-018 | E — separate final-release reference only |
| GAP-019 | A — form revision synchronization (Document Control) |
| GAP-020 | E — hold, NCR, restart verification |