QUALITY CONTROL  /  TECHNICAL ASSESSMENT

# REVISED WORK INSTRUCTION

Production & In-Process Control Procedure for Extension Tube — TF-11 | Proposed revision, source Rev.6

| Product / programme | Extension Tube — TF-11 |
| --- | --- |
| Existing WI | WI-8-2-2-11, Rev.6 (15 Mar 2023) |
| In-process sources | F-823-i60 Rev.7; F-823-i15 Rev.16; F-823-i4 Rev.19 |
| Standards | ISO 8536-10:2015 (partial preview); 2004 full legacy copy; ISO 8536-9:2015 scope check |
| Evidence availability | BLANK controlled templates only; no executed batch/qualification evidence |
| Status | CONTROLLED DRAFT — NOT APPROVED / NOT RELEASED |
| Review date | 15 October 2026 |

Critical limitation: ISO 8536-10:2015 Annex A/B/C and the full ISO 8536-9:2015, approved TF-11-PS/PDBM, actual process-control plan and execution records were not provided. The 2004 annex methods are historical context, not current approved acceptance methods.

Document no.: WI-8-2-2-11 | Revision: PROPOSED (no approved revision assigned) | Supersedes: none until approval | Status: REVIEW DRAFT | Change rationale: GAP-001–GAP-020

# 1. Purpose

To define raw-material verification, injection molding, tubing extrusion, assembly, in-process quality control, packaging, sterilization linkage and authorized product-release checks for TF-11 Extension Tubes. The controls prevent use or transfer of unverified or nonconforming product. [GAP-011,020]

# 2. Scope

This draft covers TF-11 sterile extension tube configurations produced by Saudi Mais Co. for Medical Products. It governs IN-PROCESS checks, not final inspection. Confirm whether each product configuration is a pressure-infusion connecting line, accessory, or other intended-use variant before applying standard-specific clauses. [GAP-001]

# 3. References and applicability

Existing controlled WI-8-2-2-11 Rev.6, 15.03.2023 (source baseline); this document is a DRAFT.

ISO 8536-10:2015, accessories for fluid lines with pressure equipment, applicable only where justified. Full authorized standard and annexes required before approval.

ISO 8536-9:2015, pressure-fluid lines, pending product configuration / applicability review; full text required.

TF-11-PS, TF-11-PDBM, product drawing and approved manufacturing/test control plan: OPEN TECHNICAL INPUT REQUIRED.

F-823-T1 (incoming), F-823-T90 (male Luer component), F-823-i15 (molding/extrusion), F-823-i60 (extension tube), F-823-i4 (packaging), F-823-F47 (final release): exact current issue/revisions per document control.

WI-7-5-1-1 (sterilization), related validation, laboratory and deviation / material hold SOPs: controlled copies not supplied.

Approved Standards Applicability Matrix must list product code, intended use, operating pressure, included stopcock/injection site/check valve/adapter, reference standards, risk rationale, method and evidence; hold approval where unanswered. [GAP-001,007]

# 4. Definitions

IP: in-process production control performed during manufacturing; not a final disposition test.

TF-11-PS: controlled specification containing approved limits and material/dimensions.

PDBM: approved product description/BOM identifying variants and components.

Routine screen: a controlled production check not automatically equivalent to normative ISO verification.

Qualification: performance demonstration using approved method/equipment and representative configuration; frequency per approved control plan.

Pending: acceptance or release cannot be determined from incomplete technical input/evidence.

# 5. Responsibilities

| Function | Required action |
| --- | --- |
| QC Controller | Perform required measurements/tests, time and sample ID logging, identification and holds; sign each IP control. |
| QC Supervisor | Review completeness, result validity, response to abnormalities and transfer authorizations. |
| Production/Operators | Correct setup, trace materials, follow approved parameters, segregate nonconforming work and support rechecks. |
| Maintenance/Engineering | Prepare mould/die/mandrel; establish approved machine/tool/solvent/packaging processes and valid technical limits. |
| QA/QCM | Approve technical changes, review normative applicability, qualification/lab/sterilization data; authorize final release under QMS. |

# 6. Incoming raw material — receipt and release

6.1 Verify the supplier receipt/PO, approved material specification and grade, material bag condition (no damage), appearance/colour/formation, legible lot and storage identification before use. Retain Rev.6 visual controls: condition of bags, material grade/description, granules appearance, granule formation (no lumps), lot marking and arrangement. Apply approved incoming inspection sampling plan, not an assumed universal AQL level. Record in F-823-T1 or controlled successor. [GAP-015]

6.2 Verify supplier/COA and receiving lot with approved TF-11-PS/PDBM and link each approved material lot to the subsequent production lot. If specification/grade/status is unavailable, place material on hold. [GAP-002,015]

# 7. Injection-molding control

7.1 Production shall notify Maintenance to prepare the correct mould and issue specified materials for female connector, male connector, slide clamp, male and female luer-lock caps. The operator shall identify machine/mould/cavity and material lot, use the approved loading/feeding process, and establish authorized settings. [Source WI §5.2; GAP-012]

7.2 QC shall inspect material appearance and moulding attributes against the approved component PS: flash, deformation, cracks, holes/blockage, weight where applicable, fitting/gauge, dimensional and other part-specific attributes. Record results, actual measurements, time, lot and cavity on F-823-i15. A broad generic defect label or weight band is not evidence of an approved part-specific limit. [GAP-012]

7.3 Maintain the original coordination rule: upon product irregularity notify Production/Maintenance and implement documented containment, approved adjustment and restart check before stamping/transfer. [GAP-020]

# 8. Extrusion and tube control

8.1 Prepare the correct die/mandrel and raw material; clean the extruder core as per the approved machine SOP, heat to the validated/authorized extrusion parameters and feed resin as approved. The existing source requires core cleaning with a steel brush; use this only when compatible with current controlled machine SOP. [Source WI §5.3; GAP-012]

8.2 Inspect: clear transparent tubing, no visible contamination, black spots, fish eyes, holes, cracks, bubbles, irregularity, bumps or blocked lumen; smooth surface, adequate flexibility/kink status and controlled length. See approved TF-11-PS for quantitative limits. [Source WI §§5.3–5.4; F-823-i60 rows 10,13; GAP-006,009]

8.3 Measure OD, ID and tube length and record actual readings with units and calibrated instruments, sample/lot and acceptance against the approved limits. Nominal product variants from baseline WI are retained in §8.4; no tolerance is inferred. [GAP-002,009]

| Original TF-11 configuration | Nominal OD (mm) | Nominal ID (mm) | Tube lengths (cm) | Tolerance status |
| --- | --- | --- | --- | --- |
| Male to Male | 4.1 | 2.9 | 25 / 30 / 50 / 100 / 180 | OPEN — TF-11-PS |
| Female to Female | 4.1 | 2.9 | 25 / 30 / 50 / 100 / 180 | OPEN — TF-11-PS |
| Male to Female | 4.1 | 2.9 | 25 / 30 / 50 / 100 / 180 | OPEN — TF-11-PS |
| High Pressure (Male & Female) | 4.9 | 1.7 | 25 / 30 / 50 / 100 / 180 | OPEN — TF-11-PS |

8.4 Values above are source WI Rev.6 nominal figures, not acceptance limits. Exact type, code, pressure claim and approved tolerances are required for a valid Pass/Fail decision. [GAP-002,009]

# 9. Manual assembly and bonding

9.1 Assembly: prepare approved cyclohexanone material/grade and clean authorized dispensing equipment. Dip the designated tubing end into solvent and immediately seat within the corresponding approved male/female connector, controlling solvent excess and lumen occlusion. Fit connector caps, check secure engagement and place finished assembled items in a clean identified protective bag before packaging; preserve valid baseline process steps. [Source WI §6]

9.2 No bonding lot may proceed solely under the informal “strong gluing” wording. Record solvent lot/expiry where relevant, workstation/process settings, handling/assembly delay, required cure/conditioning time, process-control findings and acceptance as established by approved manufacturing validation/specification. OPEN TECHNICAL INPUT REQUIRED: approved solvent/bond process parameters. [GAP-008]

9.3 Check part fit, no damage, no loose cap/threads, no flash/blocked lumen, clamp closure/integrity, and applicable injection-site bond/appearance only for configurations identified by PDBM. Mark absent optional features N/A with BOM justification. [GAP-005,007]

# 10. In-process physical verification

10.1 Tensile: where ISO 8536-10:2015 cl.4.3 or the approved TF-11 test method applies, subject each defined joint/component to an axial static force of at least 15 N for 15 s (subject to final verification of current Annex A.3 method). Accept no pullout/separation or disqualifying damage. Record N, seconds, sample, lot, joint and calibrated device; link alternate qualification evidence if not performed in routine IP. [GAP-004]

10.2 Leakage: follow the approved ROUTINE in-process leak screen and record applied medium, actual positive pressure in kPa, dwell in seconds, observation, sample and equipment ID. The current F-823-i60 setting of 50 kPa/15 s is a SOURCE TEMPLATE VALUE only and requires technical authorization; it shall not be asserted to satisfy ISO 8536-10:2015 cl.4.4 unless the full 2015 Annex A.4 method or approved equivalence is verified. OPEN TECHNICAL INPUT REQUIRED. [GAP-003]

10.3 Accessory-specific evidence, if applicable: flow-path bubble/particle testing, conical Luer connector qualification and injection-site leak/drop performance shall follow controlled normative/qualification methods and have linked evidence. ISO 8536-10:2015 cl.4.8 specifies ≤10 drops per batch and ≤2 per unit for injection-site units under Annex A.6, but the method and product applicability are pending verification. Do not impose such qualification tests every production hour without approved plan. [GAP-005,006,007,017]

# 11. In-process sampling and record completion

11.1 Baseline WI §§5.2/5.3 require hourly injection and extrusion records. The supplied F-823-i15 / F-823-i60 sheets contain two-hour time boxes. Until a different frequency is approved, fulfill current WI frequency and log actual hourly checks on controlled continuation sheets. For any change obtain a signed characteristic-based control plan, documented reason and synchronized form/WI. Include startup, shift handover, adjustments and restarts per approved plan. [GAP-010]

11.2 The QC Controller shall record actual measurements, device IDs/calibration status, characteristic/limits from controlled PS, sample/lot, time, visual category, P/F/N/A/Pending and abnormality disposition on designated F-823-i15 / i60 worksheets. The QC Supervisor reviews results; approved QC-tag stamping authorizes transfer only to the next in-process stage, not market release. [GAP-011,018]

# 12. Packaging and label in-process checks

12.1 Prior to packaging, QC/Production shall verify issue/lot, appropriate machine and validated blister/packaging material and selected process recipe. Check film forming condition, cutting, seal width/temperature/time, visual sealed gap, peel/seal-integrity test, component count/condition and tube geometry per approved method. The generic F-823-i4 temperature bands must be mapped to the actual machine and product; do not assume interchangeability. [GAP-013]

12.2 Verify intact blister, legible correct product name/code, size, lot, product type, manufacturer details, sterilization/expiry claims, artwork/IFU and market-specific marks against the approved label master. Verify changeover/line clearance and inspect printed label samples. Where applicable, map ISO 8536-10:2015 §8.2/§8.3 label content to the approved unit/shelf packages. Retain original small-box and master-carton arrangement steps and label control. [GAP-014]

12.3 Complete F-823-i4 and review critical parameter measurements with recipe limits; checks must have objective Pass/Fail, actual value and method where required. [GAP-013,014]

# 13. Sterilization, laboratory and final product disposition

13.1 Finished packed batches undergo sterilization per controlled WI-7-5-1-1 and applicable validated cycle. Trace production lot to sterilization load/cycle, packaging lot and laboratory/qualification records. [GAP-016]

13.2 QC checks unloaded items after sterilization; QA/QCM reviews cycle parameters, exceptions, required laboratory evidence and in-process/final records, followed by the final authorization mandated by company procedures. [Source WI §§9–10]

13.3 Required analytical/biological evidence (e.g. chemical, endotoxin/pyrogen or haemolytic assessment where applicable) shall be controlled in an approved separate verification register with test method, valid scope, approved sampling frequency, linked test report, result and disposition. A draft or blank worksheet is not proof of testing. [GAP-017]

13.4 Final release requires signed F-823-F47 (or approved successor) and complete applicable cycle/lab/packaging/NCR records. QC in-process stamp alone is NOT final product release. Production records finished quantity and transfers to finished goods only after authorized disposition. [GAP-018]

# 14. Nonconforming product / restart

On failure or missing critical result, hold and identify affected quantity dating back to the last acceptable checkpoint; record lot, sample, machine/cavity, time, defect and NCR/hold ID. Notify responsible functions, implement authorized containment/adjustment and perform documented recheck before continued production or approved rework. Never backfill absent results as Pass. [GAP-020]

# 15. Records and document control

Raw material F-823-T1; connector incoming F-823-T90; injection/extrusion F-823-i15; Extension Tube IP F-823-i60; packaging F-823-i4; final release F-823-F47. All numbers/revisions to be checked against controlled master list.

Retain approved TF-11-PS/PDBM, process validation, qualified ISO test method reports, calibration, lab results, sterilization records, nonconformities and training/authorization.

The F-823-i15 source tab still displays Rev.15 while sheet header/file says Rev.16; document control must reconcile before approval.

Record retention, signatures, revision approval and ownership are governed by company QMS; not invented in this draft.

# 16. Proposed revision history and approval

| Status | Baseline | Summary | Approval |
| --- | --- | --- | --- |
| PROPOSED / not in force | Rev.6 dated 15.03.2023 | In-process test controls, acceptance/evidence mapping, source applicability, record integrity, release linkage | QC / QA / Engineering formal review required |

# 17. Open technical inputs (approval blockers)

OPEN TECHNICAL INPUT REQUIRED — TF-11-PS/PDBM/line classifications for all four original tube configurations (and injection site variants).

OPEN TECHNICAL INPUT REQUIRED — Full authorized ISO 8536-10:2015 Annex A/B/C and ISO 8536-9:2015 if relevant; ISO 8536-4 dependent methods.

OPEN TECHNICAL INPUT REQUIRED — Approved calibration/test methods, tensile scope, leakage screen-to-normative evidence, sample/frequency plan.

OPEN TECHNICAL INPUT REQUIRED — Validated bonding settings/cure and packaging recipe / seal acceptance criteria.

OPEN TECHNICAL INPUT REQUIRED — Controlled sterilization, chemistry/biology, label artwork, actual QC and final release records.

# 18. Gap / requirement traceability

The following controlled-draft section links ensure every master gap ID is traceable to a proposed WI or in-process record change; some verifications remain outside routine in-process control.

| Gap ID | Proposed location / evidence route |
| --- | --- |
| GAP-001 | WI 2–3 |
| GAP-002 | WI 3, 6.2, 8.3–8.4 |
| GAP-003 | WI 10.2 |
| GAP-004 | WI 10.1 |
| GAP-005 | WI 9.3, 10.3 |
| GAP-006 | WI 8.2, 10.3 |
| GAP-007 | WI 9.3, 10.3 |
| GAP-008 | WI 9.1–9.2 |
| GAP-009 | WI 8.3–8.4 |
| GAP-010 | WI 11.1 |
| GAP-011 | WI 11.2 |
| GAP-012 | WI 7–8 |
| GAP-013 | WI 12.1, 12.3 |
| GAP-014 | WI 12.2 |
| GAP-015 | WI 6 |
| GAP-016 | WI 13.1 |
| GAP-017 | WI 10.3, 13.3 |
| GAP-018 | WI 13.4 |
| GAP-019 | WI 15 |
| GAP-020 | WI 7.3, 14 |