# ISO CLOSURE REGISTER — PROPOSED

Part 5 product applicability unresolved | Part 4 conditional reference mapping | GAP-001 to GAP-019

Review status: CONTROLLED PROPOSAL - pending technical verification, QC Manager/QA approval and operational evidence.

| Product | Infusion Burette Set / IV Burette Set |
| --- | --- |
| Scope | In-process manufacturing controls (not a final inspection release report) |
| Original WI | WI-8-2-2-30-1, Rev. 0 (13 Jan 2024) |
| Original in-process record | F-823-i75, Rev. 2 (01 Mar 2024) |
| Supplied ISO | BS EN ISO 8536-4:2010 |
| Product-specific ISO needed | ISO 8536-5:2004 (full controlled text NOT provided; applicability requires approval) |
| Assessment date | 15 October 2026 |
| Reviewed source scope | 11-image-page WI, 28-page ISO file, 1-page photographed form |

Document owner/reviewer/signatures: NOT ASSIGNED - proposed working draft. No original or external document has been modified.

# 1. Status definitions / closure gate

NOT ASSESSED - PART 5 MISSING: cannot identify all product-applicable requirements or exact ISO 8536-5 acceptance clauses/pages without controlled text.

CONDITIONAL REFERENCE - PART 4: the quoted ISO 8536-4 clause is verifiable in supplied file but cannot automatically be applied to the burette product.

DOCUMENTARY COVERED - OPERATIONAL EVIDENCE PENDING: proposed WI/form text exists, but relevant actual testing and approvals remain unverified; NOT treated as closed.

VERIFIED CLOSED: only approved technical control plus verified, traceable executed evidence can close a requirement. No such closure demonstrated for the uploaded documents.

# 2. Product-specific ISO 8536-5 primary register

| Requirement group | ISO Says / source state | Gap link | Required evidence | Actual evidence | Documentary / operational status |
| --- | --- | --- | --- | --- | --- |
| ISO 8536-5 complete review | OPEN TECHNICAL INPUT REQUIRED - full controlled ISO 8536-5:2004 text missing. ISO public catalog confirms burette-specific scope for 50, 100 and 150 ml. | GAP-001 | Approved full ISO copy + product scope and clause matrix. | None submitted. | NOT ASSESSED / OPEN |
| Burette capacity and graduated scale | Part 5 exact clause/page, scale tolerances and method unknown; must not be invented. | GAP-004 / GAP-005 | Controlled specification, scale test plan and actual results. | None submitted. | OPEN - TECHNICAL INPUT / EVIDENCE |
| Tubing and fluid delivery | Part 5 applicable limits for specific burette configuration, drop factor and throughput not verified. | GAP-008 / GAP-014 / GAP-015 | Variant BOM/drawing, ISO clauses, qualification/results. | None submitted. | OPEN - TECHNICAL INPUT / EVIDENCE |
| Assembly/air/connector and leakage | Part 5 exact cross-applicability of Parts 4 and other component standards unresolved. | GAP-006 / GAP-010 / GAP-011 / GAP-012 / GAP-016 | Approved risk-based test matrix, methods, results. | None submitted. | OPEN - TECHNICAL INPUT / EVIDENCE |

# 3. ISO 8536-4:2010 requirements - verified text, unverified product applicability

The rows below cover the reference ISO 8536-4 clauses as a screening checklist only. "ISO Says" refers to the supplied Part 4 and is not a Part 5 conformity assertion. The original Part 4 has 3.1, 3.2, 4.1-4.2 and other non-test provisions; this register focuses on the listed technical sections and must be revisited after product-scope authorization.

| Part 4 clause and source page | ISO Says (Part 4; concise) | Gap/change ID | Required control/evidence route | Actual verified evidence | Status |
| --- | --- | --- | --- | --- | --- |
| 3.2 ¦ ISO p.2 / PDF p.10 | Protective caps maintain sterile internal path | GAP-001 / GAP-007 | Visual + sterile barrier qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 5 ¦ ISO p.4 / PDF p.12 | Materials support physical/chemical/biological requirements | GAP-001 / GAP-018 | Material supplier + biocompatibility; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.1 ¦ ISO p.5 / PDF p.13 | Controlled particulate cleanliness; Annex A.1 | GAP-017 | Qualification / particulate; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.2 ¦ ISO p.5 / PDF p.13 | No air leakage under A.2 conditions | GAP-011 | Test evidence; in-process observations linked; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.3 ¦ ISO p.5 / PDF p.13 | Static tensile resistance 15 N for 15 s per A.3 | GAP-010 | Qualified tensile route; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.4 ¦ ISO p.5 / PDF p.13 | Spike dimension/piercing and no coring | GAP-007 | Spike supplier/qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.5 ¦ ISO p.5 / PDF p.13 | Air inlet filter and functionality; A.4 | GAP-006 | Vent component + qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.6 ¦ ISO p.6 / PDF p.14 | Tube visibility and 1500 mm standard-set length | GAP-008 / GAP-009 | Drawing/variant dimensional verification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.7 ¦ ISO p.6 / PDF p.14 | Fluid filter and retention per A.5 | GAP-017 | Fluid filter qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.8 ¦ ISO p.6 / PDF p.14 | Drop chamber visibility/geometry and 20 or 60 drops to approx. 1 ml for Part 4 | GAP-014 | Drop factor / type test; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.9 ¦ ISO p.6 / PDF p.14 | Flow regulator from zero to maximum | GAP-013 | Clamp function; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.10 ¦ ISO p.6 / PDF p.14 | At least 1000 ml 9 g/l NaCl in 10 min at head 1 m for Part 4 | GAP-015 | Throughput test; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.11 ¦ ISO p.6 / PDF p.14 | Self-sealing injection site where provided; Annex A.6 | GAP-012 | Port reseal/pressure testing; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.12 ¦ ISO p.6 / PDF p.14 | Male conical connector reference ISO 594 in Part 4 | GAP-016 | Connector standards decision; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 6.13 ¦ ISO p.6 / PDF p.14 | Protective caps secure, maintain sterility | GAP-007 / GAP-018 | Visual / sterile barrier; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 7.1 ¦ ISO p.7 / PDF p.15 | Reducing/oxidizable matter; Annex B.2 | GAP-001 / GAP-018 | Material/product chemical qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 7.2 ¦ ISO p.7 / PDF p.15 | Metal ions limits; Annex B.3 | GAP-001 / GAP-018 | Chemical qualification / supplier; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 7.3 ¦ ISO p.7 / PDF p.15 | Acidity/alkalinity extract titration; Annex B.4 | GAP-001 / GAP-018 | Chemical qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 7.4 ¦ ISO p.7 / PDF p.15 | Residue on evaporation; Annex B.5 | GAP-001 / GAP-018 | Chemical qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 7.5 ¦ ISO p.7 / PDF p.15 | UV absorbance of extract; Annex B.6 | GAP-001 / GAP-018 | Chemical qualification; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 8.1 ¦ ISO p.7 / PDF p.15 | Biological compatibility assessment | GAP-001 / GAP-018 | Biological evaluation; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 8.2 ¦ ISO p.7 / PDF p.15 | Validated sterilization process | GAP-018 | Validated sterilization evidence; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 8.3 ¦ ISO p.7 / PDF p.15 | Pyrogenic assessment | GAP-018 | Biological/product evidence; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 8.4 ¦ ISO p.7 / PDF p.15 | Haemolysis evaluation | GAP-018 | Biological evaluation; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 8.5 ¦ ISO p.8 / PDF p.16 | Toxicity assessment | GAP-018 | Biological evaluation; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 9.1 ¦ ISO p.8 / PDF p.16 | Unit container labeling particulars | GAP-018 | Packaging/labeling controlled record; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 9.2 ¦ ISO p.8 / PDF p.16 | Shelf/multiunit labeling particulars | GAP-018 | Packaging/labeling controlled record; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 10.1 ¦ ISO p.9 / PDF p.17 | Sealed sterile/tamper-evident unit pack | GAP-018 | Packaging / sterile barrier; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |
| 10.2 ¦ ISO p.9 / PDF p.17 | Packaging avoids flattened/kinked tubes | GAP-018 | Packaging / visual; establish approved relevance to burette + applicable source | None supplied | CONDITIONAL / OPEN - APPROVAL & EVIDENCE |

# 4. Cross-reference: every gap and status

| Gap ID | Priority | Issue | Draft coverage | Required evidence | Status |
| --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 | Incorrect product-specific ISO reference | Proposed wording exists; not approved. | controlled ISO 8536-5 copy, approved applicability/edition assessment, configuration matrix and authorized reference update. | OPEN |
| GAP-002 | P2 | WI and record control-stage ambiguity | Proposed wording exists; not approved. | approved control-plan flow, controlled form list, linked completed records. Actual evidence: other referenced records not su | OPEN |
| GAP-003 | P2 | Sampling, frequency and traceability not defined | Proposed wording exists; not approved. | approved control plan and sampling rationale, executed checkpoint sheets, machine/lot traceability, equipment list. Actual: | OPEN |
| GAP-004 | P1 | Burette scale accuracy not verifiable | Proposed wording exists; not approved. | applicable Part 5 clauses, approved drawing/tolerance, validated method, calibrated volume equipment, completed scale accura | OPEN |
| GAP-005 | P2 | Burette assembly/hanging features lack defined criteria | Proposed wording exists; not approved. | approved component drawing, assembly acceptance catalog, any qualified hanger/valve retention data, sampled observations. | OPEN |
| GAP-006 | P1 | Air vent filter verification insufficient | Proposed wording exists; not approved. | selected standard, approved vent design, filter qualification evidence and traceable assembly check. Actual NOT PROVIDED. | OPEN |
| GAP-007 | P2 | Spike performance not defined beyond appearance | Proposed wording exists; not approved. | component drawing, supplier report/qualification or actual piercing data if applicable, linked in-process inspection. | OPEN |
| GAP-008 | P1 | Tubing length and dimension conflict | Proposed wording exists; not approved. | controlled drawing, current product configuration, ISO 8536-5 clause, documented disposition for 100 cm option, calibrated m | OPEN |
| GAP-009 | P2 | Subjective visual and surface cleanliness terms | Proposed wording exists; not approved. | controlled visual defect standard and executed checks; separate particle-test route if applicable. | OPEN |
| GAP-010 | P1 | Joining integrity/tensile result not documented | Proposed wording exists; not approved. | approved tensile/control plan, calibrated force gauge, actual test record or qualified linked evidence. | OPEN |
| GAP-011 | P1 | Leak-test method and recording gap | Proposed wording exists; not approved. | confirmed Part 5 approach, approved leakage protocol, calibrated equipment, executed test data, linked NCR if fail. | OPEN |
| GAP-012 | P1 | Injection valve/bulb closure integrity underspecified | Proposed wording exists; not approved. | approved port-specific test basis, method and actual reseal/leakage result, traceability. | OPEN |
| GAP-013 | P2 | Clamp shutoff and regulation acceptance unclear | Proposed wording exists; not approved. | clamp functional test method, approved acceptance, QC data. | OPEN |
| GAP-014 | P1 | Drop factor contradictory and mislabeled as flow rate | Proposed wording exists; not approved. | Part 5 full text, approved product drop factor, revised specification, actual calibrated drop measurement, approved discrepa | OPEN |
| GAP-015 | P1 | Infusion throughput not assessed by drop-factor row | Proposed wording exists; not approved. | ISO 8536-5 flow method/limit and engineering/risk route, executed flow test or approved linked qualification. | OPEN |
| GAP-016 | P2 | Luer connector interface verification weak | Proposed wording exists; not approved. | approved connector standard applicability, qualified gauge/test records, part traceability. | OPEN |
| GAP-017 | P1 | Fluid filter/particle-control route not evidenced | Proposed wording exists; not approved. | filter BOM/spec, appropriate ISO 8536-5/part4 applicability, approved supplier or qualification record and executed visual c | OPEN |
| GAP-018 | P2 | Linked packaging, sterilization and release records unverified | Proposed wording exists; not approved. | actual F-823-T38/i4/i15 controlled copies, approved sterilization and release records, lot linkage. None supplied; Open - Ev | OPEN |
| GAP-019 | P1 | Failure containment and disposition trail not captured | Proposed wording exists; not approved. | approved nonconforming process, sample executed NCR, hold/release evidence and review; actual evidence absent. Open - Approv | OPEN |

# 5. Closure signoff (not completed)

| Gate / authority | Required document or verification | Review status |
| --- | --- | --- |
| Engineering + QA | Confirm current standard Part 5 and Part 4 cross-application; set specifications/tolerances. | PENDING |
| QC Manager | Authorize WI revision, sampling/verification route, form fields and revised critical acceptance criteria. | PENDING |
| QC operational representative | Review actual manufacturing evidence and complete measurements and checks. | PENDING |
| QA / Document Control | Approve and issue controlled WI and F-823-i75 revision; ensure linked forms/release controls. | PENDING |
| Final ISO closure | Verify each applicable clause with approved control and actual evidence. | NOT VERIFIED |

Sources: ISO 8536-4:2010 supplied PDF: PDF pp.8-25; photographed WI WI-8-2-2-30-1 Rev0 pp.1-11; photographed F-823-i75 Rev2 p.1; ISO catalog https://www.iso.org/standard/20921.html. No missing Part 5 technical clause/limit was guessed.