# QC MANAGER — MASTER GAP ANALYSIS

Infusion Burette Set | ISO applicability, existing WI and F-823-i75 in-process record

Review status: CONTROLLED PROPOSAL - pending technical verification, QC Manager/QA approval and operational evidence.

| Product | Infusion Burette Set / IV Burette Set |
| --- | --- |
| Scope | In-process manufacturing controls (not a final inspection release report) |
| Original WI | WI-8-2-2-30-1, Rev. 0 (13 Jan 2024) |
| Original in-process record | F-823-i75, Rev. 2 (01 Mar 2024) |
| Supplied ISO | BS EN ISO 8536-4:2010 |
| Product-specific ISO needed | ISO 8536-5:2004 (full controlled text NOT provided; applicability requires approval) |
| Assessment date | 15 October 2026 |
| Reviewed source scope | 11-image-page WI, 28-page ISO file, 1-page photographed form |

Document owner/reviewer/signatures: NOT ASSIGNED - proposed working draft. No original or external document has been modified.

# 1. Executive summary

Total evidence-based observations: 19 | P1: 11 | P2: 8 | P3: 0. These are documentation/control observations, NOT a measured product defect rate or operational nonconformity count.

Core blocker: The supplied standard is ISO 8536-4:2010, but the specific burette product standard is ISO 8536-5:2004. The full Part 5 text and approved product drawings/specifications are missing. As a result, 100% ISO requirement coverage, final numeric acceptance criteria and verified conformity cannot be asserted. ISO 8536-4 values below are explicitly marked conditional; do not implement them blindly.

The source record is an IN-PROCESS manufacturing control form, not a conventional final inspection or product release form. The strongest discrepancies visible now are the 100/150 cm tubing length vs the baseline >=1500 mm WI text, the "10 drops / 1 ml" entry vs the WI 20/60-drop description, burette scale accuracy without measurable criteria, and test fields without actual values/methods.

Manager decision: authorize a controlled Part 5 applicability study; resolve product variants, drawing/PO limits and drop factors; approve which tests remain qualification/periodic rather than routine; align the WI and in-process record, then validate on actual production evidence. No closure is verified from the supplied files.

# 2. Documents and source status

| Source | Identity | Evidence state |
| --- | --- | --- |
| WI (photographed PDF) | WI-8-2-2-30-1; Rev. 0; 13 Jan 2024; 11 printed pages | Controlled-copy marking is visible; formal current revision/supersession not independently verified. |
| In-process control form (image) | F-823-i75; Rev. 2; 01 Mar 2024; 1/1; 18 checks | Form photographed blank; no lot data or actual completed evidence. |
| Supplied standard | BS EN ISO 8536-4:2010; 28 PDF pages; ISO pp.1-17 at PDF pp.9-25 | Source fully readable; product-specific applicability NOT established. |
| Required missing standard | ISO 8536-5:2004; future 2026 FDIS is draft | Full controlled text not provided; exact clauses/limits/page references must remain OPEN. |
| Company controls | Product drawing, spec, BOM, PO, control plan, F-823-T38 / i4 / i15, equipment & qualification records | Not submitted; their existence/approval/execution cannot be verified. |

# 3. Audit method and priority basis

Sources were compared line-by-line from photographed WI printed pages 1-11, one photographed form with Rows 1-18, and the supplied Part 4 PDF. The ISO Says column distinguishes directly verified Part 4 clauses from pending applicability of Part 5. P1 means plausible patient/product-safety or critical acceptance impact; P2 means traceability/method/documentary verification impact; P3 means non-critical editorial defect.

# 4. Initial findings

| ID | Priority | Short finding | Key source |
| --- | --- | --- | --- |
| GAP-001 | P1 | Incorrect product-specific ISO reference | WI p.11, Section 9 References: "BS EN ISO 8536-4 infusion equipment for medical use, infusion s |
| GAP-002 | P2 | WI and record control-stage ambiguity | WI pp.1-3 Sections 5.1-5.2 (receiving), p.6 Section 5.4 (inspection/test report), p.10 Section |
| GAP-003 | P2 | Sampling, frequency and traceability not defined | WI p.10 Section 7 states QC takes samples from packed quantity every 2 h based on AQL |
| GAP-004 | P1 | Burette scale accuracy not verifiable | WI p.2 (Burette Chamber): "clear and transparent with printed graduation scale" and good fittin |
| GAP-005 | P2 | Burette assembly/hanging features lack defined criteria | WI p.2 describes burette chamber, bottom cap, shut-off valve and hanger as proper fittings/no d |
| GAP-006 | P1 | Air vent filter verification insufficient | WI p.4 Section 5.3 states "Air inlet port should be with an air filter to prevent the ingress o |
| GAP-007 | P2 | Spike performance not defined beyond appearance | WI p.3 Section 5.2 "Spike": sharp tip, no tip damage, clear fluid channels, properly fixed air- |
| GAP-008 | P1 | Tubing length and dimension conflict | WI p.4 Section 5.3: "The tubing from the distal end to the drip chamber shall be not less than |
| GAP-009 | P2 | Subjective visual and surface cleanliness terms | WI pp.2-3 Sections 5.2 and p.10 Section 6: "no black spots", "no dust particle", clear and kink |
| GAP-010 | P1 | Joining integrity/tensile result not documented | WI p.4 Section 5.3 "Tensile Test": force of 15 N for 15 seconds. WI pp.2-3 require strong gluin |
| GAP-011 | P1 | Leak-test method and recording gap | WI p.4 Section 5.3: "Immerse the infusion set into water and apply an internal air pressure of |
| GAP-012 | P1 | Injection valve/bulb closure integrity underspecified | WI p.5 Section 5.3 describes puncturing the injection site with a 0.8 mm outside diameter hypod |
| GAP-013 | P2 | Clamp shutoff and regulation acceptance unclear | WI p.5 Section 5.3 "Flow Regulator" explains flow between zero and maximum, usable during infus |
| GAP-014 | P1 | Drop factor contradictory and mislabeled as flow rate | WI p.8 Section 5.3 (Drip chamber and drip tube): specifies "20 drops ... or 60 drops ... (1 +/- |
| GAP-015 | P1 | Infusion throughput not assessed by drop-factor row | WI p.9 Section 5.3 "Flow rate of infusion fluid": at least 1000 ml of 9 g/l NaCl in 10 min unde |
| GAP-016 | P2 | Luer connector interface verification weak | WI p.5 "Male Conical Fitting": compatibility with ISO 80369-7 (as written) |
| GAP-017 | P1 | Fluid filter/particle-control route not evidenced | WI p.2 describes "Fluid Filter" fitted in chamber outlet and free of spots/fibers/damage. WI p. |
| GAP-018 | P2 | Linked packaging, sterilization and release records unverified | WI p.10 Section 7 packaging/labeling includes 2-hour packed quantity sampling |
| GAP-019 | P1 | Failure containment and disposition trail not captured | WI p.11 Section 8 references QC inspection/release to QCM/QAM |

# 5. MASTER GAP & REVISION TABLE

MANDATORY SEVEN-COLUMN COMPARISON. Each row is a proposed correction, not a released instruction. "Inspection" in the required comparison headings refers here to the photographed IN-PROCESS CONTROL form; no change of process classification is intended.

| Gap / Our Problem | ISO Says | Existing WI Text | Revised WI Text | Existing Inspection Text | Revised Inspection Text | Evidence (ISO Closure) |
| --- | --- | --- | --- | --- | --- | --- |
| GAP-001 [P1] — Incorrect product-specific ISO reference The supplied ISO 8536-4 addresses conventional gravity-feed infusion sets, whereas the product has a volumetric burette. The WI uses Part 4 without a verified applicability assessment of the product-specific Part 5. A complete product ISO compliance conclusion is therefore blocked. | ISO 8536-4:2010 Foreword, PDF p.8 / printed iv, explicitly lists Part 5 for burette infusion sets. ISO 8536-5:2004 scope, as described by ISO, covers gravity-feed burette infusion sets of 50, 100 and 150 ml nominal capacity. Exact Part 5 clauses/pages are OPEN TECHNICAL INPUT REQUIRED. | WI p.11, Section 9 References: "BS EN ISO 8536-4 infusion equipment for medical use, infusion sets for single use, gravity feed." No Part 5 applicability matrix is shown. | [References and technical applicability] Before approval, QA shall obtain a controlled copy of the applicable edition of ISO 8536-5; document product capacity/configuration and its clause-by-clause application. ISO 8536-4 requirements may be used only where a documented, approved technical basis supports cross-application. Open requirements shall not be represented as verified. | Form F-823-i75 p.1, header: "In-process Control for IV Burette Set" with WI/TF references; no verified ISO applicability field. | Applicable standard/edition: [approved reference]; product capacity/model: [ ]; ISO applicability-matrix ID/revision: [ ]; approval status: Pending / Approved; verification evidence ID: [ ]. Keep in-process checks separate from ISO conformity decisions. | Required: controlled ISO 8536-5 copy, approved applicability/edition assessment, configuration matrix and authorized reference update. Actual evidence: NOT PROVIDED. Closure: Open - Technical Input / Approval. |
| GAP-002 [P2] — WI and record control-stage ambiguity The baseline WI combines receiving inspection, in-process controls, packaging and final release. The photographed F-823-i75 is an in-process log and does not itself document the final inspection/testing/release steps defined by the WI. | No direct Part 4 numeric requirement. This is a source-based process/document-control alignment issue; ISO 8536-4 Section 1, PDF p.9 / printed p.1, does not prescribe company in-process form architecture. | WI pp.1-3 Sections 5.1-5.2 (receiving), p.6 Section 5.4 (inspection/test report), p.10 Section 7 (packaging), pp.10-11 Section 8 (final inspection and release). | [Scope / records] Classify F-823-i75 as in-process manufacturing control only. Maintain distinct and traceable records for receiving, qualified testing, packaging verification and final release. A satisfactory in-process entry shall not authorize lot release. Cross-reference the approved WI and release instruction. | F-823-i75 p.1 contains 18 hourly observation rows, "Checked by Q.C." and "Reviewed by Q.C.M."; no final-release authorization. | Record category: In-Process Control (not Final Inspection). References to receiving/qualification/packaging/final-release records: [document ID, revision, lot, record ID]. Final release status: "Not determined by this form." | Required: approved control-plan flow, controlled form list, linked completed records. Actual evidence: other referenced records not supplied. Closure: Open - Evidence / Approval. |
| GAP-003 [P2] — Sampling, frequency and traceability not defined Clock-hour boxes record only acceptability; the inspected quantity, sampling basis, production line/machine, inspected units, examiner and equipment traceability are not captured. It is unclear whether every box represents a sampled check or a required test. | No ISO 8536-4 sampling frequency is established for this in-process log. Any frequency/AQL must be sourced from the approved company control plan; the Part 4 standard cannot be used to invent hourly sampling. | WI p.10 Section 7 states QC takes samples from packed quantity every 2 h based on AQL; WI p.11 Section 8 cites AQL general/special inspection. These do not establish the frequency/sample size for all 18 in-process rows. | [Sampling and documentation] At each in-process checkpoint record actual time, shift, line/machine, lot, sample identifier, sample quantity, applicable approved sampling-plan ID/revision, inspected parameter(s), measured result or observation, disposition and operator/QC initials. Apply frequency only as stated by the approved control plan; do not interpret the printed hourly headers as test-frequency requirements. | F-823-i75 p.1 header: Size #, Lot #, Date, Shift; 12 clock time columns; legend Acceptable / Not Acceptable / Read Remarks / N/A; equipment column unfilled by design. | Header fields: size/capacity, lot, line/machine, date, shift, WO/PO and control-plan ID/rev. Checkpoint log: actual time, process step, sample IDs/qty, applicable test IDs, result, P/F/NA, inspected by, equipment/calibration ID, NCR if failed. | Required: approved control plan and sampling rationale, executed checkpoint sheets, machine/lot traceability, equipment list. Actual: NOT PROVIDED. Closure: Open - Technical Input / Evidence. |
| GAP-004 [P1] — Burette scale accuracy not verifiable The record asks for "clear and accurate printing" without defining scale-volume accuracy, acceptance tolerance, test setup or actual measurement; a visual yes/no cannot prove burette metering accuracy. | Product-specific graduated burette requirements must be taken from controlled ISO 8536-5 and approved product specification. Part 4:2010 Clause 6.8, PDF p.14 / printed p.6, relates to ordinary drip-chamber drop formation and does NOT establish burette scale accuracy. Numeric Part 5 limit: OPEN TECHNICAL INPUT REQUIRED. | WI p.2 (Burette Chamber): "clear and transparent with printed graduation scale" and good fitting. No verifiable graduations tolerance or calibration method is visible. | [Burette graduation verification] Confirm nominal capacity/model against the approved drawing. Verify printed graduation legibility and alignment. Where required by the applicable approved standard/control plan, fill/measure at specified scale points under a documented method and record indication, reference volume, deviation and applicable tolerance: [OPEN TECHNICAL INPUT REQUIRED - controlled ISO 8536-5 / product specification]. Reject or quarantine results outside established limits. | F-823-i75 p.1 Row 10 "Chamber graduation scale"; acceptance "Clear and accurate printing"; only timed P/F cells. | Graduation inspection: legibility [P/F]; nominal capacity [ml]; designated scale point(s) [ml]; indicated volume [ml]; reference volume [ml]; deviation [ml or %]; approved tolerance/ref [ ]; method/equipment ID [ ]; status Pending/P/F; evidence [ ]. Use actual measurement only at the approved checkpoint/frequency. | Required: applicable Part 5 clauses, approved drawing/tolerance, validated method, calibrated volume equipment, completed scale accuracy results. Actual: NONE. Closure: Open - Technical Input / Evidence. |
| GAP-005 [P2] — Burette assembly/hanging features lack defined criteria "Straight soft", "strong fitting" and "no pull out" are subjective without a defined inspection or retention specification for chamber body, hanger, vent/inlet interfaces and connections. | ISO 8536-5 applicability and component requirements must be verified from the complete standard; ISO 8536-4 figures/p.10-12 illustrate standard sets, not performance limits for burette hanging strips. | WI p.2 describes burette chamber, bottom cap, shut-off valve and hanger as proper fittings/no damage; WI p.3 includes hanger/valve and chamber. Numeric strength criterion not verified. | [Burette assembly] Confirm chamber clarity/integrity, secure top/bottom cap and hanger, unrestricted prescribed valve movement and absence of visible cracks or detachment. Where retention-force testing is required, use the approved drawing/control plan force and method [OPEN TECHNICAL INPUT REQUIRED]. Record assembly defects and identify affected lots. | F-823-i75 Rows 8 "Chamber Quality" (straight soft, no leakage) and 9 "Chamber Hanging Strip" (strong fitting, no pull out, flexible). | Chamber/hanger checks: condition [ ]; attachment inspection [ ]; test method/spec revision [ ]; measured retention (if applicable) [N]; acceptance [ ]; result [P/F/NA]; related defect/NCR [ ]. | Required: approved component drawing, assembly acceptance catalog, any qualified hanger/valve retention data, sampled observations. Closure: Open - Technical Input / Evidence. |
| GAP-006 [P1] — Air vent filter verification insufficient A good fit of air filter/cap does not verify filtration performance, vent function or air path integrity where such performance is required for the applicable configuration. | ISO 8536-4:2010 Clauses 6.5 and 3.2, PDF pp.13 and 10 / printed pp.5 and 2, define air-inlet filtration and protective aspects for sets in scope. Clause 6.5 describes a flow reduction limit of 20% for a specified configuration (Annex A.4). Application to burette assembly remains UNCONFIRMED pending Part 5. | WI p.4 Section 5.3 states "Air inlet port should be with an air filter to prevent the ingress of microorganisms into the container"; no in-process verification method beyond visual confirmation. | [Air vent] Inspect correct filter/cap part number, secure placement, no visible damage or blockage. Confirm functional/filtration requirements and associated test route from the approved ISO 8536-5 applicability matrix, design file and risk plan. Record inspection plus cross-reference qualified air-path/flow/filter validation where relevant. Do not apply Part 4 20% acceptance limit without approved applicability. | F-823-i75 p.1 Row 6 "Air vent filter status": "Good fitting of filter and cap". | Vent/filter component ID [ ]; filter/cap inspection [P/F]; vent condition [ ]; qualification test ID/revision and certificate [ ]; verified control route [ ]; applicable limit [OPEN]; evidence ref [ ]; result [P/F/Pending]. | Required: selected standard, approved vent design, filter qualification evidence and traceable assembly check. Actual NOT PROVIDED. Closure: Open - Technical Input / Evidence. |
| GAP-007 [P2] — Spike performance not defined beyond appearance Visual sharpness, open tip and no blockage do not independently substantiate functional penetration or dimensional compatibility when required by the approved product design. | ISO 8536-4:2010 Clause 6.4, PDF p.13 / printed p.5, requires a specified closure-piercing device geometry and penetration without pre-piercing; scope for this burette model must be assessed under Part 5 and controlled drawings. | WI p.3 Section 5.2 "Spike": sharp tip, no tip damage, clear fluid channels, properly fixed air-inlet port, no blockage/crack/leakage. | [Spike] Verify specified spike model, intact tip, clean/open channels and correct assembly. Verify geometry/penetration by the approved incoming, type or periodic test method as assigned in the control plan; link its test report. Do not assert that visual fit proves dimensional or piercing conformity. | F-823-i75 Row 4 "Spike Status": "Spike sharp tip, open tip No blockage protector good fitting". | Spike visual [P/F]; component/drawing ID [ ]; critical dimensions [mm, if sampled under plan]; penetration test reference [ ]; equipment [ ]; result [P/F/Pending]; remarks [ ]. | Required: component drawing, supplier report/qualification or actual piercing data if applicable, linked in-process inspection. Closure: Open - Technical Input / Evidence. |
| GAP-008 [P1] — Tubing length and dimension conflict WI gives at least 1500 mm for a section of tubing; form Row 12 accepts 100 cm or 150 cm as PO. Without configuration mapping, a 100 cm acceptance can conflict with the current WI. Tubing OD/ID row likewise provides no tolerance. | ISO 8536-4:2010 Clause 6.6, PDF p.14 / printed p.6, states at least 1500 mm for standard-set tubing from distal end to drip chamber. This is not automatically a burette limit: ISO 8536-5 requirement and approved configuration must be verified. | WI p.4 Section 5.3: "The tubing from the distal end to the drip chamber shall be not less than 1500 mm in length". WI p.3 states tubing dimensions as per specification. | [Tubing dimensional check] Use the approved model-specific drawing and PO-to-product mapping. Confirm total/segment length definition, nominal length and tolerance, inner and outer diameter with calibrated equipment and record actual values. Flag and segregate any 100 cm variant unless Engineering/QA documents that it is permitted under the applicable standard and approved product configuration. Do not silently waive the existing WI criterion. | F-823-i75 Row 11 "Tube Dimension": "3.0 X 4.1 mm, as per P.O"; Row 12 "Tube Length": "100 cm, 150 cm, as per P.O". | Model/variant [ ]; controlled drawing/PO [ ]; approved segment length [mm] / tolerance [ ]; measured length [mm] / result; ID [mm] / OD [mm] / tolerances / actual [ ]; instrument calibration ID [ ]; affected-lot hold decision [ ]. | Required: controlled drawing, current product configuration, ISO 8536-5 clause, documented disposition for 100 cm option, calibrated measurement records. Closure: Open - Technical Input / Evidence. |
| GAP-009 [P2] — Subjective visual and surface cleanliness terms Generic clarity/appearance, black spots, flashing and flexible/kink observations are not tied to a defect catalog, sample/lighting method or quantified acceptance where needed; components and fluid pathway cleanliness are conflated. | ISO 8536-4:2010 Clauses 6.1 and 6.6, PDF pp.13-14 / printed pp.5-6, address particulate control and tubing visibility, for sets in scope. ISO 8536-5 application must be confirmed; visual check alone cannot represent particulate qualification. | WI pp.2-3 Sections 5.2 and p.10 Section 6: "no black spots", "no dust particle", clear and kink resistance; inspection lighting, defect categories and rejection threshold are not given. | [Visual checks] Examine each sampled component under the controlled visual inspection procedure [OPEN: reference], with approved lighting, magnification if specified, and defect definitions. Record location, type and count for nonconformities; maintain separate evidence for particle-control tests where required. | F-823-i75 Rows 1-3 "Clarity, Appearance", "Free from irregularities", "Tube Clarity"; acceptance includes transparent, no flashing, no black spots, no bending. | Actual defect code/quantity [ ]; component/zone [ ]; visual WI / defect catalog ID [ ]; light/magnification [if applicable]; sample size [ ]; result [P/F]; photo/NCR evidence [ ]. | Required: controlled visual defect standard and executed checks; separate particle-test route if applicable. Closure: Open - Technical Input / Evidence. |
| GAP-010 [P1] — Joining integrity/tensile result not documented "Strong gluing" and "no pullout" do not record actual pull force or traceable qualified test, whereas baseline WI already requires a static tensile performance check. | ISO 8536-4:2010 Clause 6.3, PDF p.13 / printed p.5, and Annex A.3, PDF p.19 / printed p.11, specify 15 N static tension for 15 s (excluding caps), for sets within its scope. Part 5 applicability to assembled burette variant is OPEN. | WI p.4 Section 5.3 "Tensile Test": force of 15 N for 15 seconds. WI pp.2-3 require strong gluing/no pullout, without defined joint-specific sampling route. | [Joint/tensile integrity] Visually inspect adhesive and connector joints for complete bonding and damage. Perform tensile testing only at the approved sampling/qualification frequency under the applicable control method. Record force [N], hold time [s], point of application, failure mode and Pass/Fail. Until applicability is established, the 15 N/15 s value remains a baseline WI requirement subject to technical approval, not a presumed Part 5 limit. | F-823-i75 Rows 9, 13 and 16 require strong attachment/gluing/fitting; no force or duration fields. | Joint/part ID [ ]; visual result [ ]; tensile/retention method and approval [ ]; force [N]; duration [s]; actual break/slip [ ]; instrument ID [ ]; sample ID [ ]; test report ref [ ]; status [ ]. | Required: approved tensile/control plan, calibrated force gauge, actual test record or qualified linked evidence. Closure: Open - Technical Input / Evidence. |
| GAP-011 [P1] — Leak-test method and recording gap The in-process log captures only "No Leakage At All Areas" without pressure, medium, temperature, time, method, tested region or result. The current WI mentions positive pressure but the Part 4 annex includes conditioning and a second negative-pressure step. | ISO 8536-4:2010 Clause 6.2, PDF p.13 / printed p.5, and Annex A.2.1-A.2.3, PDF p.19 / printed p.11, specify no air leakage; at (40 +/- 1) deg C, +20 kPa for 15 s in water and -20 kPa for 15 s with degassed water. Applicability to burette model needs approved ISO 8536-5 review. | WI p.4 Section 5.3: "Immerse the infusion set into water and apply an internal air pressure of 20 kPa for 15 second"; no 40 deg C condition shown. WI p.7 describes separate connector leak tests. | [Leakage verification] Identify the test object (assembled burette set / specific subassembly / Luer), approved method and revision, equipment, test medium, required pressure, conditioning temperature and exposure duration. Record all actual values and observed leakage. Apply the product-approved Part 5/engineering criteria. Perform formal leak testing at the approved control-plan frequency; routine assembly checks shall not substitute for verified leak data. | F-823-i75 Row 18 "Leak Test": "No Leakage At All Areas" and Rows 5/7/8/15 also say no leakage; only P/F boxes; no measurement/equipment field. | Leak test scope [ ]; method/ref [ ]; pressure [kPa] actual [ ]; temperature [deg C] actual [ ]; hold [s] actual [ ]; leak observation [ ]; test medium [ ]; unit/sample [ ]; calibrated gauge ID [ ]; evidence/worksheet [ ]; result P/F/Pending. | Required: confirmed Part 5 approach, approved leakage protocol, calibrated equipment, executed test data, linked NCR if fail. Closure: Open - Technical Input / Evidence. |
| GAP-012 [P1] — Injection valve/bulb closure integrity underspecified "Rubber valve good fitting" or "injection bulb good fitting" does not show resealing function after puncture or leak performance as described by the baseline WI for injection sites. | ISO 8536-4:2010 Clause 6.11, PDF p.14 / printed p.6, addresses self-sealing injection site and leakage of not more than one falling drop under Annex A.6, PDF p.21 / printed p.13, where the site is fitted. Only conditionally informative until ISO 8536-5 design applicability is verified. | WI p.5 Section 5.3 describes puncturing the injection site with a 0.8 mm outside diameter hypodermic needle, pressurizing to 50 kPa and observing during 1 min for excess leakage; check exact scope against product design. | [Injection port integrity] Verify the installed port/bulb design, correct placement and damage-free condition. If intended as a puncturable self-sealing port, verify resealing and leakage using the approved test method, needle definition, test pressure and duration. Record observed drops/leak rate with source-derived acceptance, test sample and evidence reference. Do not assume two different port types are equivalent. | F-823-i75 Row 7 "Injection valve": "Rubber Valve good fitting No leakage"; Row 15 "Injection bulb status": "Good Fitting, No leakage". | Port type/model [ ]; condition [ ]; functional test applicability [ ]; needle/method [ ]; actual pressure [kPa]; time [s]; observed drops/volume [ ]; acceptance [ ]; equipment/sample [ ]; outcome P/F/Pending; report ID [ ]. | Required: approved port-specific test basis, method and actual reseal/leakage result, traceability. Closure: Open - Technical Input / Evidence. |
| GAP-013 [P2] — Clamp shutoff and regulation acceptance unclear Clamp fit and roller-clamp strength are recorded but not the actual ability to regulate flow from open to closed, prevent unintended restriction or maintain closure under an approved condition. | ISO 8536-4:2010 Clause 6.9, PDF p.14 / printed p.6, requires flow adjustment from zero to maximum; specific application to burette model requires Part 5 review. The standard does not define a universal lot sampling frequency. | WI p.5 Section 5.3 "Flow Regulator" explains flow between zero and maximum, usable during infusion without tubing damage. WI p.3 mentions white clamp complete flow block. | [Clamp/function] Inspect the specified clamp/roller type, full travel and absence of tubing damage. Verify the fully closed and open positions using the approved functional method and acceptance criteria [controlled specification]. Record clamp type, observed flow/blocking result and sample ID. Separate inlet white clamp and downstream roller clamp. | F-823-i75 Row 5 "Clamp status": "No Leakage after clamping, flexible"; Row 14 "Roller Clamp": "Strong Clamping No leakage". | Clamp type/location [ ]; full travel [P/F]; shutoff result/criterion [ ]; flow adjustment [ ]; tubing damage [ ]; method ID [ ]; sample [ ]; observation [ ]; P/F [ ]. | Required: clamp functional test method, approved acceptance, QC data. Closure: Open - Technical Input / Evidence. |
| GAP-014 [P1] — Drop factor contradictory and mislabeled as flow rate Form Row 17 labels "Flow Rate" but its limit "10 drop, 1 ml" is a drop-factor statement and conflicts with the baseline WI drop-volume requirement. The inspector cannot determine which approved target applies. | ISO 8536-4:2010 Clause 6.8, PDF p.14 / printed p.6, gives alternative conventional drop factors of 20 or 60 drops to 1 +/- 0.1 ml at 23 +/- 2 deg C and 50 +/- 10 drops/min. This Part 4 criterion is NOT automatically the permitted burette drop factor under Part 5. | WI p.8 Section 5.3 (Drip chamber and drip tube): specifies "20 drops ... or 60 drops ... (1 +/- 0.1) ml", with specified water temperature and drop rate. | [Drop-volume / drop-factor verification] Verify the approved drop factor and tolerance for each product variant against the applicable Part 5 requirement and design specification [OPEN TECHNICAL INPUT REQUIRED]. Define the test fluid, temperature, count interval/rate and measured volume. Record number of drops and actual measured volume; calculate drop factor if required. Reconcile the WI and F-823-i75 before further use. | F-823-i75 Row 17 "Flow Rate": "10 drop, 1 ml"; only timed P/F columns, no temperature, volume or counting data. | Rename row "Drop factor / drop-volume check"; approved nominal [drops/ml] and allowed measured volume [ml]: [OPEN]; actual number of drops [ ]; volume [ml]; temperature [deg C]; drop rate [drops/min]; test report/sample ID [ ]; P/F/Pending. | Required: Part 5 full text, approved product drop factor, revised specification, actual calibrated drop measurement, approved discrepancy disposition. Closure: Open - Technical Input / Evidence. |
| GAP-015 [P1] — Infusion throughput not assessed by drop-factor row The throughput test described by the WI assesses a volume delivered over time at a defined static head. The form instead documents a drop-count equivalence, which is a different characteristic. | ISO 8536-4:2010 Clause 6.10, PDF p.14 / printed p.6, requires delivery of at least 1000 ml of 9 g/l NaCl in 10 min at 1 m static head, for its product scope. Part 5 applicability and any different limits must be verified; do not automatically require this volume for a burette. | WI p.9 Section 5.3 "Flow rate of infusion fluid": at least 1000 ml of 9 g/l NaCl in 10 min under 1 m static head. | [Infusion flow performance] Identify and apply the approved burette-specific liquid-flow test, setup, fluid and acceptance threshold [OPEN TECHNICAL INPUT REQUIRED]. Record static head, delivered volume, elapsed time and actual flow rate, where the test is applicable. Document whether verification is qualification, periodic or routine according to the approved control plan. | F-823-i75 Row 17 "Flow Rate" contains only "10 drop,1 ml" (drop factor). No volume/time/static-head field. | Separate reference-only field "Flow performance qualification / periodic record": method, fluid, head [m], elapsed time [min], delivered volume [ml], acceptance [ ], actual flow [ml/min], record ID/revision, applicability [ ], outcome [ ]. Do not add hourly test without plan. | Required: ISO 8536-5 flow method/limit and engineering/risk route, executed flow test or approved linked qualification. Closure: Open - Technical Input / Evidence. |
| GAP-016 [P2] — Luer connector interface verification weak A "good luer fitting" observation cannot by itself establish required dimensional/mechanical interoperability and leakage performance. The WI and Part 4 use different connector-standard references requiring an applicability decision. | ISO 8536-4:2010 Clause 6.12, PDF p.14 / printed p.6, refers to ISO 594-1/594-2 male conical fittings. WI p.5 references ISO 80369-7; selection of currently applicable fitting specification is an OPEN TECHNICAL INPUT. Do not combine limits from different editions without approval. | WI p.5 "Male Conical Fitting": compatibility with ISO 80369-7 (as written); WI pp.6-7 describe connector-specific leakage, stress cracking and separation tests. | [Luer interface] Verify male connector and cap part IDs and mating compatibility to the company-approved connector standard/edition. Record visual thread/fitting condition at routine checkpoints; link approved dimensional gauge and leak/separation test evidence under assigned inspection/qualification route. Do not treat cap fit alone as conformity. | F-823-i75 Row 16 "Male Connector": "Good luer fitting No flashing tight cap fitting". | Connector/cap ID [ ]; mating/gauge condition [ ]; selected standard/edition [ ]; dimensions/functional result if checked [ ]; approved method [ ]; qualification report/ref [ ]; P/F/Pending. | Required: approved connector standard applicability, qualified gauge/test records, part traceability. Closure: Open - Technical Input / Evidence. |
| GAP-017 [P1] — Fluid filter/particle-control route not evidenced The in-process record includes an air-vent filter fitting row but no identifiable fluid-filter check or traceable route to actual particle or fluid-filter efficiency verification. These are distinct parts/functions. | ISO 8536-4:2010 Clauses 6.1 and 6.7, PDF pp.13-14 / printed pp.5-6, and Annexes A.1/A.5, PDF pp.18-21 / printed pp.10-13, address particle contamination and at least 80% latex retention for standard-set fluid filters. Applicability/limits for this burette product remain pending Part 5. | WI p.2 describes "Fluid Filter" fitted in chamber outlet and free of spots/fibers/damage. WI p.3 tubing clean; no executed analytical particulate/filter qualification record supplied. | [Fluid pathway/filter] Identify and inspect the fluid-filter component (different from the air-vent filter). Record installed presence, correct positioning, integrity and absence of visual debris. Define analytical particulate and retention verification using approved product standards, qualified supplier/engineering evidence and frequency; record link to actual test results. | F-823-i75 p.1 Row 6 is "Air vent filter status"; no separate "Fluid filter" row among 18 characteristics. | Add "Fluid filter - presence, location, integrity and visual debris" with P/F and component ID; evidence link for particulate/filter retention qualification [report ID, applicable acceptance, approved control route]. Do not require 10-set particle test at every hourly inspection. | Required: filter BOM/spec, appropriate ISO 8536-5/part4 applicability, approved supplier or qualification record and executed visual checks. Closure: Open - Technical Input / Evidence. |
| GAP-018 [P2] — Linked packaging, sterilization and release records unverified The WI references separate packaging and final-release activities; those forms and associated sterilization/lot evidence were not provided. This is an evidence-verification gap, not proof these activities are absent or that they must be copied into the hourly record. | ISO 8536-4:2010 Clauses 8.2, 9 and 10, PDF pp.15-17 / printed pp.7-9, discuss validated sterilization, labeling and packaging for sets in scope. Part 5 and applicable regulatory/approved specification controls must be assessed before asserting any product-specific requirements. | WI p.10 Section 7 packaging/labeling includes 2-hour packed quantity sampling; p.11 Sections 8-9 final inspection/release and references to F-823-T38, F-823-i75, F-823-i4 and F-823-i15. | [Record linkage] Retain separate approved packaging, labeling, sterilization and final release records. The in-process dossier shall contain record IDs, lot numbers, revisions and status for each applicable route. QC/QA shall verify documented completion before actual release, under the existing authorization procedure. | F-823-i75 p.1 contains no packaging/sterilization/release record references; this is expected if covered by controlled alternatives but their existence is unverified. | Document references only: packaging form [ ]; final inspection [ ]; sterilization batch/release [ ]; label verification [ ]; status Pending/Linked; receiving record [ ]; no independent release decision. | Required: actual F-823-T38/i4/i15 controlled copies, approved sterilization and release records, lot linkage. None supplied; Open - Evidence. |
| GAP-019 [P1] — Failure containment and disposition trail not captured The form permits X (Not Acceptable) with remarks but has no mandatory lot containment, escalation, nonconformance identifier, recheck/disposition or authorization trail; an unacceptable in-process result could remain unaddressed in the record. | This is a quality-system control gap identified from the actual form/WI, not a claim that ISO 8536-4 prescribes a specific NCR form or a particular QMS clause. | WI p.11 Section 8 references QC inspection/release to QCM/QAM; no step linking an in-process X entry to batch hold, investigation and authorized disposition is visible in the provided WI. | [Nonconforming in-process results] On Fail, identify affected work-order/lot, isolate or place material on hold under approved nonconforming-product procedure, notify authorized QC/Production/QA personnel, record NCR/deviation ID and disposition authorization. Record any re-inspection with date, result and link. No in-process form approval overrides required release controls. | F-823-i75 p.1 legend "X = Not Acceptable" and free-text Remarks; "Checked by Q.C." and "Reviewed by Q.C.M." without NCR/hold/disposition fields. | For each Fail: discrepancy/defect [ ]; impacted lot and quantity [ ]; containment action / time [ ]; NCR ID [ ]; disposition authority [ ]; correction/retest ID [ ]; reviewed-by and date [ ]; close only after approved evidence. | Required: approved nonconforming process, sample executed NCR, hold/release evidence and review; actual evidence absent. Open - Approval / Evidence. |

# 6. Verification strategy - keep tests in correct route

| Verification route | Examples from sources | Implementation and evidence rule |
| --- | --- | --- |
| Routine in-process | Component visual condition, fitting, kink, assembly and clamp observations; dimensional checks where control plan requires | Log actual timestamp/sample/results; use controlled product acceptance specs; do not stamp pass without observation. |
| Incoming/supplier | Spike/fluid filter/cap/connector design conformity; incoming material documentation | Reference approved purchase/specification and supplier qualification records for affected parts. |
| Qualification / type / periodic | Burette graduation volume accuracy, drip factor, air-vent flow, infusion throughput, tensile, leak performance, particulate/filter retention | Perform on a frequency supported by product-specific Part 5 and approved risk/control plan. Store complete actual test results. |
| Sterilization/packaging/release | Validated sterilization; approved labeling and pack record; release authorization | Keep in distinct controlled forms with traceable IDs; do not require all test details on in-process hourly form. |

# 7. QC Manager corrective action plan

| Gap ID | Priority | Required action | Function | Objective closure evidence | Status |
| --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 | Incorrect product-specific ISO reference — implement controlled revision and confirm acceptance basis. | QA / Regulatory / Engineering | controlled ISO 8536-5 copy, approved applicability/edition assessment, configuration matrix and authorized reference update. Actual evidence: NOT PROVIDED. | OPEN |
| GAP-002 | P2 | WI and record control-stage ambiguity — implement controlled revision and confirm acceptance basis. | QC / QA / Document Control | approved control-plan flow, controlled form list, linked completed records. Actual evidence: other referenced records not supplied. | OPEN |
| GAP-003 | P2 | Sampling, frequency and traceability not defined — implement controlled revision and confirm acceptance basis. | QC / Production / QA | approved control plan and sampling rationale, executed checkpoint sheets, machine/lot traceability, equipment list. Actual: NOT PROVIDED. | OPEN |
| GAP-004 | P1 | Burette scale accuracy not verifiable — implement controlled revision and confirm acceptance basis. | Engineering / QC / QA | applicable Part 5 clauses, approved drawing/tolerance, validated method, calibrated volume equipment, completed scale accuracy results. Actual: NONE. | OPEN |
| GAP-005 | P2 | Burette assembly/hanging features lack defined criteria — implement controlled revision and confirm acceptance basis. | Engineering / QC | approved component drawing, assembly acceptance catalog, any qualified hanger/valve retention data, sampled observations. | OPEN |
| GAP-006 | P1 | Air vent filter verification insufficient — implement controlled revision and confirm acceptance basis. | Engineering / Supplier Quality / QC | selected standard, approved vent design, filter qualification evidence and traceable assembly check. Actual NOT PROVIDED. | OPEN |
| GAP-007 | P2 | Spike performance not defined beyond appearance — implement controlled revision and confirm acceptance basis. | Supplier Quality / QC | component drawing, supplier report/qualification or actual piercing data if applicable, linked in-process inspection. | OPEN |
| GAP-008 | P1 | Tubing length and dimension conflict — implement controlled revision and confirm acceptance basis. | Engineering / QC / QA | controlled drawing, current product configuration, ISO 8536-5 clause, documented disposition for 100 cm option, calibrated measurement records. | OPEN |
| GAP-009 | P2 | Subjective visual and surface cleanliness terms — implement controlled revision and confirm acceptance basis. | QC / QA | controlled visual defect standard and executed checks; separate particle-test route if applicable. | OPEN |
| GAP-010 | P1 | Joining integrity/tensile result not documented — implement controlled revision and confirm acceptance basis. | QC / Engineering / QA | approved tensile/control plan, calibrated force gauge, actual test record or qualified linked evidence. | OPEN |
| GAP-011 | P1 | Leak-test method and recording gap — implement controlled revision and confirm acceptance basis. | QC Lab / Engineering / QA | confirmed Part 5 approach, approved leakage protocol, calibrated equipment, executed test data, linked NCR if fail. | OPEN |
| GAP-012 | P1 | Injection valve/bulb closure integrity underspecified — implement controlled revision and confirm acceptance basis. | QC / Engineering | approved port-specific test basis, method and actual reseal/leakage result, traceability. | OPEN |
| GAP-013 | P2 | Clamp shutoff and regulation acceptance unclear — implement controlled revision and confirm acceptance basis. | QC / Production | clamp functional test method, approved acceptance, QC data. | OPEN |
| GAP-014 | P1 | Drop factor contradictory and mislabeled as flow rate — implement controlled revision and confirm acceptance basis. | Engineering / QC / QA | Part 5 full text, approved product drop factor, revised specification, actual calibrated drop measurement, approved discrepancy disposition. | OPEN |
| GAP-015 | P1 | Infusion throughput not assessed by drop-factor row — implement controlled revision and confirm acceptance basis. | Engineering / QC Lab | ISO 8536-5 flow method/limit and engineering/risk route, executed flow test or approved linked qualification. | OPEN |
| GAP-016 | P2 | Luer connector interface verification weak — implement controlled revision and confirm acceptance basis. | Engineering / QC / Supplier Quality | approved connector standard applicability, qualified gauge/test records, part traceability. | OPEN |
| GAP-017 | P1 | Fluid filter/particle-control route not evidenced — implement controlled revision and confirm acceptance basis. | Engineering / Supplier Quality / QC | filter BOM/spec, appropriate ISO 8536-5/part4 applicability, approved supplier or qualification record and executed visual checks. | OPEN |
| GAP-018 | P2 | Linked packaging, sterilization and release records unverified — implement controlled revision and confirm acceptance basis. | QC / QA / Document Control | actual F-823-T38/i4/i15 controlled copies, approved sterilization and release records, lot linkage. None supplied; Open - Evidence. | OPEN |
| GAP-019 | P1 | Failure containment and disposition trail not captured — implement controlled revision and confirm acceptance basis. | QC / QA / Production | approved nonconforming process, sample executed NCR, hold/release evidence and review; actual evidence absent. Open - Approval / Evidence. | OPEN |

# 8. Approval gates / release decision

Gate A - QA/Engineering verifies correct edition and obtains the controlled ISO 8536-5 full text (and approved application of any Part 4 requirements).

Gate B - QC Manager approves product configurations, nominal drop factor, dimensions/tolerances, graduation checks and control plan/frequencies.

Gate C - Document Control issues approved WI and the F-823-i75 revision only after review; old copies handled by approved change control.

Gate D - At least one complete controlled production execution provides traceable actual measurements/checks, documented failures/disposition where applicable, and all external test/release linkages.

No verified ISO conformity, 100% documentary coverage or lot release conclusion is possible with current inputs.

# 9. Source traceability

WI: photographed PDF "Wi 2.pdf": user-visible WI printed page numbers 1-11 (PDF order reversed: PDF p.11 = WI p.1; PDF p.1 = WI p.11). Identity: WI-8-2-2-30-1 Rev.0.

In-process record: IMG_2903(1).jpeg, F-823-i75 Rev.2 p.1, visual Rows 1-18; original form is blank.

BS EN ISO 8536-4:2010 supplied PDF: Foreword printed iv/PDF p.8; printed ISO pp.1-17/PDF pp.9-25; normative Annex A printed pp.10-13/PDF pp.18-21.

ISO catalog public metadata (checked 15 Oct 2026): https://www.iso.org/standard/20921.html (ISO 8536-5:2004); https://www.iso.org/standard/88576.html (ISO/FDIS 8536-5 in draft, not approved as a published standard).

The full ISO 8536-5 controlled text, all supplier/operational records and approved drawings were not provided. Neither clause numbers nor acceptance limits for Part 5 have been invented.