# IN-PROCESS CONTROL RECORD — PROPOSED

F-823-i75 | Infusion Burette Set | Working draft; replaces neither controlled form nor Final Inspection

Review status: CONTROLLED PROPOSAL - pending technical verification, QC Manager/QA approval and operational evidence.

| Product | Infusion Burette Set / IV Burette Set |
| --- | --- |
| Scope | In-process manufacturing controls (not a final inspection release report) |
| Original WI | WI-8-2-2-30-1, Rev. 0 (13 Jan 2024) |
| Original in-process record | F-823-i75, Rev. 2 (01 Mar 2024) |
| Supplied ISO | BS EN ISO 8536-4:2010 |
| Product-specific ISO needed | ISO 8536-5:2004 (full controlled text NOT provided; applicability requires approval) |
| Assessment date | 15 October 2026 |
| Reviewed source scope | 11-image-page WI, 28-page ISO file, 1-page photographed form |

Document owner/reviewer/signatures: NOT ASSIGNED - proposed working draft. No original or external document has been modified.

# 1. Identification and execution instructions

| Traceability input | Entry |
| --- | --- |
| Product, model and nominal burette capacity | _______________________________________________ |
| Batch/Lot / WO / PO / line & machine | _______________________________________________ |
| Date / shift / inspected production stage | _______________________________________________ |
| Approved WI + revision / control plan ID+revision | _______________________________________________ |
| Approved product drawing / equipment calibration list | _______________________________________________ |
| Applicable product standard + edition + approved matrix ref | _______________________________________________ |
| Inspector / reviewer | _______________________________________________ |

Record actual results; PASS only if the approved acceptance criterion was checked. FAIL requires containment/NCR; N/A must be justified; PENDING identifies unavailable technical criteria or test records. The frequency and sample size are set only by the current approved control plan. A blank is not PASS.

# 2. Routine in-process check sheet — original 18 characteristics preserved

| No. / Original characteristic | Approved acceptance / open criterion | Check method | Actual result / observation (enter value when measured) | Sample ID & qty / equipment | P / F / N/A / Pending | Evidence / NCR / remarks |
| --- | --- | --- | --- | --- | --- | --- |
| 1. Clarity, Appearance GAP-009 | Tube transparent; good appearance | Visual | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 2. Free from irregularities GAP-009 | No flashing; no damage; no black spots | Visual | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 3. Tube Clarity GAP-009 | Clear, flexible; kink resistance; no bending | Visual/functional | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 4. Spike Status GAP-007 | Sharp tip, open tip; no blockage; protector fit | Visual | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 5. Clamp status GAP-013 | No leakage after clamping; flexible | Visual/functional | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 6. Air vent filter status GAP-006 | Vent filter/cap correctly installed; performance requires qualified evidence if applicable. | Visual + evidence | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 7. Injection valve GAP-012 | Rubber valve good fitting; no leakage | Visual + evidence | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 8. Chamber Quality GAP-005 | Straight, soft chamber; no leakage | Visual | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 9. Chamber Hanging Strip GAP-005 / GAP-010 | Strong fitting; no pullout; flexible | Visual/retention | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 10. Chamber graduation scale GAP-004 | Print legibility; graduated-scale accuracy per approved Part 5/drawing (OPEN tolerance). | Visual/measurement | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 11. Tube Dimension GAP-008 | ID/OD 3.0 x 4.1 mm nominal from original record; tolerance & model: OPEN. | Calibrated measurement | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 12. Tube Length GAP-008 | 100/150 cm source options conflict with WI >=1500 mm; do not auto-PASS; approved model variant required. | Calibrated measurement | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 13. Tube Gluing status GAP-010 | Strong; no pullout at all gluing areas | Visual/qualified retention | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 14. Roller Clamp GAP-013 | Strong clamping; no leakage | Visual/functional | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 15. Injection bulb status GAP-012 | Good fitting; no leakage | Visual/functional | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 16. Male Connector GAP-016 | Good Luer fitting; no flashing; tight cap | Visual/linked evidence | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 17. Flow Rate (original label) GAP-014 / GAP-015 | Do NOT use 10 drops/1 ml as final criterion; approved drop factor and test method PENDING. | See separate test route | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |
| 18. Leak Test GAP-011 | No leak observed under approved method (pressure/temp/time/equipment OPEN). | Approved method/evidence | ____________________ ____________________ | ID:_________ Qty:______ Equipment:_________ | [ ] P [ ] F [ ] N/A [ ] Pending | ____________________ ____________________ |

# 3. Supplemental required characteristic and external evidence links

| New check / linking field | Actual entry / method / applicable evidence | Status |
| --- | --- | --- |
| Fluid filter (distinct from vent filter) | Component presence/position/condition, BOM ID and qualification/filter performance report: ________________________ | P / F / Pending |
| Burette scale verification | Scale point [ml], indicated [ml], actual [ml], deviation [ml/%], tolerance ref, calibrated tool: __________________ | P / F / Pending |
| Drop-factor qualification/periodic | Approved target [drops/ml], count, delivered water volume, temperature, drop rate and report ID: ____________ | P / F / Pending |
| True fluid throughput (separate from drop factor) | Method, fluid, head [m], delivered ml in min, limit, linked test ID: __________________ | P / F / Pending |
| Assembly leak / joint/tensile records | Pressure/force, exposure duration, temperature, equipment, tested lot/unit, approved independent worksheet: ____________ | P / F / Pending |
| Packaging / sterilization / release linkage | F-823-i4, F-823-i15, F-823-T38 and approved controlled lot release evidence ID: ______________ | Linked / Pending |

# 4. Sampling / timepoint event log — clock times are not prescribed frequency

| Actual time | Step / characteristic IDs | Sample identification | Qty | Result and record ID | Inspector initials | Line/lot / NCR |
| --- | --- | --- | --- | --- | --- | --- |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |
| ____________ | ___________________ | __________________ | ______ | ___________________ | ________ | ______________ |

# 5. Nonconformance / disposition / review

| Required item | Completed value |
| --- | --- |
| Failed characteristic / actual failure / NCR ID | ___________________________________________________ |
| Affected lot/unit quantity / date-time of containment | ___________________________________________________ |
| Hold location / notified function | ___________________________________________________ |
| Approved disposition and re-inspection reference | ___________________________________________________ |
| Checked by QC / date | ___________________________________________________ |
| Reviewed by QCM / date | ___________________________________________________ |
| Final release decision | NOT A RELEASE RECORD; link separate approved release document ID: ____________ |

The source F-823-i75 contains an hour-by-hour grid (7, 8, 9, 10, 11, 12, 1, 2, 3, 4, 5, 6). This proposal replaces the implicit hourly checkboxes with event-based records until an approved checkpoint frequency is provided. Document Control must decide whether to preserve fixed-hour fields in the authorized future revision.