# Infusion Set | ISO Closure Register

Requirement-by-requirement coverage and truthful closure status

DRAFT FOR QC/QA TECHNICAL REVIEW - NOT APPROVED; NO VERIFIED OPERATIONAL CLOSURE

Primary standard:  BS EN ISO 8536-4:2010

Assessment documents:  WI-8-2-2-20 (Rev. 4, 01 May 2022) / F-823-i4 (Rev. 19, 22 Sep 2026)

Output type:  Controlled DRAFT register; contains no executed release or test data

Audit boundary: process control during manufacture and packaging is distinct from incoming, final QC testing, qualification and release. A missing test on F-823-i4 does not prove ISO nonconformance where another controlled record is appropriate.

# 1. Mandatory Closure Gate

Proposed coverage is NOT approved and is NOT verified operational conformity. Evidence Reference = "Not supplied" for each requirement as supplied; missing F-823-T38, F-8-2-4 and sterilization records prevent an evidence-based closure conclusion. Product claims and applicable variants require review before final release. Valid alternative evidence routes may satisfy requirements without putting each test on the packaging in-process form.

# 2. ISO Applicability & Closure Matrix

| Requirement ID / ISO Clause | ISO Says / Source | Related Gap(s) | Existing / Revised Documentary Coverage | Applicability | Evidence Route | Actual Evidence | Closure Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-001 1 | PDF 9, ISO 1 Gravity-fed single-use infusion sets; verify product configuration | GAP-008 | Configuration under review; proposed WI draft addresses linkage | Review applicable for scoped set | Approved product specification / BOM / design records | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-002 3.1 | PDF 9, ISO 1 Vented / unvented configurations and container compatibility | GAP-008 | Needs configuration evidence; proposed WI draft addresses linkage | Applicable per variant - confirmation OPEN | Approved product specification / BOM / design records | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-003 3.2 | PDF 10, ISO 2 Protect fluid-path components with suitable caps | GAP-013 | WI partially covered; proposed WI draft addresses linkage | Review applicable for scoped set | Packaging in-process + approved release controls | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-004 4.1 | PDF 12, ISO 4 Designation of conforming gravity-feed set IS-G | GAP-020 | Applicability / product label not proven; proposed WI draft addresses linkage | Review applicable for scoped set | Approved product specification / BOM / design records | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-005 4.2 | PDF 12, ISO 4 Designation of compliant separate air inlet IS-AD, where applicable | GAP-008 | Conditional / variant to confirm; proposed WI draft addresses linkage | Applicable per variant - confirmation OPEN | Approved product specification / BOM / design records | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-006 5 | PDF 12, ISO 4 Component material requirements of clauses 6-8 as applicable | GAP-024; GAP-014; GAP-015 | Insufficient WI linkage; proposed WI draft addresses linkage | Review applicable for scoped set | Approved product specification / BOM / design records | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-007 6.1 | PDF 13, ISO 5 Clean fluid path, Annex A.1 index limit | GAP-003 | Visual only; assay not evidenced; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-008 6.2 | PDF 13, ISO 5 No leakage per Annex A.2 | GAP-001 | Incomplete WI test; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-009 6.3 | PDF 13, ISO 5 15 N for 15 s axial, excluding caps | GAP-002 | Wrong Annex in WI; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-010 6.4 | PDF 13, ISO 5 Fig 4 dimensional conformity and puncture capability | GAP-009 | No measurable WI check; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-011 6.5 | PDF 13, ISO 5 Microbial air filtration, ≤20% flow impact if applicable | GAP-008 | Partly visual; no qualified record; proposed WI draft addresses linkage | Applicable per variant - confirmation OPEN | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-012 6.6 | PDF 14, ISO 6 Clear visual air/water interface; functional length ≥1500 mm | GAP-007 | As per spec only; proposed WI draft addresses linkage | Review applicable for scoped set | Approved in-process / finished QC plan | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-013 6.7 | PDF 14, ISO 6 Filter installed, ≥80% latex-particle retention A.5 | GAP-004 | Not documented in WI; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-014 6.8 | PDF 14, ISO 6 Chamber geometry and correct 20/60-drop calibration | GAP-006 | Not documented; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-015 6.9 | PDF 14, ISO 6 Adjust zero-to-max flow | GAP-005 | Partial roller-clamp visual; proposed WI draft addresses linkage | Review applicable for scoped set | Approved in-process / finished QC plan | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-016 6.10 | PDF 14, ISO 6 ≥1000 ml / 10 min, 9 g/l NaCl, 1m head | GAP-005 | No corresponding WI limit; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-017 6.11 | PDF 14, ISO 6 Reseals by A.6 when present | GAP-012 | No defined reseal method; proposed WI draft addresses linkage | Applicable per variant - confirmation OPEN | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-018 6.12 | PDF 14, ISO 6 Conforms ISO 594-1 or -2 | GAP-011 | Visual mating only; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-019 6.13 | PDF 14, ISO 6 Cap maintains sterile fluid pathway | GAP-013 | Partial / not explicit; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-020 7.1 | PDF 15, ISO 7 ≤2.0 ml 0.005 mol/l Na2S2O3 difference | GAP-014 | Separate lab evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-021 7.2 | PDF 15, ISO 7 Ba+Cr+Cu+Pb+Sn ≤1 µg/ml total; Cd ≤0.1 µg/ml and Pb colour comparator | GAP-014 | Separate lab evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-022 7.3 | PDF 15, ISO 7 ≤1 ml standard titrant | GAP-014 | Separate lab evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-023 7.4 | PDF 15, ISO 7 ≤5 mg | GAP-014 | Separate lab evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-024 7.5 | PDF 15, ISO 7 ≤0.1 | GAP-014 | Separate lab evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-025 8.1 | PDF 15, ISO 7 Biological compatibility evaluated | GAP-015 | Evidence not supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-026 8.2 | PDF 15, ISO 7 Validated process for unit-container sterile product | GAP-016 | Referenced process; no validation evidence; proposed WI draft addresses linkage | Review applicable for scoped set | Sterilization / lab + final release record | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-027 8.3 | PDF 15, ISO 7 Suitable assessment shows freedom from pyrogens | GAP-017 | Evidence not supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Sterilization / lab + final release record | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-028 8.4 | PDF 15, ISO 7 Freedom from haemolytic constituents | GAP-015 | Evidence not supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-029 8.5 | PDF 16, ISO 8 Suitable assessment shows freedom from toxicity | GAP-015 | Evidence not supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-030 9.1 | PDF 16, ISO 8 All a-j required fields as applicable | GAP-020; GAP-010; GAP-017 | Partial WI/form; proposed WI draft addresses linkage | Review applicable for scoped set | Packaging in-process + approved release controls | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-031 9.2 | PDF 16, ISO 8 All a-g on shelf or multiunit pack when used | GAP-021 | General WI only; proposed WI draft addresses linkage | Review applicable for scoped set | Packaging in-process + approved release controls | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-032 10.1 | PDF 17, ISO 9 Individual packs maintain sterility, tamper evidence | GAP-013; GAP-018; GAP-019; GAP-023 | Unverified seal process/control; proposed WI draft addresses linkage | Review applicable for scoped set | Packaging in-process + approved release controls | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-033 10.2 | PDF 17, ISO 9 No flattened or kinked portions ready for use | GAP-022 | Generic deformity check only; proposed WI draft addresses linkage | Review applicable for scoped set | Packaging in-process + approved release controls | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-034 A.1 | PDF 18, ISO 10 A.1 procedure and blank/count calculation | GAP-003 | No test worksheet supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-035 A.2 | PDF 19, ISO 11 Positive+negative tests at 40±1°C | GAP-001 | WI incomplete; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-036 A.3 | PDF 19, ISO 11 15 N axial for 15 s | GAP-002 | WI incorrect cross-reference; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-037 A.4 | PDF 20, ISO 12 Air-inlet performance relative to unrestricted vent | GAP-008 | Separate evidence not supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-038 A.5 | PDF 20, ISO 12 Particle retention in duplicate; percentage calculation | GAP-004 | Evidence not supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-039 A.6 | PDF 21, ISO 13 Needle/pressure/duration water leak observation | GAP-012 | Conditional evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-040 B.1-B.6 | PDF 22, ISO 14-15 Extract preparation and 5 category tests | GAP-014 | No chemical files supplied; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |
| REQ-041 C.1-C.2 | PDF 24, ISO 16 Pyrogenicity and biological evaluation test guidance | GAP-015; GAP-017 | Evidence absent; proposed WI draft addresses linkage | Review applicable for scoped set | Qualification / lab / controlled design verification | NOT SUPPLIED | Proposed draft only; OPEN - approval / technical evidence |

# 3. OPEN Technical Inputs

Current standard/regulatory applicability and product/needle/vented configuration matrix

Approved BOM/drawings/PS and numerical ID/OD/design specs, labels and needle dimensions

Packaging machine/material validated recipe (ordinary/PFM), temperature, dwell and width windows

Validated package peel method, actual acceptance and sampling plan (not defined in ISO 8536-4:2010)

Approved in-process sample size, startup/changeover/periodic frequency and action limits

ISO physical testing frequency, qualification/periodic/supplier verification rationale

Chemical/biological/pyrogenicity reports, method selection and scope to all material/configuration variants

Complete F-823-T38, F-8-2-4, WI-7-5-1-1, sterilization qualification and execution records

# 4. Review / Authorization

QC Manager decision: ________    QA authorization: ________    Engineering package review: ________    Document Control effective revision: ________

No criterion may be marked "Verified Closed" until relevant objective evidence has been located, checked for validity, linked to product scope and formally approved. Evidence absent means OPEN, not proof of test failure.