# Infusion Set - Revised In-Process Control Record

Suggested controlled replacement/annex to F-823-i4 (NOT F-823-T38)

DRAFT FOR QC/QA TECHNICAL REVIEW - NOT APPROVED; NO VERIFIED OPERATIONAL CLOSURE

Baseline:  F-823-i4 (Rev. 19, 22 Sep 2026)

Proposed ID:  F-823-i4 [next revision / appendix ID to be assigned]

Scope:  Packaging Machine & Final Checking - at-line production

Status:  DRAFT - generic company values pending engineering and QC approval

Audit boundary: process control during manufacture and packaging is distinct from incoming, final QC testing, qualification and release. A missing test on F-823-i4 does not prove ISO nonconformance where another controlled record is appropriate.

# A. Line / Product Identification & Approved Setup

| Field | Entry | Field | Entry |
| --- | --- | --- | --- |
| Product name / product code | ________________________ | Controlled specification / revision | ________________________ |
| Lot / production date / shift | ________________________ | Previous product / previous lot | ________________________ |
| Machine / line / operator | ________________________ | Pack material lot / BOM ID / rev | ________________________ |
| With needle? / gauge | Y / N ____ | Vented? / injection site? | Y / N / Y / N |
| Start-up / changeover / adjustment reason | ________________________ | Approved packaging validation / recipe ID | OPEN TECHNICAL INPUT REQUIRED |
| Control plan ID / revision | OPEN TECHNICAL INPUT REQUIRED | Sample size / check timing | OPEN TECHNICAL INPUT REQUIRED |
| Sealing path | STANDARD / PFM / other | Validated width / dwell / temp limit | OPEN TECHNICAL INPUT REQUIRED |

# B. Time-Point Packaging Process Measurements

Record an ACTUAL result for each applicable measured parameter. Do not apply both sealing-temperature bands to one recipe unless engineering validation explicitly requires it. The old form printed 4-12 mm width; 100-150°C ordinary temperature; 2-5 sec dwell; 250-300°C for PFM. Those are unverified company defaults, not ISO numerical requirements.

| Date/Time | Machine / recipe rev | Sample ID / n | Seal width, mm actual / pass? | Seal temp, °C actual / pass? | Dwell, s actual / pass? | Peel/integrity actual & method | Decision / inspector |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |

# C. Product-Specific In-Process Controls

Repeat Section C for EACH required check event or sample set, referencing the Date/Time and Sample ID in Section B. Each checklist must be completed and reviewed per the approved control plan. Date/Time: __________  Sample ID(s): __________  Inspector: __________

| Check | In-process characteristic | Approved requirement / acceptance | Actual observation / measurement | Sample n / defect n | P/F/N/A | Remarks / spec / evidence |
| --- | --- | --- | --- | --- | --- | --- |
| 01 | Clarity and appearance | No visible spots/particles/contamination; approved sample acceptance plan | ________________ | _______ / ______ | _____ | F-823-i4 row 1; GAP-003 |
| 02 | Cutting status | Smooth edges, no sharp edges or fibers; approved sample plan | ________________ | _______ / ______ | _____ | Row 2 |
| 03 | Plastic forming status | No hole, burn mark, unacceptable forming/deformation | ________________ | _______ / ______ | _____ | Row 3 |
| 04 | Peel / seal integrity | No gaps, burns, fibers; validated inspection method and criterion [INPUT] | ________________ | _______ / ______ | _____ | Row 4; GAP-019 |
| 05 | Unit pack label / printing | Approved artwork; ALL ISO 9.1 applicable elements including gravity-feed, sterile, LOT, expiry, manufacturer, single-use, IFU and drop factor | ________________ | _______ / ______ | _____ | Row 5; GAP-020 |
| 06 | Outer box label if used | All ISO 9.2 applicable fields; count and LOT/expiry match | ________________ | _______ / ______ | _____ | GAP-021 |
| 07 | Product status | No damage/flash/pullout/broken/kink/flattening; correct components | ________________ | _______ / ______ | _____ | Row 6; GAP-022 |
| 08 | Components / caps | BOM-linked filter, caps, regulator, conical connector, optional vent/Y-site and needle/gauge if fitted | ________________ | _______ / ______ | _____ | Row 7; GAP-004; GAP-008; GAP-010; GAP-013; GAP-024 |
| 09 | Tubing ID / OD (mm) | Actual readings against controlled product spec [INPUT] | ________________ | _______ / ______ | _____ | Row 8; GAP-007; GAP-024 |
| 10 | Tubing ISO-defined length (mm) | ≥1500 mm plus controlled design nominal/tolerance where applicable | ________________ | _______ / ______ | _____ | Row 9; GAP-007 |
| 11 | Graduation / marking, if applicable | Clear/readable, matches approved design / marking plan | ________________ | _______ / ______ | _____ | Row 10 (N/A permitted) |
| 12 | Line clearance, product/lot change | Prior-product remnants cleared and correct current labels/components loaded | ________________ | _______ / ______ | _____ | GAP-023; GAP-024 |

# D. Failure, Containment, and Restart

| Failed check/time | Suspect output range / quantity | NCR/hold ID | Root cause / corrective adjustment | Recheck evidence | Authorized restart / initials |
| --- | --- | --- | --- | --- | --- |
| ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ |

# E. Required Signoff / Evidence References

| QC first shift / date-time | QC second shift / date-time | QC Supervisor reviewer / date-time | Controlled drawing/recipe IDs | Referenced batch / final records |
| --- | --- | --- | --- | --- |
| ________________ | ________________ | ________________ | ________________ | ________________ |

THIS FORM DOES NOT contain routine ISO 8536-4 Annex A/B/C qualification assays, chemical / biological evaluation, sterilization records or final product release. These belong to approved F-823-T38, F-8-2-4, laboratory/validation/technical-file routes (subject to company assignment), each with traceable objective evidence.

OPEN TECHNICAL INPUT REQUIRED: validated sealing limits, peel method/acceptance, sample sizes and frequency, product-specific dimensions/tolerances, approved spec IDs and applicable company signoff. No Pass/Fail may be assumed until entries are completed.