# OXYGEN TUBE | EVIDENCE & CLOSURE REGISTER

In-Process gap traceability | WI page 3 included | ISO not assessed

EVIDENCE-BASED REASSESSMENT | IN-PROCESS, NOT FINAL INSPECTION | ISO NOT ASSESSED

| Field | Verified / current assessment |
| --- | --- |
| Source review scope | WI-8-2-2-33-1 Rev 07 pp.1–6; Oxygen Tube form F-823-138-1 Rev 3 |
| ISO | Excluded because no standard is supplied; no ISO-based closure is claimed |
| Closure rule | Only approved methods/limits and verified executed records demonstrate operational control |
| Reassessment | 6 open issues / 2 verification items / 4 covered WI controls; zero operationally verified closures |

# A. Open gap closure register

| Gap | Priority | Characteristic / requirement | Evidence necessary to close | Documentary status | Actually verified evidence | Operational status |
| --- | --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 (provisional) | Approved dimensional limits and specification revisions not verified | Approved TF-33-PS and TF-33-PDBM, confirmed effective revisions, applicable dimensional limit matrix, reviewed check records. | Documentary proposed / approval pending | No relevant executed record or approved technical inputs verified | OPEN — Evidence / approval |
| GAP-002 | P2 | Numerical in-process measurement evidence is not captured on the displayed form | Executed numeric log OR demonstrated approved alternative record, source limits, method and calibrated equipment ID. | Documentary proposed / approval pending | No relevant executed record or approved technical inputs verified | OPEN — Evidence / approval |
| GAP-003 | P2 | Subjective visual and kink/retention acceptance methods require verification | Approved visual standard/approved samples, approved flexibility and retention methods if applicable, verified results and training/qualification as required. | Documentary proposed / approval pending | No relevant executed record or approved technical inputs verified | OPEN — Evidence / approval |
| GAP-004 | P2 | Two-hour checks require stronger time, sample and shift traceability | Approved sampling plan, training, time-stamped execution records across overnight shifts, exception records when applicable. | Documentary proposed / approval pending | No relevant executed record or approved technical inputs verified | OPEN — Evidence / approval |
| GAP-005 | P2 | Equipment ID and calibration linkage not visible | Instrument register and current status, equipment ID on measurement records, issue disposition where relevant. | Documentary proposed / approval pending | No relevant executed record or approved technical inputs verified | OPEN — Evidence / approval |
| GAP-006 | P2 | Failed/missed in-process check handling and disposition not evidenced | Approved nonconforming product procedure, linked failed-event and hold/NCR examples if any, disposition and recheck evidence. | Documentary proposed / approval pending | No relevant executed record or approved technical inputs verified | OPEN — Evidence / approval |

# B. Covered controls — do not reopen as “missing WI”

| Control | Source characteristic | Source | Evidence still needing verification |
| --- | --- | --- | --- |
| COV-01 | Tube ID, OD, and length listed as inspection characteristics | WI p.3 §§8.1–8.1.1; form rows 3–4 | Documentary control exists; numeric limit/evidence verification remains open. |
| COV-02 | Universal connector surface, fitting and ID/OD | WI p.3 §7.4; form rows 5–6 | Documentary control exists; criteria and measurement details to verify. |
| COV-03 | Tube surface appearance and typical defects | WI p.3 §8.1.1; form rows 1–2 | Documentary control exists; specific acceptance guidance to verify. |
| COV-04 | Two-hour tubing check and recording | WI p.3 §8.1.1; form time grid | Written frequency exists; event traceability and real execution not verified. |

# C. Verification items (not gaps)

| Reference | Issue | Open evidence request |
| --- | --- | --- |
| VER-001 | In-Process / WI document scope | Form header cites WI-8-2-2-33-1. WI p.3 §8.1–8.1.1 explicitly covers tubing, and WI p.5 §10.4 covers the glued joint. The previously alleged complete absence of a tube WI link is NOT supported. Check current approved form matrix and applicability; treat as verification, not a confirmed gap. |
| VER-002 | Sampling scheme / external forms | WI p.3 §8.1 refers to AQL general visual and special physical inspection; WI pp.5–6 list additional in-process forms. Applicable AQL tables, inspection level/sample size, current forms and batch evidence are not supplied. No sampling defect or absent company record can be confirmed from this submission. |

# D. Source characteristics linked to inspection data

| Original row | Characteristic | WI source / confirmation | Current open item |
| --- | --- | --- | --- |
| 1 | Appearance | WI p.3 §8.1.1 | GAP-003 |
| 2 | Tube status / kink | Form row 2; WI p.3 §8.1.1 appearance | GAP-003 |
| 3 | Tube length | WI p.3 §8.1–8.1.1 | GAP-001, GAP-002 |
| 4 | Tube ID/OD | WI p.3 §8.1–8.1.1 | GAP-001, GAP-002 |
| 5 | Connector appearance / fit | WI p.3 §7.4 | GAP-003 |
| 6 | Connector ID/OD | WI p.3 §7.4 | GAP-001, GAP-002 |
| 7 | Connector length | Form row 7; approved connector specification unverified | GAP-001, GAP-002 |
| 8 | Glue / connector-tube joint | WI p.5 §10.4 | GAP-003, GAP-006 |
| Frequency | Every-two-hour tubing check | WI p.3 §8.1.1 | GAP-004 |
| Tools | Instrument type check boxes | Original form header | GAP-005 |

# E. Closure and approval rules

“Documentary Covered” means wording has been mapped/proposed in these drafts. It does not mean the revised WI/form has been officially approved or the required tests performed. Mark VERIFIED CLOSED only after checking actual current specifications, approved methods/control plans, controlled issued documents, appropriate measured/observed results and QC/QA review. Draft status remains PENDING for all proposed changes.

QC Manager decision: ______________________ Date: __________  QA technical review: ______________________  Document Control issue/revision: PENDING