# IN-PROCESS CONTROL | EVIDENCE CLOSURE REGISTER

Oxygen Tube / Oxygen Mask document interface • ISO not assessed

SOURCE STATUS  •  IMAGE-BASED REVIEW  |  NOT AN APPROVED CONTROLLED DOCUMENT

| Scope | WI-8-2-2-33-1 Rev 07 (01.06.2024), pp. 1, 2, 4, 5, 6 of 6; Oxygen Tube in-process form F-823-138-1 Rev 3 (29.09.2026), p. 1/1. |
| --- | --- |
| Closure rule | Drafts, proposed limits and empty records are NOT closure evidence. |
| ISO | Not assessed; no standard supplied |
| Status basis | Only photographic source material available; no live executed test or approval evidence supplied. |

# Gap closure and proof needed

| Gap | Priority | Control topic | Closure criteria / required records | Evidence actually verified | Draft documentary state | Operational closure | Owner function |
| --- | --- | --- | --- | --- | --- | --- | --- |
| GAP-001 | P2 | Tube-to-mask process interface not verified | Document Control-approved scope and process map; current parent WI; controlled form-to-WI matrix; traceable tube and assembly batch records. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-002 | P2 | WI page 3 of 6 not supplied | Certified complete WI Rev 07 including page 3; complete technical redline and approval record. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-003 | P2 | Unverified specification identification and acceptance limits | Approved TF-33-PS and TF-33-PDBM (correct effective revisions), limits matrix, authorized inspection plan. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-004 | P1 | Tube length and ID/OD actual measurements not captured | Approved dimensional specification; defined measurement method and sampling; filled measurement records and instrument verification. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-005 | P1 | Connector ID/OD and length actual measurements not captured | Approved connector drawing/tolerances; completed numeric inspection log; suitable calibrated tool evidence. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-006 | P2 | Flexibility / kink-resistance verification undefined | Controlled method and criteria; verified training and completed test record. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-007 | P1 | Glued joint acceptance not objectively defined | Authorized joining process/method/criteria; executed inspection or validation evidence as applicable; records for failed joints. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-008 | P2 | Sampling, per-check identity, and overnight traceability incomplete | Approved sampling/control plan; complete dated per-check traceability; executed records across shift change. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-009 | P2 | Instrument identification and metrology status missing | Equipment register/status, check records, completed instrument IDs and any impact assessment. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-010 | P2 | Visual defect criteria and result rules require control | Approved visual acceptance guide; executed records; review of subjective criteria and training. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-011 | P1 | Failed in-process result response and disposition not evidenced | Approved nonconformance procedure; linked NCR/hold records; verified disposition and reinspection before use. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |
| GAP-012 | P2 | Distinct oxygen-mask assembly checks are not evidenced by tube record | Current approved mask assembly/packaging in-process forms, production traveler/control plan and actual executed records. | Not submitted or not verified | Partial draft / source limitation | Open - evidence / technical input | QC / QA / Eng. / Document Control |

# Source-based characteristic checklist

| Control | Characteristic | Source | Related gap | Evidence / assessment limit |
| --- | --- | --- | --- | --- |
| C-01 | Clarity/appearance | Form row 1 | GAP-010 | Specification/visual guide not verified |
| C-02 | Tube status / kink | Form row 2 | GAP-006 | Test method and limits not verified |
| C-03 | Tube length | Form row 3 | GAP-003,004 | Numerical result and limits not verified |
| C-04 | Tube ID/OD | Form row 4 | GAP-003,004 | Numerical result and limits not verified |
| C-05 | Universal connector appearance / fit | Form row 5 | GAP-010 | Criteria reference not verified |
| C-06 | Connector ID/OD | Form row 6 | GAP-003,005 | Numerical result and limits not verified |
| C-07 | Connector length | Form row 7 | GAP-003,005 | Numerical result and limits not verified |
| C-08 | Connector/tube gluing | Form row 8 | GAP-007 | Reproducible retention method not verified |
| W-01 | Oxygen mask molding/assembly | WI §§7,10 | GAP-001,012 | Separate execution forms not provided |
| W-02 | Reservoir bag assembly | WI §10.3 | GAP-012 | Separate execution forms not provided |
| W-03 | Packaging / final handling | WI §§11-12 | GAP-012 | Separate execution forms not provided |
| W-04 | WI original page 3 | WI p.3/6 | GAP-002 | Page not supplied |

# Sign-off and approval integrity

No gap may be marked VERIFIED CLOSED until its authorized source, control method, applicability and appropriate executed evidence (where required) have been checked. Full original WI, signed/issued controlled specifications and operation records must be reconciled. Proposed drafts must not be used to assert actual ISO or product conformity.

QA / QC Manager approval for final controlled revisions: PENDING. Register review date: ______  Reviewer: ______  Reference to approved actions: ______