# REVISED IN-PROCESS CONTROL | OXYGEN TUBE

Proposed replacement for F-823-138-1 Rev 3 • Unapproved editable form

SOURCE STATUS  •  IMAGE-BASED REVIEW  |  NOT AN APPROVED CONTROLLED DOCUMENT

| Source | F-823-138-1 Rev 3 dated 29.09.2026; Gg.pdf p.1 |
| --- | --- |
| Parent refs (source form) | WI-8-2-2-33-1 / TF-33-PS / TF-33-PDBM ¦ verify exact currently approved revisions |
| Draft status | NOT APPROVED; numerical tolerances, test methods, frequencies and sampling OPEN |
| Record ID | [Controlled number/revision pending Document Control] |

# A. Production and traceability

| Date(s) | Shift | Tube/product variant | Tube length/configuration | Production order | Lot | Line / machine | QC inspector |
| --- | --- | --- | --- | --- | --- | --- | --- |
| __________ | __________ | __________ | __________ | __________ | __________ | __________ | __________ |

| Spec ID & Rev | Tube WI / Rev | Connector spec & Rev | Tube control plan & Rev | Nominal/limits verified by | Verified on date | Route / Mask assembly lot (if applicable) |
| --- | --- | --- | --- | --- | --- | --- |
| ____________ | ____________ | ____________ | ____________ | ____________ | ____________ | ____________ |

# B. Source-based control characteristics - acceptance input required

| No. | Characteristic | Baseline control (source text paraphrase) | Approval / open technical input | Method route | Record result | Gap ID |
| --- | --- | --- | --- | --- | --- | --- |
| 01 | Clarity / appearance | Clear, good appearance; free from irregularities | Visual criteria guide required | Visual | Observed / defect count | GAP-010 |
| 02 | Tube status | Flexible and kink resistant; smooth internal channel; no black spots; PVC per source; appropriate post-winding condition | Approved material/functional criteria & test method | Visual + approved method | Observed + criterion | GAP-006,010 |
| 03 | Tube length | As per approved dimension specification, cm | Approved nominal/tolerance not provided | Caliper/scale per approved method | Actual numeric cm | GAP-003,004 |
| 04 | Tube ID / OD | As per approved dimension specification, mm | Approved ID/OD limits not provided | Approved measuring method | Actual numeric ID, OD mm | GAP-003,004 |
| 05 | Universal connector appearance / fitting | No flashing, damage or black spots; suitable fitting to tube and nebulizer jar | Controlled fit/visual requirement to verify | Visual / approved fit check | Observed / pass-fail | GAP-010 |
| 06 | Connector ID / OD | As per approved dimension specification, mm | Approved ID/OD limits not provided | Approved measuring method | Actual numeric ID, OD mm | GAP-003,005 |
| 07 | Connector length | As per approved dimension specification, mm | Approved length limits not provided | Approved measuring method | Actual numeric mm | GAP-003,005 |
| 08 | Tube / connector glued joint | Tube inserted correctly; joint secure, no pullout and no excess adhesive | Approved retention method and criterion not provided | Visual + approved retention method | Observation & actual test result | GAP-007 |

This draft preserves the eight source checks. No source-derived numeric dimensions, kink/retention methods, acceptance limits or sample quantities were supplied. All such cells must be populated from controlled references BEFORE active use. [GAP-003 to GAP-007]

# C. Dimensional measurement log - actual numeric results

One line per sampled unit/check event. Record the approved limits in section A or attach an approved characteristic-limit matrix. Pass requires all applicable documented limits to be met.

| Date / time | Shift / line | Sample ID | Tube L (cm) | Tube ID (mm) | Tube OD (mm) | Conn. ID (mm) | Conn. OD (mm) | Conn. L (mm) | Equipment ID | Pass / Fail / Pending | Inspector |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |

Controlled limits reference: ___________________  Revision: ______  Sample size per approved plan: ______  Frequency per approved plan: ______

# D. Visual / functional / adhesive control log

Use codes: P = Pass; F = Fail; N/A = documented not applicable; PEN = Pending (missing criterion/evidence). A check mark without supporting measured data is insufficient for dimensional characteristics.

| Date / time | Sample ID / count | Appearance 01 | Tube status 02 | Connector 05 | Glue 08 | Kink method / result | Joint method / result | Defect / NCR ref | Inspector |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |
| ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ | ________ |

# E. Exception, containment and disposition log

| Event / time | Failed feature | Affected interval / quantity | Hold/NCR reference | Action + disposition (approved) | Recheck result + evidence | Authorization |
| --- | --- | --- | --- | --- | --- | --- |
| __________ | __________ | __________ | __________ | __________ | __________ | __________ |
| __________ | __________ | __________ | __________ | __________ | __________ | __________ |
| __________ | __________ | __________ | __________ | __________ | __________ | __________ |
| __________ | __________ | __________ | __________ | __________ | __________ | __________ |

# F. Equipment and final review

| Equipment type | Equipment ID | Verification / calibration status | Checked by | Review signature / date |
| --- | --- | --- | --- | --- |
| ___________ | ___________ | ___________ | ___________ | ___________ |
| ___________ | ___________ | ___________ | ___________ | ___________ |
| ___________ | ___________ | ___________ | ___________ | ___________ |

QC Shift 1 signature/date: ________________________    QC Shift 2 signature/date: ________________________    QC Supervisor review/date: ________________________

Release / disposition is NOT authorized by this blank form. Relevant approved operating procedures, technical acceptance data, sample-size plan and the parent document relationship must be confirmed prior to controlled issuance.