# ISO / Technical Closure Register

Requirements & Evidence • In-Process Control Scope

────────────────────────────────────────────────────

| DOCUMENT ID | DRAFT-M52-ISO-CLOSURE |
| --- | --- |
| PRODUCT | Silicone Peritoneal Catheter / Peritoneal Catheter Kit |
| ASSESSMENT DATE | 27 October 2026 |
| SOURCE WI | WI-8-2-2-52-1 Rev.0 (02.10.2025) |
| SOURCE IN-PROCESS RECORD | F-823-i99 Rev.7 (11.01.2026) |
| DRAFT STATUS | PROPOSED • NOT APPROVED • DO NOT USE FOR LOT RELEASE |
| ASSESSMENT BASIS | 27 screened requirement groups / process controls |
| CLOSED WITH VERIFIED EVIDENCE | 0 |
| APPROVAL STATUS | All controls open pending controlled decisions and executed evidence |

Source documents remain unchanged. No operational test results, technical approvals or effective revisions are asserted.

# 1  How to Read the Register

“Covered in draft” means the proposal has wording or a field, not that it has been released or carried out. The existing documents and product standards do not by themselves prove any result. ISO 20697 entries are conditional on formally approved product/intended-use applicability; ISO 80369-6 and ISO 8536-4 are initially not applied without specific technical justification. Status N/A requires authorized evidence-based justification, not omission.

# 2  Requirement Coverage & Objective Evidence

| ID | ISO / ref | ISO says / assessed requirement | Related gap | Required evidence | Verified available evidence | Documentary coverage | Operational status | Open action |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-001 | ISO 20697 §1 | Intended-use applicability to peritoneal dialysis catheter | GAP-001 | Approval of scope/variant matrix | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-002 | ISO 20697 §4 | Access and drainage functional performance (if in scope) | GAP-001; GAP-007 | Intended use + functional verification evidence | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-003 | ISO 20697 §5.1 | Risk management file review (if applicable) | GAP-001; GAP-009 | Controlled risk management file | None supplied / not verified | Draft proposal only | Open — Evidence | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-004 | ISO 20697 §5.2 | Biocompatibility (if applicable) | GAP-016 | Approved biological evaluation | None supplied / not verified | Draft proposal only | Open — Evidence | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-005 | ISO 20697 §5.3 | Detectability based on risk assessment | GAP-012 | Risk decision + detectability qualification if required | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-006 | ISO 20697 §5.4 | Surface finish and extraneous matter | GAP-011 | Approved visual control & completed records | None supplied / not verified | Draft proposal only | Open — Evidence | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-007 | ISO 20697 §§5.5.1–5.5.3 | OD size designation and effective length | GAP-003; GAP-005 | Approved drawing / size records | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-008 | ISO 20697 §5.5.4 | Balloon inflation designation IF balloon present | GAP-001 | Configuration evidence / balloon N/A justification | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-009 | ISO 20697 §5.6 | Connector misconnection risk | GAP-009; GAP-010 | Risk file + connector compatibility/type test | None supplied / not verified | Draft proposal only | Open — Evidence | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-010 | ISO 20697 §5.7 | MRI hazards IF applicable | GAP-001 | Risk applicability and MRI assessment if required | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-011 | ISO 20697 §5.8 | Sterility assurance for sterile drainage catheter | GAP-016 | Approved sterilization validation and release records | None supplied / not verified | Draft proposal only | Open — Evidence | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-012 | ISO 20697 §6.1 | Kink performance by intended use and risk | GAP-001 | Approved design verification/risk-based criterion | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-013 | ISO 20697 §6.2 / Annex B | Corrosion only for relevant exposed metal | GAP-001 | Metal configuration assessment; qualified evidence if applicable | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-014 | ISO 20697 §6.3 / Annex C | Negative-pressure deformation if intended | GAP-001 | Pressure use case and test evidence if applicable | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-015 | ISO 20697 §6.4.1 Table 1/Annex D | External-connection tensile strength if applicable | GAP-008 | Diameter classification and peak-force study | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-016 | ISO 20697 §6.4.2 Table 2/Annex E | Catheter/tube junction strength if applicable | GAP-008 | Diameter classification and peak-force study | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-017 | ISO 20697 §6.5 / Annex F | Collection-device impact resistance if included | GAP-001 | Kit configuration and evidence/N/A | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-018 | ISO 20697 §6.6 / Annex G | Claimed drainage flow performance | GAP-007 | Approved flow claim + method + execution evidence | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-019 | ISO 20697 §6.7 / Annex H informative | Non-balloon retention strength if applicable | GAP-006 | Controlled retention assessment and claim | None supplied / not verified | Draft proposal only | Open — Technical Input | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-020 | ISO 20697 §§6.8–6.10 / Annexes I–L | Balloon safety, lumen and traction IF balloon present | GAP-001 | Variant BOM; balloon evidence or N/A justification | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-021 | ISO 20697 §6.11 / Annex M | Vacuum air leakage if designed for suction | GAP-010 | Pressure use case/test evidence or N/A | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-022 | ISO 20697 §7 | Marking and instructions supplied by manufacturer | GAP-016 | Approved labels, IFU, packaging checks | None supplied / not verified | Draft proposal only | Open — Evidence | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-023 | ISO 80369-20 §4 / Annexes B–I | Connector test methods; only with applicable product connector requirements | GAP-002; GAP-010 | Authorized method, type-test results; applicability | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-024 | ISO 80369-6 §1 | Neuraxial connector standard; ordinarily not peritoneal dialysis | GAP-001; GAP-009 | Formal excluded-scope justification unless design proves otherwise | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-025 | ISO 8536-4 §1 | Gravity infusion set standard; not catheter by designation | GAP-001; GAP-002 | Formal excluded-scope justification or valid component link | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-026 | BS EN 1618 §1/Annexes A–F | Optional controlled catheter test methods; not all are lot requirements | GAP-002; GAP-007; GAP-008 | Method applicability / qualified procedure & results | None supplied / not verified | Draft proposal only | Open — Applicability | Approve applicability/criteria; verify supporting records and completed observations. |
| REQ-027 | Internal F-823-i99 + i15 | Component genealogy / hourly record validity | GAP-013; GAP-014; GAP-015; GAP-017 | Document control and completed records | None supplied / not verified | Draft proposal only | Open — Approval / Evidence | Approve applicability/criteria; verify supporting records and completed observations. |

# 3  Evidence / Document Control Register

| Evidence ID | Specific source required | Current record status |
| --- | --- | --- |
| EVD-01 | Approved IFU / intended use | Missing |
| EVD-02 | Approved product TF (TF-M52-PS / TF-52-PS / TF-52-PDBM reconciliation) | Missing |
| EVD-03 | Approved risk file and ISO/connector applicability matrix | Missing |
| EVD-04 | Approved sampling plan and measurable test methods | Missing |
| EVD-05 | Executed F-823-i99 lot results | Not provided |
| EVD-06 | Executed linked F-823-i15 lot results | Not provided |
| EVD-07 | Calibrated instruments and flow/tensile/leak worksheets | Not provided |
| EVD-08 | Connector qualification / cross-connection assessment | Not provided |
| EVD-09 | Biocompatibility and EtO / packaging qualification | Not provided |
| EVD-10 | Controlled QA / QC Manager approvals and actual final release evidence | Not provided |

# 4  Closure Criteria and Decisions

Verified Closed requires an APPROVED source, objectively completed and traceable verification, controlled record match to applicable device/configuration, review/authorization and no unresolved exceptions.

Documentary Covered — Operational Evidence Pending can be used only after applicable controlled wording and methods are approved, but execution proof is not verified; no such claim is presently made.

Open — Technical Input: controlled limits, design drawings, intended-use claims or methods are not available. Open — Approval: correct drafts await required authorized approvals. Open — Evidence: approved executed records have not been supplied.

QC Manager decision required: approve applicability and owners, secure approved TF, validate limits and control-plan routes, authorize proposed WI and in-process form after technical review, and re-run the closure register against executed evidence.