# ISO / WI Technical Gap Review

QC Manager • IN-PROCESS CONTROL • Master Gap & Revision Table

────────────────────────────────────────────────────

| DOCUMENT ID | DRAFT-QC-M52-GAP-2026-10 |
| --- | --- |
| PRODUCT | Silicone Peritoneal Catheter / Peritoneal Catheter Kit |
| ASSESSMENT DATE | 27 October 2026 |
| SOURCE WI | WI-8-2-2-52-1 Rev.0 (02.10.2025) |
| SOURCE IN-PROCESS RECORD | F-823-i99 Rev.7 (11.01.2026) |
| DRAFT STATUS | PROPOSED • NOT APPROVED • DO NOT USE FOR LOT RELEASE |
| STANDARD APPLICABILITY | ISO 20697:2018 = CONDITIONAL; BS EN 1618:1997 = common methods; ISO 80369-20 = connector methods; ISO 80369-6 & ISO 8536-4 = not automatically applicable |

Source documents remain unchanged. No operational test results, technical approvals or effective revisions are asserted.

# 1  Executive Summary

17 source-based documentary/control gaps: 6 high (P1), 10 medium (P2), 1 low (P3). Technical scope has not been approved, so this is NOT a certified ISO compliance or product conformity conclusion.

Key decisions for QC Manager: approve intended-use and standard applicability matrix; request the missing approved TF-M52-PS / TF-52-PS / TF-52-PDBM and variant drawings; resolve quantitative flow, tensile and cuff limits; authorize a separate F-823-i99 in-process form with measurement fields; approve controlled verification / sampling routes; retain sterilization, supplier, type and final-release evidence in their existing distinct records.

Not supplied: product technical file/drawings, IFU, risk file, approved method sheets, calibration and validation evidence, historical executed records, approved sampling plan, ISO 80369-1 text. Consequently every claimed operational closure remains open.

# 2  Documents Reviewed

| Document | Identity | Length | Assessment |
| --- | --- | --- | --- |
| WI-8-2-2-52-1 | Rev 0; 02 Oct 2025 | 5 pages | Primary source; p1–p5 |
| F-823-i99 | Header Rev 7; 11 Jan 2026 | 1 worksheet; A1:Q29 | In-process assembly form; worksheet tab says Rev 6 |
| F-823-i15 | Header Rev 16; 26 Aug 2026 | 1 worksheet; A1:W32 | Upstream molding/extrusion; worksheet tab says Rev 15 |
| ISO 20697:2018 corrected edition | 2018-09 | 44 PDF pp | Conditional drainage product standard |
| BS EN 1618:1997 | 1997 | 12 PDF pp | Non-intravascular catheter test methods (scanned) |
| ISO 80369-20:2015 | 2015 | 36 PDF pp | Common small-bore connector test methods |
| ISO 80369-6:2016 | Corrected 2016-11-15 | 52 PDF pp | Neuraxial; scope not shown to apply |
| ISO 8536-4:2019 | 2019 | 24 PDF pp | Gravity infusion set; scope not shown to apply |

# 3  Scope & Critical Applicability Rules

This assessment addresses the in-process assembly and component workflow for a silicone peritoneal dialysis catheter/kit, NOT a standalone routine final inspection report. Column names “Existing Inspection Text” and “Revised Inspection Text” in the mandatory master table below refer EXCLUSIVELY to existing/proposed F-823-i99 in-process entries. Incoming records, design/type tests, sterilization records and final release records remain independent.

ISO 20697 applies to sterile drainage catheters removing fluid/air to the exterior; the uploaded WI describes infusion into the peritoneal cavity. Dialysate drainage may occur, but the approved intended use/IFU and risk file were not available. Clause-by-clause ISO 20697 applicability is therefore a CONDITIONAL technical question. ISO 80369-6 concerns neuraxial connectors, while ISO 8536-4 concerns gravity infusion sets; neither shall be applied automatically.

# 4  Applicable Requirement Reference Map

| Reference / clause | Verified source meaning / decision |
| --- | --- |
| ISO 20697 §1 (PDF 7 / printed 1) | Drainage-catheter scope; conditional for this peritoneal dialysis device. |
| ISO 20697 §§5.3–5.6 (PDF 10–11 / printed 4–5) | Surface, conditional detectability, size designations, connector risk. |
| ISO 20697 §§6.4.1–6.4.2 (PDF 12–13 / printed 6–7) | Different tensile tables for external connections versus tube/junctions; not interchangeable. |
| ISO 20697 §6.6 (PDF 13 / printed 7) | If claimed drainage flow: ≥80%/≥90% claimed by OD category, subject to applicability. |
| ISO 80369-20 §4 (PDF 7–8 / printed 1–2) | Lists common test methods; does not by itself set all connector acceptance limits. |
| BS EN 1618 §1; Annexes A–F (PDF 8–11 / printed 3–6) | Common catheter test methods, not an automatic lot-testing program. |
| ISO 80369-6 §1 (PDF 7 / printed 1) | Neuraxial-only connector product requirements; not general catheter requirement. |
| ISO 8536-4 §1 (PDF 5 / printed 1) | Single-use gravity infusion sets; not generic peritoneal catheter specification. |

# 5  Initial Findings

GAP-001 | P1 | Unverified intended-use / normative scope

GAP-002 | P2 | Incorrect standard identity and method/criterion conflation

GAP-003 | P1 | Controlled product specification and variant mapping not verified

GAP-004 | P2 | Hole count and geometry under-defined across variants

GAP-005 | P2 | OD/ID/length measurements lack measurable results

GAP-006 | P2 | Cuff dimension and spacing values conflict or are unverified

GAP-007 | P1 | Flow performance method and units are insufficient

GAP-008 | P1 | Tensile acceptance and controlled method missing

GAP-009 | P1 | Connector application and misconnection controls not demonstrated

GAP-010 | P1 | Gravity-flow leak check is not an established standardized connector test

GAP-011 | P2 | In-process visual criteria not standardized by defect and observation

GAP-012 | P2 | Radiopacity cannot be verified only by visual color

GAP-013 | P2 | Hourly record lacks sample size, readings and disposition traceability

GAP-014 | P2 | Extrusion/molding to kit genealogy and calibration reference omitted

GAP-015 | P2 | Missing explicit in-process Hold / Reject / corrective disposition

GAP-016 | P2 | In-process versus final release evidence boundaries are not explicit

GAP-017 | P3 | Revision tab labels and review fields inconsistent

# 6  MASTER GAP & REVISION TABLE

MANDATORY SEVEN-COLUMN COMPARISON. IMPORTANT: “Inspection” column labels in this table mean IN-PROCESS F-823-i99 only. Every revised statement is proposed wording, not an approved operational control. ISO 20697 requirements are stated as conditional unless scope approval is available.

| Gap / Our Problem | ISO Says | Existing WI Text | Revised WI Text | Existing Inspection Text | Revised Inspection Text | Evidence (ISO Closure) |
| --- | --- | --- | --- | --- | --- | --- |
| GAP-001 ¦ P1 Unverified intended-use / normative scope Problem: Scope cannot be established from available source documents. | ISO 20697:2018 §1 (PDF p7; printed p1) applies to single-use sterile drainage catheters removing fluid/air from a body cavity or wound to the exterior. ISO 80369-6:2016 §1 (PDF p7; printed p1) is specifically neuraxial. ISO 8536-4:2019 §1 (PDF p5; printed p1) is for gravity infusion sets. | WI p1 §§1–3 describes a silicone peritoneal catheter intended to infuse dialysis solution; p2 §5.3.3 cites ISO 20697 while §5.3.2 cites ISO 8536. | Proposed §2.1: Before assigning product-standard acceptance limits, QA/Engineering shall approve the intended-use and design-configuration applicability matrix for every catheter variant. Classify ISO 20697 as applicable, partially applicable, or not applicable by clause with justification; ISO 80369-6 and ISO 8536-4 shall not be claimed applicable to the catheter without a documented scope rationale. | i99 header references WI and TF-52-PS / TF-52-PDBM, but the form does not record intended-use configuration or applicable standard route. | Header: Product variant / intended use / configuration / approved applicability-matrix ID and revision. Result: Matrix verified / Pending. No additional routine ISO test is created. | Approved IFU/technical file; risk management; signed standard applicability matrix identifying product/variants. Verification route: Technical-file review. Status: Open — Technical Input / Approval. |
| GAP-002 ¦ P2 Incorrect standard identity and method/criterion conflation Problem: Incorrect standard identity and method/criterion conflation. | BS EN 1618:1997 §1 and Annexes A–F (PDF pp8–11; printed pp3–6) specifies test methods for non-intravascular catheters. ISO 80369-20:2015 §1 and §4 / Annexes B–I (PDF pp7–8; printed pp1–2) supplies common connector test methods; application-specific performance criteria come from the applicable connector/product standard. | WI p2 §5.3.1 calls “1618-1997” a standard without identifying BS EN; §5.3.2 cites “8536”; p3 §5.3.4 cites ISO 80369-1 and -20 as if a gravity saline check establishes connector standard conformity. | Proposed §3: Identify standards by exact number, title and edition. Use BS EN 1618 only for its selected test method(s), where justified; use ISO 80369-20 only with an approved applicable connector performance specification. Remove unsupported statements that a visual flow test proves compliance with an unrelated product or connector standard. | i99 rows 10, 11 and 19 use generic “strong” or “no leakage” wording without method identity or authorized verification route. | Header: Approved method / control-plan ID. Tensile, flow and leakage entries: record actual WI method or controlled verification record; do not print an unsupported ISO-conformity assertion. | Approved reference register, applicability matrix, and controlled method sheets. Route: Document control / technical verification. Status: Open — Approval. |
| GAP-003 ¦ P1 Controlled product specification and variant mapping not verified Problem: Controlled product specification and variant mapping not verified. | ISO 20697:2018 §5.5.1–5.5.3 (PDF p10; printed p4), if applicable, requires OD size designation and effective length identification, with OD expressed in mm and ±1 French tolerance on the stated OD; the standard does not set a general effective-length tolerance. | WI p3 §6 cites “TF-M52-PS”; p3 §5.3.6 repeatedly states “as per specification”. The controlled TF-M52-PS and drawings were not supplied. | Proposed §6.1: QC shall use the currently approved specification and drawing identified by product code, catheter configuration and revision. Document Control shall reconcile TF-M52-PS versus TF-52-PS / TF-52-PDBM. No dimension, tolerance, hole count or cuff dimension shall be accepted from an uncontrolled/unknown reference. OPEN TECHNICAL INPUT REQUIRED: signed current product-specific specification and drawing. | i99 header cites “TF-52-PS / TF-52-PDBM”; rows 15, 16 and 17 state only “as per specification”. Supplied source cannot establish whether TF-M52-PS and TF-52-PS are identical. | Header: Product code, variant, specification number / revision, drawing number / revision. Rows requiring numeric specification shall carry target, tolerance and measured results from the controlled specification. | Signed controlled TF / approved dimensions and mapping, document revision control. Route: Technical-file / Document control. Status: Open — Technical Input. |
| GAP-004 ¦ P2 Hole count and geometry under-defined across variants Problem: Hole count and geometry under-defined across variants. | ISO 20697:2018 §5.4 (PDF p10; printed p4), only if scope applies, requires externally visible effective-length surfaces free of extraneous matter and potentially harmful process defects; no universal 72-hole count is specified. | WI p3 §5.3.6 states “72 holes in total, clear, no blocks” for a 3-cuff model and “smooth edges, round shape”. Applicability to other variants is not defined. | Proposed §7.4: For each approved variant verify total hole count, approved positions/diameters (if specified), lumen patency and absence of blocked holes, fibers, flash and sharp edges against the corresponding drawing. The 72-hole target is applicable only to a variant for which the approved drawing explicitly specifies 72. | i99 rows 3–4 (sheet rows 12–13) say holes must be clear and dimensionally accurate but omit the actual number and measurable criteria. | Characteristic: Hole count / pattern / edge condition. Approved target and drawing ref: __. Actual holes counted: __. Diameter / position where required: __ mm. Pass/Fail/Pending; sample ID and time. | Approved variant drawings, completed in-process dimensional/hole records. Route: In-process. Status: Open — Technical Input / Evidence. |
| GAP-005 ¦ P2 OD/ID/length measurements lack measurable results Problem: OD/ID/length measurements lack measurable results. | ISO 20697:2018 §§5.5.2–5.5.3 (PDF p10; printed p4), conditional applicability: OD expression in mm; effective length recorded in mm or cm based on length. No specific ID or all-product length tolerances are stipulated. | WI p3 §5.3.6 lists “Catheter length / OD / ID — According to specification” and “Length of tubing — As per specification” without which size/configuration or measurement points. | Proposed §7.5: Measure OD, ID, effective length and overall length at drawing-defined locations using identified calibrated devices, and record individual readings in mm (or approved length unit). Acceptance criteria shall be the exact controlled variant-specific dimensions/tolerances. OPEN TECHNICAL INPUT REQUIRED for target values and sampling points. | i99 rows 14 and 17 (sheet rows 23, 26) have generic “Accourding to specification / As per specification”; hourly cells do not define actual values or units. | Characteristics: OD, ID, effective length, overall length. Fields: Spec ref/rev; nominal; lower/upper limits; instrument ID; measured values; unit; Pass/Fail; time/shift/sample. | Approved drawing, calibration certificates, traceable dimensional records. Route: In-process. Status: Open — Technical Input / Evidence. |
| GAP-006 ¦ P2 Cuff dimension and spacing values conflict or are unverified Problem: Cuff dimension and spacing values conflict or are unverified. | ISO 20697:2018 §6.7 (PDF p13; printed p7), if the product is within scope and uses non-balloon retention means, requires ability to prevent undesired dislodgement; the clinically relevant retention value is manufacturer-defined by intended use/risk. The standard does not prescribe the WI cuff spacings. | WI p3 §5.3.6 says “Distance of cuff into cuffs — 6 cm, 13 cm distance” and “Cuffed adhesive status — Accurate, strong, fixed not move”. | Proposed §7.6: QC shall measure cuff width, cuff count, cuff locations and inter-cuff distances against the specific approved variant drawing; inspect adhesion and displacement against a validated company-defined method/limit. Retention verification shall follow the approved risk-based evidence route if relevant. The existing 6 cm, 13 cm and 6.46 ±3 mm values are UNVERIFIED until reconciled with the signed drawing. | i99 row 15 (sheet row 24) says “Cuff width is 6.46±3mm” and row 16 (sheet row 25) says “Distance between cuffs — As per specification”. Neither source supplies an approved drawing confirming these limits. | Characteristic: Cuff number, width, placement, inter-cuff spacing, adhesion. Record nominal/tolerances, individual measurements, specimen ID, observations and approved drawing/method references. | Controlled drawings and retention assessment when applicable; measured and reviewed lot records. Route: In-process + qualification as applicable. Status: Open — Technical Input / Evidence. |
| GAP-007 ¦ P1 Flow performance method and units are insufficient Problem: Flow performance method and units are insufficient. | ISO 20697:2018 §6.6 and Annex G (PDF p13 / annex from PDF p26; printed p7 / p20), IF the drainage-catheter standard applies and a flow rate is claimed: measured flow must be at least 80% of claimed for OD <1.0 mm, or 90% for OD ≥1.0 mm. No fixed 540/450 mL figure is imposed. BS EN 1618:1997 Annex E (PDF pp10–11; printed pp5–6) is a common flow-test method. | WI p2 §5.3.2 describes constant-head water setup and three repetitions; p3 §5.3.6 says “540 ±15ml / 450 ±15ml as per size” without time interval, head, temperature or specimen identification. | Proposed §8.2: Use the approved flow method specifying fluid, head/pressure, temperature, stabilization, timed collection, number of repeats and expression in mL/min (or the exact approved unit). Record the three individual values, their assessment and catheter size; apply the manufacturer-approved acceptance band. If ISO 20697 §6.6 is applicable, verify performance against the applicable percentage of approved manufacturer-claimed flow using Annex G. OPEN TECHNICAL INPUT REQUIRED: claim, time basis and company limits; do not use 540/450 mL alone as a pass limit. | i99 flow row 11 (sheet row 20) only reads “As per specification”, with hour-mark cells and no numeric result/time unit. | Flow rate: method / head / fluid / temperature / duration / sample ID / measured volume ×3 / calculated rate and unit / acceptance limit / evidence ID / Pass-Fail-Pending. | Approved test method & product-specific flow claim, completed raw results with equipment calibration. Route: In-process if approved plan; otherwise qualification/periodic. Status: Open — Technical Input / Evidence. |
| GAP-008 ¦ P1 Tensile acceptance and controlled method missing Problem: Tensile acceptance and controlled method missing. | ISO 20697:2018 §6.4.1 Table 1 (PDF p12; printed p6), if applicable, sets external connections ≥2–≤4 mm: 5 N; >4 mm: 15 N. §6.4.2 Table 2 (PDF p13; printed p7) sets tubular portions/junctions ≥2–≤4 mm: 10 N; >4 mm: 20 N. Tubing <2 mm requires manufacturer-defined values based on risk. Annexes D/E are normative. | WI p2 §5.3.1 references “1618-1997” and p3 §5.3.3 applies 500 mm/min without an approved method source or acceptance force; p3 §5.3.6 says “Strong, no breakage” with “Cometech machine”. | Proposed §8.3: Use an approved tensile method that identifies whether an external connection (ISO 20697 §6.4.1) or catheter-body/tube junction (§6.4.2) is tested, calibrated equipment, fixture, gauge length and controlled crosshead rate. Record peak force (N), mode/location of failure and comparison with the exact approved criterion. Apply ISO table limits only after technical approval that ISO 20697 applies and OD category is verified. Existing 500 mm/min shall not be assumed normative without source verification. | i99 row 19 (sheet row 28) says “Strong, no breakaged”; no force, loading rate, break location, sample ID, test identification or actual peak force field. | Tensile: test-piece/joint ID; OD category (mm); method; machine/calibration ID; speed (mm/min); measured peak N; criterion N; failure mode; time, sample; Pass/Fail/Pending. | Applicability approval, tensile test plan, equipment calibration and raw force/time traces or signed worksheet. Route: In-process/qualification per control plan. Status: Open — Technical Input / Evidence. |
| GAP-009 ¦ P1 Connector application and misconnection controls not demonstrated Problem: Connector application and misconnection controls not demonstrated. | ISO 20697:2018 §5.6 (PDF p11; printed p5), when applicable, does not prescribe a standard drainage connector but requires avoidance of misconnection based on risk assessment and ISO 80369-1 general requirements. ISO 80369-6:2016 §1 (PDF p7; printed p1) covers neuraxial, not peritoneal-dialysis connections. | WI p3 §§5.3.3–5.3.4 controls FLL/MLL fit but does not identify approved connector interface category, drawing or compatible transfer-set configuration. | Proposed §7.8: Engineering/QA shall approve connector type, intended compatible mating devices and interface drawing; evaluate wrong-route misconnection risk and identify applicable connector-standard requirements. During production QC shall verify presence, integrity and fit only using authorized counterpart(s). Do not apply ISO 80369-6 neuraxial geometry by inference. | i99 row 10 (sheet row 19) “FLL & MLL Status — Strong matching, no leakage”; no mating-part IDs or controlled cross-connection review. | Connector type/ref/rev; approved counterpart part #; locking/fitting status; sample ID; observed defect; record of design / misconnection assessment evidence ID. | Approved connector drawing, IFU/compatible set matrix, risk assessment, conformity/type-test report if required. Route: Design/qualification + In-process fit. Status: Open — Technical Input / Evidence. |
| GAP-010 ¦ P1 Gravity-flow leak check is not an established standardized connector test Problem: Gravity-flow leak check is not an established standardized connector test. | ISO 80369-20:2015 §4 & Annexes B–I (PDF pp7–8; printed pp1–2) defines different common connector methods. ISO 20697:2018 §6.11 / Annex M (PDF p14; printed p8), IF applicable to negative-pressure design, requires no air leakage at manufacturer-defined maximum negative pressure. Methods/limits depend on relevant product and connector applicability. | WI p3 §5.3.4 describes normal saline flowing by gravity while checking for visible leakage and claims conformity to ISO 80369-1 and -20. | Proposed §8.4: Clearly separate the routine gravity functional joint leak check (observed leakage: none) from connector type-verification under an approved applicable ISO 80369 method. If the device is intended for negative pressure, maintain separate approved qualification evidence per ISO 20697 §6.11 (subject to scope decision). A gravity visual check alone shall not be reported as conformity to ISO 80369-20. | i99 row 10 does not identify pressure, test duration, approved liquid, counterpart, measurable result or evidence record. | Functional joint leakage: mating part ID; approved functional method; test conditions / observation duration; observed leakage Yes/No; result. Separate field: Connector qualification/test report ID (if required), not repeated every lot. | Approved functional leak WI and lot records; risk-based type-test report and pressure method as applicable. Route: In-process + Qualification. Status: Open — Approval / Evidence. |
| GAP-011 ¦ P2 In-process visual criteria not standardized by defect and observation Problem: In-process visual criteria not standardized by defect and observation. | ISO 20697:2018 §5.4 (PDF p10; printed p4) requires drainage catheter surface to be free of extraneous matter and unsafe process defects where applicable; minimum 2.5× magnification is conditional on risk, not always compulsory. | WI p3 §5.3.6 states “Strong without any defect”, “Smooth”, “Good”, “clear, no blocks”, “no bumps” without defect classification, inspection lighting or rework action. | Proposed §7.2: Inspect tubing, surface, coil shape, punch holes, particulate matter, cracks, cuts, collapse, flash, protrusions and blocked lumens according to an approved defect catalogue and visual conditions; record characteristics individually. Any visible defect affecting integrity, patency or patient-contact surface shall cause immediate Hold and defined disposition, not an unsupported generic “Good”. | i99 rows 1, 2, 3, 4, 5, 7, 13 record appearance and faults but generic “Good surface / Proper coiled formed” without variant-controlled acceptance references. | Existing visual lines retained; provide exact defect category / acceptance source / Pass-Fail and remarks/defect code per sample/time. Specify magnification only where approved risk assessment requires. | Approved visual defect catalog / QC control plan, training, completed records and hold actions. Route: In-process. Status: Open — Approval / Evidence. |
| GAP-012 ¦ P2 Radiopacity cannot be verified only by visual color Problem: Radiopacity cannot be verified only by visual color. | ISO 20697:2018 §5.3 (PDF p10; printed p4) requires detectable effective length by X-ray or other means if required by the risk assessment. Visual blue/white colour alone does not establish radiographic detectability. | WI p3 §5.3.6: “X-ray opaque — Clear, good, blue color / White color” (visual). | Proposed §7.9: For variants requiring radiographic detectability, confirm correct specified radiopaque material/marker during in-process assembly against controlled BOM. Maintain separately approved detectability verification evidence from qualification/material validation; do not infer detectability from colour. If not required, document risk-based N/A. | i99 row 18 (sheet row 27): “X-Ray Radio opaque — Clear, good, blue color, white color” without technical evidence. | Radiopaque marker/material presence: component lot & spec, identification check, result; qualification evidence ref / N/A reason. Do not add routine radiography absent an approved control plan. | Approved risk assessment/detectability study or justified N/A; incoming material traceability and hourly assembly checks. Route: Qualification/Supplier + In-process. Status: Open — Technical Input / Evidence. |
| GAP-013 ¦ P2 Hourly record lacks sample size, readings and disposition traceability Problem: Hourly record lacks sample size, readings and disposition traceability. | Documentary operational-control gap (no ISO 20697 clause specifies the F-823-i99 hourly schedule). The source i99 currently provides a time grid; a frequency/quantity cannot be invented from product ISO. | WI pp2–3 §§5.3.1–5.3.6 says randomly selected samples and three repeats for flow, but does not specify plan ID or recording fields for actual results by sample. | Proposed §§6.3, 9: QC shall follow an approved risk-based sampling/control plan with sampling event, piece/sample ID, quantity, results and trend/disposition. Time slots are monitoring prompts only, not an ISO-mandated frequency. If product fails, record event, segregation and NCR/reject reference and notify production/QA per approved workflow. | i99 time header sheet rows 8–9 provides 7/9/11/1/3/5/7/9/11/1/3/5 and 19 rows but not sample count, lot quantity, result method, or hold/nonconformity reference. | Header: line/machine, date, shift, lot, qty/lot, approved sampling plan and sample count, QC operator, instrument IDs. Per reading: time, sample IDs, Actual/Pass/Fail/N/A/Pending, exception ref, remarks. Footer: Hold / Accepted for next production stage / Pending, supervisor review. | Approved frequency plan, completed traceable records and rejection/hold log. Route: In-process. Status: Open — Approval / Evidence. |
| GAP-014 ¦ P2 Extrusion/molding to kit genealogy and calibration reference omitted Problem: Extrusion/molding to kit genealogy and calibration reference omitted. | Traceability / product and process verification gap from WI and forms; ISO 20697 technical product standard does not prescribe an i15→i99 batch ID mapping. Company-controlled genealogy is needed to demonstrate which component observations support the kit. | WI p5 §10 lists F-823-i15 (molding/extrusion) and F-823-i99 but does not state how component lots and test equipment IDs are transferred to final assembly record. | Proposed §6.4: At kit assembly, verify the released tubing/component lot codes against F-823-i15 and record their corresponding lot and QC reference on F-823-i99. QC shall only use equipment within calibration and record instrument ID/reference for quantitative work. | i15 Rev16 logs part/code, lot, machine, specification, shift, inspection instruments; i99 Rev7 logs catheter lot/type and instrument field but no linked i15 source lot or machine trace. | Component/extrusion lot(s), i15 record ID, machine/source, incoming component release status, instrument ID and next calibration due date where applicable. | Completed matching i15 and i99 records, component genealogy, calibration status. Route: In-process. Status: Open — Evidence. |
| GAP-015 ¦ P2 Missing explicit in-process Hold / Reject / corrective disposition Problem: Missing explicit in-process Hold / Reject / corrective disposition. | Operational document-control gap evidenced in source WI/form; no specific ISO product clause requires a particular internal reject workflow or approval title. | WI p3 §5.3.6 lists inspection controls but does not define immediate action on failed in-process observation or repeat test decision. | Proposed §10: On an out-of-specification, contamination, leakage, dimensional, blocked-lumen or integrity failure, stop release of affected material to next process stage, identify scope/quantity, apply QC Hold tag, log rejection/NCR reference, and route for authorized disposition under the company-approved nonconforming product process. QC shall not overwrite a failed initial reading with a passing retest. | i99 sheet row 29 only “Checked by Q.C.” and “Reviewed by: QCM.” with quality remarks but no nonconforming product status, affected quantity or issue link. | Exception block: Defect code, sample/lot, quantity affected, status (Hold/Reject/Rework/Pending), NCR/rejection reference, disposition authority/approval ref, follow-up verification and reviewed initials. | Approved internal NC/reject workflow, completed Hold/NCR records and approval evidence. Route: In-process / QC process. Status: Open — Approval / Evidence. |
| GAP-016 ¦ P2 In-process versus final release evidence boundaries are not explicit Problem: In-process versus final release evidence boundaries are not explicit. | ISO 20697:2018 §§5.2, 5.8, 7 (PDF pp10,12,15; printed pp4,6,9), if applicable, addresses biological safety, sterile device assurance and manufacturer information; it does not require adding each validation/lab test to hourly in-process form. | WI p4 §§7–8 and p5 §9 contain packaging, ethylene oxide sterilization, biological indicators and final release workflows mixed with in-process instructions. | Proposed §§11–12: Preserve existing packaging, EtO WI-7-5-1-1 and final-release procedures as separate downstream controls. The in-process record shall only capture applicable assembly process acceptance and references to downstream hold/transfer status. No new routine endotoxin/biocompatibility/sterility/EtO test shall be added without an approved plan. | i99 has no linkage fields for separate packaging/sealing, sterilization, QC lab or final release records, but these are not intrinsically hourly form fields. | Footer: Assembly in-process status / downstream packaging record ref / transfer ref (if available). Mark sterilization, biocompatibility and final release as separate evidence routes, not Pass/Fail hourly items. | Approved sealing/EtO/biological evaluation validation references, controlled lab release and final record F-8-2-4 when available. Route: Validation/Lab/Final release (separate). Status: Open — Evidence. |
| GAP-017 ¦ P3 Revision tab labels and review fields inconsistent Problem: Revision tab labels and review fields inconsistent. | Internal document-control gap supported by available electronic files; not a specific ISO 20697 product-performance clause. | WI p1 says Rev 0 dated 02.10.2025; p5 §10 cites i99 / i15 without revision numbers. | Proposed §14: Verify effective controlled versions of WI, F-823-i99 and F-823-i15 before production use; display current document numbers, effective dates, revision and superseded versions in the master document list. Worksheet-tab names shall be reconciled with the approved header. Do not change official revision identity without Document Control approval. | Supplied i99 sheet tab says “F-823-i99 (Rev 6)” while printed document header says Rev 7 (11.01.2026). Supplied i15 sheet tab says “Rev 15” but printed header says Rev 16 (26.08.2026). This does not prove the document itself is obsolete, but needs correction/confirmation. | Use controlled header: F-823-i99 Proposed Revision — approval pending; Approved form number/revision/effective date fields; Record prepared/checked/reviewed by functions per approved matrix. | Approved master document register and authorized corrected controlled forms/revision history. Route: Document control. Status: Open — Approval. |

# 7  Required Verification Strategy

Routine in-process: approved visual assembly checks, hole/size/cuff measurements and production-stage functional fit/flow/leakage checks only when defined by a documented control plan; do not insert an unapproved ISO test or frequency into hourly F-823-i99.

Component source: F-823-i15 Rev16 for extrusion/molding genealogy. Alternative evidence: controlled product qualification, supplier controls, design connector testing, biocompatibility assessment, sterilization validation, packaging qualification and final release under their respective processes.

# 8  Proposed Revised WI & Form Overview

The companion WI is reconstructed as an IN-PROCESS CONTROL WI incorporating original raw-material, packaging, sterilization and release interfaces (not duplicate inspection forms). The companion F-823-i99 draft includes numeric readings, references and disposition fields while preserving the original 19 control subjects and source time grid as a configurable monitoring layout.

# 9  ISO / Operational Closure Summary

27 requirements/control routes screened; 0 evidenced as Verified Closed. Applicability and technical inputs are open. Proposed drafts cannot be “100% documentary coverage” until the controlled TF, IFU, risk file, approved criteria and review/approval exist.

# 10  QC Manager Actions

| Gap | Priority | Required action / decision | Owner function | Required evidence | Status |
| --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 | Unverified intended-use / normative scope | QA / Engineering | Approved IFU/technical file; risk management; signed standard applicability matrix identifying product/variants. Verification route: Technical-file review. | Open |
| GAP-003 | P1 | Controlled product specification and variant mapping not verified | QA / Engineering | Signed controlled TF / approved dimensions and mapping, document revision control. Route: Technical-file / Document control. | Open |
| GAP-007 | P1 | Flow performance method and units are insufficient | QC / QA | Approved test method & product-specific flow claim, completed raw results with equipment calibration. Route: In-process if approved plan; otherwise qualification/periodic | Open |
| GAP-008 | P1 | Tensile acceptance and controlled method missing | QC / QA | Applicability approval, tensile test plan, equipment calibration and raw force/time traces or signed worksheet. Route: In-process/qualification per control plan. | Open |
| GAP-009 | P1 | Connector application and misconnection controls not demonstrated | QA / Engineering | Approved connector drawing, IFU/compatible set matrix, risk assessment, conformity/type-test report if required. Route: Design/qualification + In-process fit. | Open |
| GAP-010 | P1 | Gravity-flow leak check is not an established standardized connector test | QC / QA | Approved functional leak WI and lot records; risk-based type-test report and pressure method as applicable. Route: In-process + Qualification. | Open |
| GAP-013 | P2 | Hourly record lacks sample size, readings and disposition traceability | QC / QA | Approved frequency plan, completed traceable records and rejection/hold log. Route: In-process. | Open |
| GAP-017 | P3 | Revision tab labels and review fields inconsistent | Document Control | Approved master document register and authorized corrected controlled forms/revision history. Route: Document control. | Open |

# 11  Final Conclusion & Source Traceability

The existing WI and F-823-i99 are not presently sufficient as a fully traceable in-process control set. An edited draft can correct phrasing and create recording fields; however verified ISO applicability, complete numerical acceptance criteria and executed proof cannot be confirmed. No original files were modified.

WI-8-2-2-52-1 Rev0: p1 §§1–5.1; p2 §§5.2–5.3.2; p3 §§5.3.3–6; p4 §§7–8; p5 §§9–11.

F-823-i99 Rev7: worksheet A1:Q29, header rows 2–5, characteristics sheet rows 10–28, footer row 29.

F-823-i15 Rev16: worksheet A1:W32, header rows 1–10, checks rows 13–26, legend rows 27–31.

ISO 20697:2018: see exact section/page citations in master table and requirement register. Other supplied standards were evaluated for scope before acceptance-rule transfer.