# In-Process Control Record

Silicone Peritoneal Catheter Kit • Proposed F-823-i99 Reconstruction

────────────────────────────────────────────────────

| DOCUMENT ID | DRAFT-F-823-i99 |
| --- | --- |
| PRODUCT | Silicone Peritoneal Catheter / Peritoneal Catheter Kit |
| ASSESSMENT DATE | 27 October 2026 |
| SOURCE WI | WI-8-2-2-52-1 Rev.0 (02.10.2025) |
| SOURCE IN-PROCESS RECORD | F-823-i99 Rev.7 (11.01.2026) |
| DRAFT STATUS | PROPOSED • NOT APPROVED • DO NOT USE FOR LOT RELEASE |
| SOURCE | F-823-i99 Header Rev7 / tab “Rev6” — discrepancy for Document Control |
| RELATED UPSTREAM | F-823-i15 header Rev16 / tab “Rev15” |
| RECORD PURPOSE | In-process assembly checks; not final Inspection & Test Report |

Source documents remain unchanged. No operational test results, technical approvals or effective revisions are asserted.

# A  Identification / Controlled Baseline

| Required header field | Actual entry |
| --- | --- |
| Product name / code | ________________________________ |
| Type / variant / no. of cuffs | ________________________________ |
| Batch / lot / planned production qty | ________________________________ |
| Date / shift / area / line / machine | ________________________________ |
| Approved TF / drawing / revision | ________________________________ |
| Approved control plan / revision | ________________________________ |
| Approved ISO applicability-matrix ID | ________________________________ |
| Component lot / i15 source record ID | ________________________________ |
| Sampling plan / quantity / sample IDs | ________________________________ |
| QC inspector / supervisor | ________________________________ |
| Instruments / IDs / calibration refs | ________________________________ |
| Linked approved test worksheets | ________________________________ |

# B  In-Process Acceptance / Measurement Matrix

Use one entry per sampling event and product variant. For measurements record limits and actual numeric values. The original 19 characteristics are retained; no automatic ISO test frequency is asserted. Blank means NOT ASSESSED, never Passed.

## 01  Appearance of tubing

| Acceptance / approved source | No dots / fish eyes / black spots; no unsafe defects | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual / approved defect atlas | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 02  Condition of catheter

| Acceptance / approved source | No collapse, cuts, unintended holes / lumen damage | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual / patency check | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 03  Number / placement of holes

| Acceptance / approved source | Approved variant drawing; clear no blockage/flash | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Count / approved dimensional method | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 04  Punching-hole edges

| Acceptance / approved source | Smooth/rounded without fibers, sharp edges or flash | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual; approved magnification if needed | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 05  Coiling condition

| Acceptance / approved source | Correct approved formed shape | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual / forming specification | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 06  Coiling width and height

| Acceptance / approved source | Approved per-variant target ± tolerance | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Calibrated scale/gauge | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 07  Trocar appearance (optional)

| Acceptance / approved source | Included only if BOM specifies; sound straight tapered | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual / approved diagram | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 08  Trocar length (optional)

| Acceptance / approved source | Approved variant drawing range | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Calibrated length measure | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 09  Injection sealing cap

| Acceptance / approved source | Secure fit, undamaged/no cracking | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Fit / visual with approved mate | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 10  FLL / MLL connector fit and gravity leak

| Acceptance / approved source | Correct mate; secure fit; no visible leak under approved functional conditions | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Controlled functional method (NOT ISO 80369 qualification) | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 11  Catheter flow rate

| Acceptance / approved source | Controlled manufacturer claim / spec with correct unit and test conditions | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Approved flow method: 3 readings if authorized | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 12  Catheter type

| Acceptance / approved source | Designated silicone short-term / long-term variant | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Controlled BOM / code | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 13  Catheter surface status

| Acceptance / approved source | No bumps / harmful defects / extraneous matter | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual / defect catalogue | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 14  Catheter OD, ID

| Acceptance / approved source | Drawing-specific target/range in mm | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Calibrated dimensional tool | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 15  Cuff number, width, adhesion

| Acceptance / approved source | Approved variant drawing; secure without movement | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Visual + numeric width/adhesion if specified | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 16  Inter-cuff distances / positions

| Acceptance / approved source | Drawing-specific dimensional range in mm or cm | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Calibrated scale/gauge | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 17  Length of tubing / effective length

| Acceptance / approved source | Drawing-specific ranges and locations | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Calibrated length tool | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 18  Radiopaque material presence

| Acceptance / approved source | Correct specified marker/material if required by risk | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Material ID / incoming certificate; not radiography | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

## 19  Peak tensile force / failure mode

| Acceptance / approved source | Approved threshold N per applicable test category and approved method | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| --- | --- | --- | --- |
| Method / instrument | Calibrated tensile test machine; result N | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec target / limit / unit | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |
| Spec or drawing / revision | ____________________________ | Actual: __________ | Result: ☐ Pass ☐ Fail ☐ Pending ☐ N/A |

Date/time: __________   Sample ID / qty: __________   Actual measured value / observation / units: _____________________________________

Remarks / defect code / evidence or worksheet ID: __________________________________________________________________________

# C  Time-Grid Record (Original Monitoring Slots Retained)

Original F-823-i99 time layout contains twelve slots (7,9,11,1,3,5,7,9,11,1,3,5). The controlled sampling plan governs whether and when these slots apply. Enter a sampling-event reference linking Section B measurements instead of a bare untraceable tick.

| Shift / date | 07 | 09 | 11 | 13 | 15 | 17 | 19 | 21 | 23 | 01 | 03 | 05 | Record ref / remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Actual / due (approved plan) | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | __________ |
| Sample/event IDs | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | __________ |
| QC result / exception IDs | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | ___ | __________ |

# D  Exception / Approval / Handover

Failure / defect / measurement out of limit: __________________________________________

Affected lot/qty and segregation / HOLD tag ID: ______________________________________

Reject / NCR / rework reference and approved disposition: ______________________________

Original failed result retained / retest authority reference: ______________________________

In-process stage: ☐ Hold  ☐ Ready for next production stage  ☐ Pending review

Downstream packaging transfer / evidence ID (if available): _____________________________

QC checked by / date / shift: ______________________________________________________

QC supervisor reviewed by / date: __________________________________________________

QCM/QA approval only where required by authorization matrix: __________________________

CONTROL NOTICE: This editable Word reconstruction is not an approved F-823-i99 revision. No filled results or authorizations are represented. The technical values must come from the signed product specification.