# ISO 23500-4 | REQUIREMENT & CLOSURE REGISTER

ISO 23500-4:2019  |  WI-8-2-2-40  |  In-Process Control

Requirement coverage includes applicable, conditional, excluded and informative sections with explicit reasons.

SOURCE REVIEWED / CONTROLLED DRAFT - not approved for production or lot release

# Evidence and closure rules

Required evidence is not available evidence. Five provided files are one ISO technical PDF, a WI and three blank in-process templates. They do not demonstrate successful lab/product testing, approved artwork, source material certificates, equipment qualification or a released batch. Verified Closed count = 0. Proposed documentary revisions are not proof of implementation or operational conformity.

Assessment: 23 requirement families, 19 identified gaps. All applicable/conditional items involving missing technical evidence remain Open - Evidence/Technical Input/Approval. No requirement is unconditionally declared compliant.

# ISO requirement coverage and outstanding evidence register

| ID | ISO clause / page | ISO requirement (paraphrased) | Applicability | Gap ID(s) | Proposed documentary route | Actual verified evidence | Operational closure | Required open action |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-01 | 4.1.1 & 5.2.1 PDF pp.8/12; printed 2/6 | Dry-powder form/visual state | Applicable | GAP-001, GAP-002 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Appearance in both powder forms; physical state partially represented |
| REQ-02 | 4.1.1.1-4.1.1.2; 5.2.2 PDF pp.8-9/13; printed 2-3/7 | Reconstituted composition/solute analytical verification | Applicable | GAP-002, GAP-012 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Study and product-specific IFU absent |
| REQ-03 | 4.1.2; 5.2.3 PDF pp.9/13; printed 3/7 | Quality of water used at manufacturer | Conditional | GAP-010 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Dry powder filling only appears in WI; factory use of water to prepare concentrate not established |
| REQ-04 | 4.1.3.2; 5.2.4 PDF pp.9/13-14; printed 3/7-8 | Microbiological quality of powder under labelled use | Applicable | GAP-003 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Qualification/monitoring records unprovided |
| REQ-05 | 4.1.4; 5.2.5 PDF pp.9/14; printed 3/8 | Endotoxin assurance and process validation | Applicable | GAP-004 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | No verified endotoxin evidence |
| REQ-06 | 4.1.5; 5.2.6 PDF pp.9/14; printed 3/8 | Net fill per correct powder usage category | Applicable | GAP-006 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Category/specified tolerance unresolved |
| REQ-07 | 4.1.6; 5.2.7 PDF pp.9/14; printed 3/8 | Chemical/pharmacopoeial grade | Applicable | GAP-001 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Raw chemistry COA not supplied |
| REQ-08 | 4.1.7; 5.2.8 PDF pp.10/14; printed 4/8 | 1 µm filtration of aqueous concentrate | Not Applicable - dry form as reviewed | - | Revised WI + In-Process templates (DRAFT) | Applicability justification only; confirmed approved SKU/portfolio master pending | N/A - provisional with stated product-scope justification | Only applies to aqueous concentrate; cannot impose 1 µm on dry-powder cartridges |
| REQ-09 | 4.1.8; 5.2.9 PDF pp.10/15; printed 4/9 | Additives/spikes | Conditional / not documented | - | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | No additives identified; verify SKU range in controlled product master |
| REQ-10 | 4.1.9; 5.2.10 PDF pp.10/15; printed 4/9 | Container compatibility/closure | Applicable | GAP-007,GAP-015,GAP-019 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Routine seal observations do not prove shelf-life |
| REQ-11 | 4.1.10; 5.2.11 PDF pp.10/15; printed 4/9 | Bulk concentrate shipments | Not Applicable - reviewed packaged SKUs | - | Revised WI + In-Process templates (DRAFT) | Applicability justification only; confirmed approved SKU/portfolio master pending | N/A - provisional with stated product-scope justification | No bulk-delivered concentrate documented |
| REQ-12 | 4.1.11; 5.2.12 PDF pp.10/15; printed 4/9 | Concentrate-generator function/retention | Conditional / likely selected SKUs | GAP-005,GAP-011,GAP-016 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Classification needed per SKU |
| REQ-13 | 4.2; 5.3 PDF pp.10/15; printed 4/9 | Product-contact equipment compatibility | Applicable | GAP-008 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Equipment BOM/cleaning evidence unprovided |
| REQ-14 | 4.3.1-4.3.8; 5.4 PDF pp.10-16; printed 4-10 | Mixers/UV/tanks/piping installed at dialysis facility | Not Applicable - manufacturer scope reviewed | - | Revised WI + In-Process templates (DRAFT) | Applicability justification only; confirmed approved SKU/portfolio master pending | N/A - provisional with stated product-scope justification | Clauses govern dialysis facility mixing/distribution systems, not dry-powder line unless such equipment supplied |
| REQ-15 | 5.1 PDF p.12; printed 6 | Appropriate verified test methods | Applicable | GAP-017 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Sampling/validation records unprovided |
| REQ-16 | 6.1-6.2 PDF pp.16-17; printed 10-11 | All applicable general labelling | Applicable | GAP-009 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Actual artwork/IFU not supplied |
| REQ-17 | 6.3 PDF p.18; printed 12 | Liquid concentrate labelling | Not Applicable - dry SKUs reviewed | - | Revised WI + In-Process templates (DRAFT) | Applicability justification only; confirmed approved SKU/portfolio master pending | N/A - provisional with stated product-scope justification | Do not impose liquid concentrate wording without liquid SKU |
| REQ-18 | 6.4 PDF pp.18-19; printed 12-13 | Powder concentrate labels/IFU | Applicable | GAP-010 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | No approved powder IFU supplied |
| REQ-19 | 6.5 PDF p.19; printed 13 | Additives labelling | Conditional / not documented | - | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | No additive SKU supplied |
| REQ-20 | 6.6 PDF p.19; printed 13 | Generator-specific labelling | Conditional / likely selected SKUs | GAP-011 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Needs SKU classification and compatible machine evidence |
| REQ-21 | 6.7 PDF pp.20-22; printed 14-16 | Mixing-system labelling/literature | Not Applicable - no such supplied system | - | Revised WI + In-Process templates (DRAFT) | Applicability justification only; confirmed approved SKU/portfolio master pending | N/A - provisional with stated product-scope justification | Only if company supplies concentrate mixing systems |
| REQ-22 | Annex A PDF pp.22-28; printed 16-22 | Informative rationale, not independent mandatory controls | Informative | - | Revised WI + In-Process templates (DRAFT) | Annex A read for rationale; no operational claim | Not a normative closure row | Used for interpretation only; no standalone findings |
| REQ-23 | Company process/document control WI pp.5-17 and three F-823 forms | Validated equipment settings, traceable controlled forms | Applicable - company controls | GAP-013,GAP-014,GAP-016,GAP-017,GAP-018,GAP-019 | Revised WI + In-Process templates (DRAFT) | ISO audit source and WI/F-823 drafts present; actual executed approval, study or test evidence NOT SUPPLIED | Open - Technical Input / Evidence | Technical reconciliation and issued records needed |

# Gap-specific closure decision / owner

| Gap | Priority | Requirement link | Responsible function | Evidence necessary (NOT PROVIDED unless explicitly stated) | Closure status |
| --- | --- | --- | --- | --- | --- |
| GAP-001 | P1 | ISO 4.1.6 | Supplier Quality / QC Laboratory | Required: approved material specification, monograph matrix, qualified COA/actual analytical tests, incoming release record. Route: Incoming / Supplier. Actual evidence: not supplied. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-002 | P1 | ISO 4.1.1.2 | R&D / QC Laboratory / QA | Required: product composition specs, controlled IFU, analytical study/raw lab results, qualified methods and applicability matrix. Route: Qualification / Periodic / Laboratory. Actual evidence: not supplied. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-003 | P1 | ISO 4.1.3.2 | Microbiology Laboratory / QA | Required: qualified microbial test reports, validation/justified methods, acceptance criteria, labelled-use evidence and change-control record. Route: Validation / Periodic. Evidence: none supplied. Closure: Open - Evidence and Technical Input. | Open - Technical Input / Evidence |
| GAP-004 | P1 | ISO 4.1.4 | QA Validation / QC Laboratory | Required: validated endotoxin study, related laboratory records and process qualification, change control. Route: Qualification / Risk-based periodic. Evidence not supplied. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-005 | P1 | ISO 4.1.11 | Engineering / QA Validation | Required: design dossier, intended-device functional study, powder-retention results, matrix and approval. Route: Qualification / Change control. Evidence not provided. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-006 | P1 | ISO 4.1.5 | Process Engineering / QC / QA | Required: qualified fill tolerances by application, calibration, approved setup specs and executed mass records. Route: In-process + Qualification. No approved matrix or completed records supplied. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-007 | P1 | ISO 4.1.9 | Engineering / QA Validation | Required: materials compatibility/biocompatibility rationale, closure integrity/transport/shelf-life study where applicable and executed seal results. Route: Qualification + In-process. Evidence absent. Closure: Open - Evidence. | Open - Evidence / Approval |
| GAP-008 | P2 | ISO 4.2 | Engineering / Production / QA | Required: equipment BOM/contact-material declarations, risk assessment, cleaning qualification and change records. Route: Qualification / Line clearance. Not supplied. Closure: Open - Evidence. | Open - Evidence / Approval |
| GAP-009 | P1 | ISO 6.1 | Regulatory / QA / QC | Required: physical or controlled artwork/IFU, completed ISO label matrix, approved changes and in-process sample label check. Route: Document qualification + In-process. Not supplied. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-010 | P1 | ISO 6.4 | Regulatory / QA | Required: approved IFU/powder artwork, storage/stability studies, use-condition evidence and label reconciliation. Route: Qualification / Controlled labelling. Evidence not provided. Closure: Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-011 | P1 | ISO 6.6 | Engineering / Regulatory | Required: SKU classification, machine compatibility/flow/time study, actual approved label/IFU and authorisation. Route: Design qualification / Label master. Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-012 | P1 | ISO 4.1.1.1/4.1.1.2 | QC Laboratory / QA | Required: approved method and product-specific specification, conductivity-composition correlation if used, executed results. Route: Lab qualification / defined monitoring. Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-013 | P1 | ISO 4.1.9 and 4.1.11 | Process Engineering / QA | Required: equipment manual, validated parameter master, reviewed deviation impact, qualified setup/pressure evidence. Route: Engineering / In-process. Open - Technical Input. | Open - Technical Input / Evidence |
| GAP-014 | P1 | ISO 4.1.9 | Production Engineering / QA Validation | Required: equipment recipe sheets, OQ/process qualification, approved reconciliation, executed setup records. Route: Process validation / In-process. Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-015 | P2 | ISO 4.1.9 | QC / Production | Required: completed leak records and calibration; qualified method and approved sample plan. Route: In-process. Evidence absent. Closure: Documentary Covered - Operational Evidence Pending (once approved). | Open - Evidence / Approval |
| GAP-016 | P2 | ISO 4.1.11 / 4.1.9 | Engineering / QC | Required: controlled TF-40-PS/PDBM/drawings, calibration and real dimensional logs. Route: Component In-process / qualification. Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-017 | P2 | ISO 5.1 | QC / QA | Required: controlled sampling/inspection plan, sampled records and evidence of training. Route: In-process / QC. Open - Technical Input and Evidence. | Open - Technical Input / Evidence |
| GAP-018 | P2 | ISO 23500-4 provides technical requirements, not the company f | Document Control / QA | Required: master document list, approved revision/supersession history, missing referenced forms incl. 4008/control records and reconciliation decision. Route: Document Control. Open - Approval and Evidence. | Open - Evidence / Approval |
| GAP-019 | P2 | ISO 4.1.9, 4.1.11 and 6.1-6.6 | QA / QC Manager | Required: executed lots, deviations, approved QC/QA release records and traceable evidence register. Route: In-process / Final release. Open - Evidence. | Open - Evidence / Approval |

# Approval sequence and management decision

QC Manager / QA must first approve controlled SKU and applicability matrix, missing technical values, source-document revision reconciliation, risk-based sampling/control route and qualification evidence. Then the rebuilt WI and each F-823 form can be reviewed, issued and used. Only after implementation, actual test execution or authorised external evidence, verification and quality disposition can an item move to Verified Closed.

No historical inspection result, supplier COA, calibration record, clinical compatibility demonstration or release approval has been invented in this register.