# CONTROLLED RECONSTRUCTION DRAFT  |  WI-8-2-2-40  |  PROPOSED REVISION: TBD  |  SOURCE HEADER: REV.14 / FILENAME: REV.13

This is a reconstructed standalone WI based on the full supplied 17-page source, retaining existing processes and controlled company criteria unless explicitly placed ON HOLD pending technical input. References to company sign-off and historical settings do not constitute approval of this draft.

- Purpose:

The purpose of this document to stipulate the production procedure for manufacturing of Sodium Bicarbonate cartridge and additional components used along with cartridge to ensure producing good quality products during production process.

- Scope

This procedure is applicable for manufacturing of all types of sodium bi carbonate cartridges and manufacturing components for sodium bi carbonate, sodium bi carbonate filling process and labelling process by Saudi Mais company.

This procedure applicable for the sodium bi carbonate cartridges supporting for different types of dialysis machines.

- Sodium Bi carbonate Cartridge for Gambro Machine (Althin, Baxter Miroclav, Braun Dialog, Gambro AK100, AK200 Hospal, Monitral, Hospal Intergra, Nikkiso DBB-03, DBB-05

- Fresenius 5008 Dialysis system.

- Fresenius 4008S Dialysis system.

- Definition:

  - Sodium Bicarbonate: A white fine granular material used for neutralizing the conductivity of hemodialyzer during the hemodialysis process.

  - Gambro Cartridges 650g / 720g / 750g: A vertical cylindrical shaped Cartridge.

  - Fresenius 5008 Cartridges 650g /900g: A vertical cylindrical shaped Cartridge.

  - Fresenius 4008 bag 650g:  A vertical bag with neck holder.

  - Fresenius 5008 650g / 720g/900g: A vertical bag with holder.

- Responsibilities:

Bulk Store: Store Keeper is responsible to receive and store the Raw Materials, and Components used for manufacturing of Sodium bi carbonate.

Production Department: Production Department is responsible for issuing and processing the Raw Materials, manufacturing cartridge and components, filling of sodium bi carbonate cartridge, labelling and packaging.

Maintenance Department: Maintenance Department is responsible for taking Corrective and Preventive actions during the Manufacturing process.

Finished Goods Store: is responsible to receive and store the Finished Products in a warehouse and release the products when received conformed delivery request.

- Procedure.

  - Receiving Sodium Bi Carbonate Granules material from supplier.

    - Bulk store incharge is responsible to receive shipment of sodium bi carbonate Granules material and make sure that the received sodium bi carbonate is same as per purchase order and expiry date is valid.

    - Inspect the material bag condition, must be no damaged of all bags with material and make sure that the bag should be free from damage and holes.

    - Quality controller is responsible to take samples and submit to chemical lab and make sure that the results are same as per COA provided by supplier.

    - Record results into receiving inspection report.

    - Inspect the material bag condition, must be no damaged at all bags of material and make sure that the bag should be free from damage and holes.

    - Quality controller is responsible to take samples and submit to chemical lab for further testing.

    - Inspection and test report for Raw Material.

| Description | Requirement | Status |
| --- | --- | --- |
| Condition of Bag / Appearance | Good without damage bag | Visual |
| Material Grade and description | As per P.O. | Visual |
| Appearance of Granules | Good / Transparent / colored | Visual / Physical |
| Granules Formation | No lobes formation | Visual |
| Lot no printed status | Clear printed | Visual |
| Arrangement of material bag | Should be specified and clear label | Visual |

    - Record results into receiving inspection report.

  - Receiving Filter Fabric (80micron) and (50micron) from supplier.

    - Bulk store incharge is responsible to receive shipment of filter fabric roll and make sure that the received fabric is same as per purchase order.

    - Inspect the roll condition, must be no damaged make sure that fabric should be free from damage, holes and stains.

    - Quality controller is responsible to take samples and make sure that the fabric thickness is same as required 80 micron and 50 micron and record the results in specified receiving inspection form.

  - Receiving bag Sheet Roll from supplier for sodium bi carbonate bag type.

    - Bulk store incharge is responsible to receive shipment of PVC bag sheet roll and make sure that the received bag sheet roll is same as per purchase order.

    - Inspect the roll condition, must be no damaged make sure that bag sheet should be transparent free from damage, scratches, holes and stains and accurate in dimension as per purchase order.

    - Record the results in a specified receiving inspection form.

  - Receiving Bulk components (Black O-ring, / white silicon round (O-ring) / Porous PE Filter.

    - Ensure the received components from supplier are correct as per purchase order.

    - Inspect the carton condition, must be in good condition, no damaged at all cartons.

    - Take samples & inspect in order to AQL general inspection procedure, product description, specification with apply the AQL sampling plans for normal inspection.

    - Record into OBM inspection report.

  - Visual Inspection.

- Ensure that the received material is same as written in the purchase order.

- Inspect the packaging for any deterioration during transportation.

- There must be no visible contamination on the products.

- The granules must be small size, pure white, and no lumps formation.

- The sticker must be clear printed, accurately marking and not in washout.

- The product / product part (s) should be clear, organize and free from any visible contamination and irregularity.

- The Sodium Bicarbonate granules should be pure and not contaminated with any other dust particles

or lumps formation and black spot. Dimension must be accurately in order to quality specification, formation, etc.

## 5.6  Controlled Material Qualification, Reconstitution and ISO Evidence (PROPOSED)

• Incoming sodium bicarbonate shall be released only against approved monograph and material-code requirements using verified COA/test evidence, qualification and QC/QA incoming disposition. Controlled reference: applicable pharmacopeia and approved material specification. GAP-001; ISO 4.1.6/5.2.7. OPEN TECHNICAL INPUT REQUIRED: monograph edition/limits and approved testing/sampling plan.

• For each powder SKU maintain authorised IFU reconstitution/water method, analytical final-concentrate composition plan and controlled acceptance criteria linked to ISO 4.1.1.2/5.2.2.2. Qualified solute analysis must use controlled methods. Neither generic appearance nor a generic conductivity range proves chemical conformity. GAP-002/012. OPEN TECHNICAL INPUT REQUIRED: formulation, concentration/conditions and actual test results.

• Use controlled qualification/monitoring plans to demonstrate packaged powder can produce dialysis fluid meeting ISO 23500-5 microbiological criteria and endotoxin requirements under intended use. Methods, limits, frequency and results shall be documented in approved separate records and linked by ID, revision and scope; no automatic per-lot test is imposed. GAP-003/004. OPEN TECHNICAL INPUT REQUIRED: validated protocols, methods and limits.

• If manufacturing uses water to make concentrate, verify ISO 23500-3 water-quality control; otherwise document the dry process and ensure the approved IFU states user mixing water. ISO 4.1.2; conditional applicability.

- Production Process

  - Manufacturing PVC Material for Blue Lock Adaptor.

    - Production controllers are responsible to issue raw material chemicals and issue production order for manufacturing of PVC granules.

    - Machine operator is responsible to make sure that all stations of machine are properly working and follow the procedure WI-8-2-7-01 (Production and Test Procedure for PVC Granule’s) for manufacturing of PVC granules.

    - Quality controller is responsible to take samples as per AQL general sampling method and perform visual inspection of material, check material hardness test and submit sample to analytical lab for material density test.

    - After received results from analytical lab record the results in specified material Inprocess control form.

    - After completed production order machine operators are responsible to close the bag and transfer to specified area.

    - Production controllers are responsible to make transfer slip and transfer to store.

  - Injection Machine Procedure (Manufacturing Components)

- The injection molding components for sodium bi carbonates are.

- BC Body cartridge

- BC Body cover

- BC Body cap

- BC Body holder for cartridge

- Holder Insert part

- Bag Holder

- Bag Holder cap

- Neck Holder for bag

- Neck Holder cap for bag

- Filter for Cartridge

- Production controllers are responsible to issue required raw materials for components and make sure that the received material is same as per issue slip and issue production order for manufacturing components.

- Maintenance department is responsible to set the required mold for components on specified machine and set the standard parameters.

- Quality controller is responsible to verify the parameters and make sure that the required standard parameters are set.

- Machine operators are responsible to fill the material into machine hopper by using vacuum pump and fill hoppers up to the required capacity.

- Quality controller is responsible to take sample as per AQL general inspection method for visual inspection and AQL special inspection method for physical testing.

- BC Body Cartridge

- Quality controller is responsible to make sure that the BC body cartridge should be clear and transparent, no scratches marks, no black spots, no crack, no damage and clear and smooth edges, no extra fibers.

- BC body Cover

- Quality controller is responsible to make sure that the BC body cover should be clear and transparent, no scratches marks, no black spots, no crack, no damage and clear and smooth edges, no extra fibers and inlet and outlet channels should not block.

- BC Body Cap

- Quality controller is responsible to make sure that the BC body cap should be clear and transparent, no scratches marks, no black spots, no crack, no damage and clear and smooth edges good fitting with cartridge, no loose.

- BC Body holder for cartridge

- Quality controller is responsible to make sure that the BC body holder clear and no black spots, no crack, no damage and clear and smooth edges good fitting with cartridge, open inlet and outlet channels and holding should not damage.

- Holder Insert part

- Quality controller is responsible to make sure that the holder insert part clear and no black spots, pencil type, no crack, no damage and clear and smooth edges can be insert properly in holder.

- Neck holder for bag

- Quality controller is responsible to make sure that the neck holder for bag should be clear and no black spots, no crack, no damage and clear and smooth edges, no blockage.

- Neck holder cap for bag

- Quality controller is responsible to make sure that the neck holder cap should be clear and no black spots, no crack, no damage and clear and smooth edges, good fitting with neck holder, no loose.

- Filter for Cartridge

- Quality controller is responsible to make sure that the filter should be clear and no black spots, no crack, no damage and clear and smooth edges, good fitting in cartridge and no loose.

  - Manufacturing process of Filters 50 microns and 80-microns

- Production controller is responsible to issue required raw material and fabric for making filters 50 and 80 microns.

- Maintenance department is responsible to set the required mold and set approved parameters.

- Quality controller is responsible to verify the parameters and make sure that the parameters should be as per approved parameters.

- Machine operators are responsible to place fabric in mold and press the button to start machine.

- Quality controller is responsible to make sure that the filter housing should be round circle shape clear and smooth edges, no black spots, no crack, no damage and clear and smooth edges and fabric should be proper assembled with housing and no pullout, no gap and record the results in specified Inprocess control form.

- Quality controller is responsible approve the and place sticker on bag.

- Machine operator is responsible to arrange good products in clean bag and transfer to next process for cutting.

    - Filter Cutting Machine 1 & 2.

- Production controller is responsible to arrange all filters for cutting.

- Maintenance department is responsible to set the required approved heating parameters for cutting filters.

| Machine. # 01(A) | Machine. #02 (B) |
| --- | --- |
| 220 - 240°C | 320 - 350°C |

- Place filters in machines die and press button and repeat process for cutting each filter.

- Quality controller is responsible to make sure that filters should be smooth cutting edges, no extra fibers on external surface of housing and record the results in specified form.

- Machine operator is responsible to place good quantity in clean PE bag and place approved inspection tag.

- Production controller is responsible to arrange in specified storage area.

## 6.4  Component, Contact-Material and Equipment Controls (PROPOSED)

• Before authorised production, QC shall verify approved drawings and applicable dimensions, actual sample values and cavity/tube lot, instrument ID and criteria. The raw component template F-823-i15 is the In-Process Control record, not an ISO laboratory test report. GAP-016.

• Engineering/QA shall document material compatibility of all product-contact surfaces and approved change/cleaning status for granule filling and assembly. Record equipment/assessment IDs without claiming any particular hazardous material is installed. GAP-008; ISO 4.2.

- Extrusion Machine Procedure.

  - Process for Manufacturing Filter Tube.

- Machine operator is responsible to fill the machine hopper with material and make sure that the material is same as per production order.

- Maintenance operators are responsible to set the required die for tube and set the standard parameters.

- Machine operator is responsible to clean the machine surface and make sure each station of machine should be clean and dust free.

- Quality controller will take sample as per AQL piece/piece, if there is no any problem then machine run.

- Quality controller is responsible to inspect the tube sample against specification and make sure that the tube is as per standard specification.

    - Filter Tube

- Quality controller is responsible to make sure that the tube should be smooth surface, clear and accurate in dimension in length ID and OD as per specification and no black spots, no crack marks, no damage and blockage.

- Every 2 hours QC will take samples from produced quantity base on AQL samples and check as per specification.

- Machine operator is responsible to arrange the good products in clean transparent polyethylene bag and make inspection tag and quality controller will approve.

- Production controller is responsible to arrange the good products in specified rack or transfer to mais components storage facility.

## 7.2  Sampling, Measurement and Inspection Traceability (PROPOSED)

• Identify the controlling approved sampling-plan ID, quantity, shift/time, sample/cavity, measured values and acceptance decision. Preserve §7.1.1 two-hour tube inspection unless approved controlled plan changes it; AQL cannot substitute for documenting an applicable sampling table. GAP-017.

• Segregate nonconforming items and any affected production interval when out-of-control results occur. Record NCR, affected lot, containment, investigation and authorised disposition; do not treat passing subsequent samples as automatic lot release. GAP-019.

- Assembly Process for Gambro Cartridges.

The process done by automatic welding machine by the following given stages.

- Granules material Filling

- Cartridge Filling

- Cartridge Sealing

## 8.0  Gambro Configuration / Fill / Leak / Recipe Reconciliation (PROPOSED)

• Verify SKU/machine and intended-use classification and approved fill net mass/tolerance matrix. The historic nominal fill values and ±10 g must not be interpreted as universally required by ISO 4.1.5. Record actual mass in grams using a calibrated/tared balance, as defined by QA. GAP-006.

• AP32 entry historically shows A=0.760 mbar (±0.2) while WI and manual leak test show 0.8 bar (±0.2). These are not reconciled. No operator shall rely on the ambiguous sensor entry as an approved parameter until Engineering/QA signs off exact physical units, tolerance and recipe. GAP-013. OPEN TECHNICAL INPUT REQUIRED.

• Existing Gambro weld time is 1.00-1.40 s in WI versus 1.0-1.50 s in F-823-i14-2. Follow approved machine recipe only after technical reconciliation; monitor actual readings and test evidence. GAP-014. OPEN TECHNICAL INPUT REQUIRED.

• Retain WI §8.2 company immersion test 0.8 bar ±0.2 bar for 8-15 s pending existing validation confirmation; record actual pressure, time, instrument, bubble observations and sample ID. This value is not set by ISO. GAP-015.

• Verify compatibility with defined Gambro equipment and that the filter path prevents undissolved powder reaching the fluid stream by approved functional studies. Routine checks confirm correct filter position and qualified configuration. GAP-005.

  - King Ultrasonic Automation Machine Welding procedure for (Gambro Type 650g / 720g / 750g Cartridges)

- Before starting machine, production controller responsible to set all required Parameters and make sure that machine and all surrounding area must be clean and previous batch should be properly removed.

- Make sure that the machine pressure should be set on required range 5.5 – 6.5 bar.

- The sealing die should be properly mounted and perfectly alignment without gap.

- Make sure that die must be free from any Granules material granules.

- Fill the machine hopper with required sodium bi carbonate Granules material up to the required capacity.

- Machine operator is responsible to fill the machine specified liner with body of Gambro cartridge and cover and make sure that cartridge and cover should be with proper inserted filters.

- Set the required approved parameters for cartridge sealing and set the required filling Granules material in gram for filling cartridge.

- Quality controller is responsible to verify the parameters and make sure that the parameters are set as per approved parameters.

  - Parameters for (Sodium Bi Carbonate Cartridge Gambro Type)

| Description | Parameters |
| --- | --- |
| Charging Time on Cartridge | 1.0 – 3.0Seconds |
| Filling Time for Sodium Granules material | 1.40 – 1.70 Seconds (650 Gram) 1.50 - 1.90 Second (750 Grams) |
| Leak Test on sealed Time | 1.0 - 3.0 Second |
| Cycle Time | 15 – 30 Second |
| Delayed Time (Sealing Cycle) | 1.00 – 1.50Sec |
| Holding Time (Sealing Cycle) | 1.0 – 1.50Sec |
| Weld Time (Sealing Cycle) | 1.00 – 1.40Sec |
| Pressure Gauge setting 0.8bar (± 0.2) Sensor Key: AP32 A = 0.760mbar (± 0.2) | Vol. 750 Grams |
| Pressure Gauge setting 0.8bar (± 0.2) Sensor Key: AP32 A = 0.760mbar (± 0.2) | Vol. 650 Grams |
| Manual leak test | - bar (± 0.2) |

  - Machine operator is responsible to start the machine, it will automatically fill the Granules material as per required weight and transfer to next station for cover assembly and final station for sealing.

  - Leak test performed by machine by applied pressure 0.8bar±0.2bar if cartridge leakage generate alarm.

  - Machine operator is responsible to monitor machine parameters, sodium bi carbonate Granules material and components stock should be available and no deviation in parameters during process.

  - Quality controller is responsible to get sample as per AQL special inspection method and make sure that the filters should be proper inserted in cartridge body and cover and make sure cover should be proper sealed with cartridge body, no extra material fibers, no crack, no damage and make sure the Granules material should not across or pass on filter side when by shaking the product and cover has no sink or streak or canal line on the cover and record the results in specified form.

  - Leak Testing (Sodium bi Carbonate Cartridge Gambro Type)

  - After filling and sealing cartridges production controllers are responsible to transfer the final filled and sealed cartridges for manual leak test.

  - Operators are responsible to set the required parameters (0.8 bar ± 0.2 bar pressure) for leak test as per approved parameters as Gambro cartridges leak test parameters are mentioned in above table (Parameters for sodium bicarbonate Gambro Type.)

  - Connect required tubing with cartridge insert the steel adaptor into the cartridge outlet and immerse the cartridge into water and apply required pressure (0.8 bar ± 0.2 for 8 – 15 second and observe any bubble appear from tubing connection site or edges of the sealed cover.

  - Quality controller is responsible to get sample from final approved products as per AQL sampling plan and retest cartridges leakage test to make sure cartridges are leakage free.

  - Observe the air bubbles and if bubbles appeared from the sealing area considered Rejected and if no bubbles, Acceptable.

Acceptance Criteria

No leakage from sealed area, Acceptable

Conductivity (historical company criterion): 13.5-14 mS/cm appears in source WI. HOLD for technical reconciliation. Apply only an approved SKU-specific tested solution specification, dilution, water quality, temperature/method and QC result field; this is NOT a universal ISO 23500-4 acceptance limit. OPEN TECHNICAL INPUT REQUIRED. [GAP-012]

Homogeneous sealed no leakage, Acceptable

No leakage from cover sealed area and from insert part, Acceptable

  - Fix cartridge cover cap with cover and bottom side and make sure cover cap no loose.

  - Production controller is responsible to transfer good products to next process for labelling.

  - Labelling Process (Sodium bi Carbonate Cartridge Gambro Type)

- Production controller is responsible to issue required sticker for Gambro sodium bi carbonate cartridge and print Product reference, lot number, production date, expiry date, UDI QR Code and any other tender information on labels.

- Quality controller is responsible to make sure label sticker information should be as per approved labels and as per packaging order and all other information’s are correct.

- Clean cartridge by dry cloth and make sure cartridge should be free form dust particles before placing labels.

- Remove sticker from paper and fix on the middle of the cartridge and make sure that the sticker should be proper adhesive on cartridge body and sticker should be bubble free.

- Production controller responsible to arrange in specified space and transfer to tertiary packaging section for final packaging.

  - Tertiary packaging (Sodium bi carbonate cartridge Gambro type)

- Production controller is responsible to issue production order for sodium bi carbonate tertiary packaging and make tertiary packaging sticker.

- Quality controller is responsible to make sure that the sticker information’s are correct as per approved labels and production order.

- Place cartridge one by one in tertiary packaging in opposite direction and make sure that cartridge should be in horizontal position in tertiary packaging.

- Close carton with tap and place approved sticker by glue.

- Production controller is responsible to arrange final finished products in as designated place.

## 8.5  Gambro Labelling and Release Linkage (PROPOSED)

• Correct printed fields and UDI must be verified against controlled SKU artwork. ISO 6.2/6.4 and where appropriate 6.6 labelling is qualified through actual artwork/IFU and a clause-by-clause checklist, including composition, warnings, machine compatibility and capacity where relevant. GAP-009/010/011.

• Routine in-process forms shall link current label master, qualification records and exception status but are not final release forms; final release remains governed by approved F-823-F10 and company SOP. GAP-019.

- Assembly Process for Fresenius all Types of cartridges.

- The process for manufacturing Fresenius cartridges by the following given stages.

- Assembly of Porous PE filter with tube and Cover.

- Assembly of blue lock adaptor and insert part with holder.

- Bicarbonate Granules material loader process.

- Cartridge Filling process.

- Cartridge Sealing process.

- Holder final assembly with cartridge by using black O ring and white flat O ring.

  - Assembly of Porous PE Filters with Tube and Cover.

- Production controllers are responsible to issue required tube and filters and issue production order for assembly.

- After receiving filters and tube make sure that the filters and tubes should be same as per issue slip and should be with proper identification labels.

- Cut tube in required length as per below table and hold tube and insert it into the filter hole and make sure tube should be proper inserted up to the required insertion hole of filter and other end of tube inset into the outlet channel of cartridge cover and tube should not pullout from filter and cover.

| Tube (ID x OD) (± 0.1 mm) | Tube Length (cm) (±1.0) | Filter ID x OD (± 0.1 mm) |
| --- | --- | --- |
| 2.4 – 3.7 mm | 16.9 cm | - x 6.4 mm |

- Quality controller is responsible to make sure that the tube proper fixed with filter and cover and tube should not pullout from filter and cover and record the results in specified form.

  - Assembly of blue lock adaptor and insert part with holder.

- Production controllers are responsible to issue required holder, blue lock adaptor and insert part and issue production order for assembly.

- After receiving components make sure that the components should be same as per issue slip and should be with proper identification labels.

    - Assembly of Insert Part with Holder by machine

- Place holder on clean assembly table and holder holes should be Infront of face direction, dip insert part in glue and insert into the holder holes and place in machine die then rotate machine handle to press the insert part into holder and make sure the insert part should be proper inserted in holder holes and no pullout from holes.

- Clean holder by dry cloth and make sure that the holder should be free from extra gluing drops.

- Quality controller is responsible to make sure that the insert part should not pullout from holder and record the results in specified form.

    - Blue Lock Adaptor Assembly by machine

- Place holder on clean assembly table and channels should be upward direction, fix blue lock adaptor to inlet and outlet port of holder, place holder in machine die and rotate handle of machine to press the lock adaptor and make sure the adaptor should be strong attached with holder inlet and outlet port no pullout from ports.

- Quality controller is responsible to make sure that the adaptors should not pullout from holder inlet and outlet ports and record the results in specified form.

  - Assembly of Porous PE Filters with Tube and Cover.

- Production controllers are responsible to issue required tube and filters and issue production order for assembly.

- After receiving filters and tube make sure that the filters and tubes should be same as per issue slip and should be with proper identification labels.

- Hold tube and insert it into the filter hole and make sure tube should be proper inserted up to the required insertion hole of filter and other end of tube inset into the outlet channel of cartridge cover and tube should not pullout from filter and cover and make sure 50 micron filer inserted in cover inlet channel.

- Quality controller is responsible to make sure that the tube proper fixed with filter and cover and tube should not pullout from filter and cover and record the results in specified form.

  - Bicarbonate Granules material Loading Process.

- Production controller is responsible to issue required sodium bi carbonate Granules material and make sure that the bag should be free from dust particles after receiving from bulk store.

- Transfer sodium bi carbonate Granules material to loading area and open all bags of materials and dispense into machine hopper up to the required capacity and close hopper by cover after filled sodium bi carbonate to required capacity.

  - Filling machine process for filling cartridges for Fresenius type.

- Machine operator is responsible to make sure that machine surface should be clean and dust free and set the approved parameters for required weight for filling cartridge of sodium bi carbonate Granules material.

- Machine operator is responsible to make sure that sensor of machine should be in working condition before start filling process and filled weight of sodium bi carbonate Granules material should be accurate as required.

- Place assemble cover with Porous PE filter in cartridge and place cartridge below the sensor of machine and machine will fill cartridge automatically the required weight and fix cover with cartridge and repeat process for each cartridge to fill.

- Quality controller is responsible to make sure that the filled Granules material should be accurate in required weight and record the results in specified form.

- After filling place cartridges in cartridge rack and transfer to next process for sealing.

- All Fresenius cartridges are sealed by Kisonic king Ultrasonic welding machine.

  - Kisonic King Ultrasonic Welding Machine process for (Fresenius5008 – Cartridges

Holder Type)

- Machine operator is responsible to set the required die for welding sodium bi carbonate Fresenius cartridge and set required sealing parameters.

- Kisonic King ultrasonic cartridge cap welding machine parameters are.

| Cycle Time | Parameters |
| --- | --- |
| Delay Time | 0.50 – 1.50 sec |
| Holding Time | 0.50 – 1.50 sec |
| Welding Time | - – 1.50 sec |

- Place cartridge in specified die and press button for start welding.

- Quality controller is responsible to make sure that the cover should be proper sealed with body, no extra fibers and cartridge should not damage and crack.

- Machine operators are responsible repeat the process for each cartridge one by one and after sealing place cartridge in specified cartridge rack.

- Production controller is responsible to transfer the sealed cartridges for next process leakage test by.

9.7     Assembly Holder / Black O ring / White flat O ring with cartridge

- Production controller is responsible to issue required holder, black O ring / white flat O ring for holder assembly with cartridge body.

- Place Black O ring on cartridge inlet and out let port and insert white flat O ring into holder inlet and out let channel then fix holder over cartridge cap by press and apply little force.

- Quality controller is responsible to make sure that holder proper fixed with cartridge no loose.

- Production controller is responsible to transfer cartridges to next process for labelling.

  - Leak Testing (Sodium Bi Carbonate Cartridges Fresenius 5008)

  - After sealing cartridges, production controllers are responsible to transfer the final filled and sealed cartridges for manual leak test.

  - Operators are responsible to set the required parameters (0.6 bar ± 0.2 bar pressure) for leak test as per approved parameters.

  - Connect required tubing with cartridge insert core into the holder inlet and insert steel road into holder outlet and immerse the cartridge into water from cover site and apply required pressure (0.6 bar ± 0.2 for

8 – 15 second and observe any bubble appear from tubing connection site or edges of the sealed cover.

  - Quality controller is responsible to get sample from final approved products as per AQL sampling plan and retest cartridges leakage test to make sure cartridges are leakage free.

  - Observe the air bubbles and if bubbles appeared from the sealing area considered Rejected and if no bubbles appear then it’s Acceptable.

  - Production controller is responsible to transfer good products to next process for fixing holder.

Acceptance Criteria

- No leakage from sealed surface, Acceptable.

- Conductivity (historical company criterion): 13.5-14 mS/cm appears in source WI. HOLD for technical reconciliation. Apply only an approved SKU-specific tested solution specification, dilution, water quality, temperature/method and QC result field; this is NOT a universal ISO 23500-4 acceptance limit. OPEN TECHNICAL INPUT REQUIRED. [GAP-012]

- Homogeneous sealed no leakage, Acceptable.

- No leakage from cover sealed area and from insert part, Acceptable.

  - Labelling process (Sodium bi carbonate Cartridge Fresenius5008 type)

- Production controller is responsible to issue required sticker for Fresenius5008 sodium bi carbonate cartridge and print Product reference, lot number, production date, expiry date, UDI QR Code and any other tender information on labels.

- Quality controller is responsible to make sure label sticker information should be as per approved labels and as per packaging order and all other information’s are correct.

- Clean cartridge by dry cloth and make sure cartridge should be free form dust particles before placing labels.

- Remove sticker from paper and fix on the middle of the cartridge and make sure that the sticker should be proper adhesive on cartridge body and sticker should be bubble free.

- Production controller responsible to arrange in specified space and transfer to tertiary packaging section for final packaging.

  - Tertiary packaging (Sodium Bi Carbonate Cartridge Fresenius 5008 Type)

- Production controller is responsible to issue production order for sodium bi carbonate tertiary packaging and make tertiary packaging sticker.

- Quality controller is responsible to make sure that the sticker information’s are correct as per approved labels and production order.

- Fix internal cartridge supporting material in tertiary packaging and place cartridge one by one in tertiary packaging and make sure that cartridge should be in vertical position in tertiary packaging.

- Close carton with tap and place approved sticker by glue.

- Production controller is responsible to arrange final finished products in as designated place.

## 9.9  Analytical and Machine-Specific Acceptance (PROPOSED)

• For Fresenius cartridge units define approved machines, filter retention and compatibility evidence; verify configuration-specific powder mass and laboratory/reconstitution results through the authorised qualification/periodic plan, not by assuming all tests are hourly. GAP-002/005/006.

• Historical conductivity 13.5-14 mS/cm is not an ISO universal acceptance limit. Approved specification must identify sample solution, preparation, dilution, temperature and method before the historical figure is allowed in acceptance decisions. GAP-012.

- Production Process Fresenius 5008 Bag Type with Holder.

- Production controllers are responsible to issue required film roll for Fresenius bag from store.

- After receiving roll make sure the bag should be free from holes and dust particles and issue production order for making sodium bi carbonate bag for Fresenius 5008.

- Cut the roll in required length and width by manual and make sure the cutting should be straight.

- The dimension of Fresenius bag as in given table.

| Bag dimension | Length (±5.0) mm | Width (±0.1) mm | Thickness (±0.01) mm |
| --- | --- | --- | --- |
| 650/720 | 350 | 280 | 0.31 |
| 900 | 400 | 280 | 0.31 |

- Quality controller is responsible to make sure that the bag dimension should be accurate as required and sheet should be straight cutting edges.

- Production controller is responsible to transfer sheets for next process for sealing.

  - Bag Layers Sealing Process

- High Power Sealing Machine (Fresenius 5008 Bag with holder Type)

Machine operator is responsible to set the required 2 cavity die and make sure die properly fixed in machine and set the approved required parameters for bag sealing.

| Bag Sealing Temperature 3-Sides | Temperature (oC) | Sealing Up Time (Sec) | Sealing Down Time (Sec) | Sealing Welding Time (Sec) | Sealing Hold Time (Sec) |
| --- | --- | --- | --- | --- | --- |
|  | 150 - 170 | 2-3 | 2-3 | 2.5 – 3.5 | - Sec |

- Fix both layers in machine die and make sure layers both sides should be proper match from edge to edge and layers should not short shot.

- Press start button then the machine weld both layers from three sides of bag.

- Quality controller is responsible to make sure that both layers should be proper welded from three sides and no open from welded edges and holder insertion site should be open.

- Repeat process to weld required quantity and make sure both layers should be proper welded.

- After completed welded process production controller is responsible to transfer sheets for next process for bag cutting.

  - Sheet Cutting Machine

- Production controller is responsible to transfer all welded sheets for cutting.

- Machine operator is responsible to set the regulator to 1.0 – 5.0 and make sure that the cutting die should be with proper inserted cutting blade.

- Fix welded sheets on cutting die, invert die and place on machine frame then close the machine.

- Cutting Machine Bag Dimension in given table.

| Bag dimension | Length (±5.0) mm | Width (±1.0) mm | Thickness (±0.01) mm |
| --- | --- | --- | --- |
| 650/720 | 340 | 120 | 0.31 |
| 900 | 390 | 120 | 0.31 |

- Quality controller is responsible to make sure that the bag should be straight cutting edges and no holes.

- Repeat the process and complete the required quantity and then transfer to next process.

  - Assembly of Holder and Tube with Porous PE Filter.

- Production controllers are responsible to issue required tube and filters and issue production order for assembly.

- After receiving filters and tube make sure that the filters and tubes should be same as per issue slip and should be with proper identification labels.

- Cut the tube in required length and hold tube and insert it into the filter hole and make sure tube should be proper inserted up to the required insertion hole of filter and other end of tube inset into the outlet channel of holder and make sure 13mm porous PE filter inserted in holder inlet and tube should not pullout from holder outlet port.

| Tube (ID x OD) (± 0.1 mm) | Tube Length (cm) (±1.0) | Filter ID x OD (± 0.1 mm) |
| --- | --- | --- |
| 2.4 – 3.7 mm | 29 cm | - x 6.4 mm |
| 2.4 – 3.7 mm | 36 cm | 3.4 x 6.4 mm |

- Quality controller is responsible to make sure that the tube proper fixed with filter and holder and tube should not pullout from filter and holder and record the results in specified form.

  - Holder Sealing process with Bag by Sealing Machine 1 and 2.

- Insert holder in bag opening site and make sure bag should be accurate fixed on holder and edges should not above then holder.

- Machine operator is responsible to set the required sealing approved parameters for sealing holder with bag as mentioned in given table.

| Bag and Holder Sealing Position | Sealing Temperature | Sealing Time |
| --- | --- | --- |
| Holder Sealing Upside Temperature | 175 – 210 °C | 08 – 20 sec. |
| Holder Sealing Downside Temperature | - – 240 °C | 08– 20 sec. |

- Insert steel adaptor in holder and fix holder in machine die press start button.

- Machine operator is responsible to make sure that bag should be proper sealed with holder, and no gap in sealing.

- Quality controller is responsible to make sure that holder should be proper sealed with holder without gap in sealing, no burn marks, and free from black spots and record the results in specified form.

- Repeat process for sealing of holder with bag until the required quantity complete.

  - Bag filling process by Sodium Bi Carbonate Granules material (Manual filling)

- Production controller is responsible to fill the container with sodium bi carbonate Granules material up to the required capacity and make sure the storage container should be with label.

- Fill sodium bi carbonate Granules material into bag, and then transfer to bag through holder top side and make sure filled bag should be accurate in weight.

- Quality controller is responsible to make sure that the filled sodium bi carbonate Granules material accurate in weight as required and record the results in specified form.

- Repeat process for filling sodium bi carbonate Granules material ed in bag until the required quantity complete.

- After complete filling production controller is responsible to transfer filled bag for next process holder top side sealing.

  - Bag Holder Sealing Top Side

- Cut the sheet for holder top side sealing and make sure sheet should be accurate in dimension in length and width.

| Description | Length (±1.0cm) | Width (±1.0cm) |
| --- | --- | --- |
| Bag Holer Top sheet for sealing | 11 | 6 |

- Machine operator is responsible to set the required parameters for sheet sealing with holder top side.

- The parameters for holder sealing top side are.

| Description | Sealing Temp. | Sealing Time |
| --- | --- | --- |
| Holder Sealing top side | 200 – 260 °C | 08 – 20 sec. |

- Place sheet on bag holder and make sure that the sheet should be cover holder edges properly and then start machine and wait for 08 – 20second to weld sheet with holder.

- Quality controller is responsible to make sure that the holder proper sealed by sheet and edges should be proper covered by sheet, no gap in sealing, no holes and burn marks.

- Make sure the Granules material  should not across or pass on filter side when by shaking the product and valve has no sink or streak or canal line on the contracting neck in order to control the leakage on dialysis machine.

- If all good / strong bond then waits for 45 minutes or keep the products cold before next process leak test.

  - Leak Testing Process for Sodium Bi carbonate Bag with holder for Fresenius 5008.

  - After sealing bag with holder, production controllers are responsible to transfer the final filled and sealed bag for manual leak test.

  - Operators are responsible to set the required parameters (0.6 bar ± 0.2 bar pressure) for leak test as per approved parameters.

  - Connect required tubing or leak test adaptor, insert core into the holder inlet and insert steel adaptor into holder outlet and immerse the bag into water from holder site and apply required pressure (0.6 bar ± 0.2 for 8 – 15 second and observe any bubble appear from tubing connection, adaptors, holder sealing side or edges of the sealed area.

  - Observe the air bubbles and if bubbles appeared from the sealing area considered Rejected and if no bubbles appear then it’s Acceptable.

  - Quality controller is responsible to get sample from final approved products as per AQL sampling plan and retest cartridges to make sure cartridges are leakage free and record the results in specified form.

Acceptance Criteria

- No leakage from sealed surface and edges, Acceptable

- Conductivity (historical company criterion): 13.5-14 mS/cm appears in source WI. HOLD for technical reconciliation. Apply only an approved SKU-specific tested solution specification, dilution, water quality, temperature/method and QC result field; this is NOT a universal ISO 23500-4 acceptance limit. OPEN TECHNICAL INPUT REQUIRED. [GAP-012]

- Homogeneous sealed no leakage, Acceptable

- No leakage from holder sealed area, Acceptable

- Production controller is responsible to transfer good products to next process for labelling.

  - Labelling process (Sodium bi carbonate Bag with holder Fresenius5008 type)

- Production controller is responsible to issue required sticker for Fresenius5008 sodium bi carbonate bag with holder and print Product reference, lot number, production date, expiry date, UDI QR Code and any other tender information on labels.

- Quality controller is responsible to make sure label sticker information should be as per approved labels and as per packaging order and all other information’s are correct.

- Clean bag by dry cloth and make sure bag should be free from dust particles before placing labels.

- Remove sticker from paper and fix on the middle of the bag and make sure that the sticker should be proper adhesive on bag body and sticker should be bubble free.

- Production controllers are responsible to arrange in specified space after labeling and transfer to tertiary packaging section for final packaging.

  - Tertiary Packaging (Sodium bi Carbonate Bag with Holder Fresenius 5008 Type)

- Production controller is responsible to issue production order for sodium bi carbonate bag with holder for tertiary packaging and make tertiary packaging sticker.

- Quality controller is responsible to make sure that the sticker information’s are correct as per approved labels and production order.

- Place cartridge one by one in tertiary packaging and make sure that cartridge should be in horizontal position in tertiary packaging.

- Close carton with tap and place approved sticker by glue.

- Production controller is responsible to arrange final finished products in as designated place.

## 10.9  Fresenius 5008 PE Bag In-Process Verification (PROPOSED)

• For bag forms F-823-i14-5 use original 24 process characteristics and record actual sample-specific net fill mass, bag weld width (historic 4-6 mm), weld parameter recipe, leak pressure/time and separate port/edge/holder leak outcomes. Keep existing historic criteria pending controlled approval, and suspend any conflicting settings. GAP-006/014/015.

• For the 5008 holder downside sealing temperature, source table lists only “- 240 °C”. This missing bound is OPEN TECHNICAL INPUT REQUIRED; the operator shall verify an authorised complete recipe rather than infer any temperature from this draft. GAP-014.

• Maintain traceability to qualified bag film/holder/filter/material compatibility and applicable machine study. Do not treat visually intact packaging as shelf-life or microbial verification. GAP-005/007.

- Production process Fresenius 4008 Bag with Neck Type Holder.

- Production controllers are responsible to issue required sheet roll for manufacturing of Fresenius bag from store.

- After receiving roll make sure the sheet should be free from holes and dust particles and issue production order for making sodium bi carbonate bag for Fresenius 4008 with neck holder.

- Cut the 2 sheets by manual for bag from roll in required length and width and make sure the cutting should be straight.

- The dimension of Fresenius bag with neck holder as in given table.

| Bag Dimension | Length (±5.0) mm | Width (± 0.1) mm | Thickness (± 0.01) mm |
| --- | --- | --- | --- |
|  | 330 - 400 | 130 | 0.31 |

- Quality controller is responsible to make sure that the bag dimension should be accurate as required and sheet should be straight cutting edges.

- Repeat process for cutting sheets until completed required quantity.

- Production controller is responsible to transfer sheets for next process sheet hole for fixing neck holder.

  - Process for making Hole in Sheet by Machine.

- Machine operator is responsible to make sure that the die is properly assembled and approved parameters

are set properly.

- The parameters for making hole in sheet in given table.

| Bag Neck Hole Parameters | Air Pressure ±0.2 Bar | Time ±0.5 Sec | Hole Diameter (±1.0mm) |
| --- | --- | --- | --- |
|  | 8.0 | 2.5 | 40 |

- Take one sheet and insert it into the machine die and make sure that the hole distance from top edge of bag should be 80±2mm and hole should be round circle shape the hole diameter should be 40±1.0mm.

- Quality controller is responsible to make sure that the distance from top to bag hole and hole diameter should be accurate and record the results in specified form.

- Repeat the process to complete required quantity and then transfer to next process for neck holder assembly with bag.

  - Assembly Filter with Holder with sheet.

- Production controller is responsible to arrange all components filters, neck holders and bag for assembly.

Insert 50-micron filter in holder inlet and make sure filter proper inserted in holder no loose insert holder into sheet cutting hole and make sure filter proper covered cutting area of holes and proper fixed.

- Quality controller is responsible to make sure that sheet hole proper covered by holder and no gap from edges and record the results in specified form.

- Repeat the process to assemble and until required quantity complete then transfer to next process for sealing holder with sheet.

11.3	Process for Welding Neck Holder with Sheet

11.3.1    High Power Sealing Machine (Fresenius 4008 Bag with Neck Type Holder)

- Machine operator is responsible to set the required die and make sure die properly fixed in machine and set the approved required parameters for neck sealing.

Parameters for neck holder sealing with sheet are.

| Neck Holder with Bag Sealing Parameters | Temperature (oC) | Sealing Down Time (Sec) | Sealing Welding Time (Sec) | Sealing Hold Time (Sec) | Up Timer Time (Sec) |
| --- | --- | --- | --- | --- | --- |
|  | 180 - 200 | 2 - 3 | - – 6.0 | 4.0 – 10.0 Sec | - – 3.0 |

- Place neck holder sheet in a machine die insert holder into machine die hole and make sure holder proper covered sheet hole close cover of machine and press start button then the machine weld neck holder with sheet.

- Quality controller is responsible to make sure holder should be proper welded with sheet and holes proper covered in sealing and no open from welded round edges of holder sealing side.

- Repeat process for required quantity and make sure neck holder should be proper welded with sheet.

- After completed welded process production controller is responsible to transfer sheets for next process to weld sheet 2nd layer.

11.3.2	High Power Sealing Machine (Welding 2nd layer)

- Machine operator is responsible to set the required die and sealing approved parameters.

| Bag Sealing Temperature 3-Sides | Temperature (oC) | Sealing Down Time (Sec) | Sealing Welding Time (Sec) | Sealing Hold Time (Sec) | Up Timer Time (Sec) |
| --- | --- | --- | --- | --- | --- |
|  | 145 - 155 | 2.0 – 3.0 | - – 6.0 | 2.0 – 5.0 Sec | - – 3.0 |

- Place holder assembled layer into die and cover by second layer and make sure that sheet edges should be fixed each other, then press start button to start welding and wait process to complete.

- Quality controller is responsible to make sure that the bag proper welded form three sides bottom, right and left and make sure layers proper welded each other, no gap between layers and no burn marks and holes record the results in specified form.

- Repeat the process to complete required quantity and then transfer to next process for cutting.

  - Bag Manual Cutting

- Production controller is responsible arrange cutting bag process by manual.

Cut unwelded sheet from bag from corner and bottom side by scissors and make sure the cutting should be straight and make sure bag should be accurate in dimension in given table.

| Bag Dimension After cutting | Length (± 1.0) cm | Width (± 1.0) cm |
| --- | --- | --- |
|  | 33.0 | 12.0 |

- Quality controller is responsible to make sure that the bag should be straight and smooth cutting edges, accurate in dimension and record the results in specified form.

- Repeat the process to complete required quantity and then transfer to next process for assembly tube and Porous PE filter with neck holder.

  - Assembly of Porous PE Filter with Tube and Neck Holder.

- Production controllers are responsible to issue required tube and filters for assembly.

- After receiving filters and tube make sure that the filters and tubes should be same as per issue slip and should be with proper identification labels.

- Hold tube and insert it into the filter hole and make sure tube should be proper inserted up to the required insertion hole of filter and other end of tube inset into the outlet channel of neck holder and make sure tube should not pullout from holder outlet port.

| Tube Dimension | Length (± 0.1) mm | O.D (± 0.1) mm | I.D (± 0.1) mm |
| --- | --- | --- | --- |
|  | 20cm | 3.7 | 2.4 |

- Quality controller is responsible to make sure that the tube proper fixed with filter and holder and tube should not pullout from filter and holder and record the results in specified form.

- After completed required quantity, production controllers are responsible to transfer bag for next process Sodium Bi carbonate filling by manual.

  - Bag Filling Process by Sodium Bi Carbonate Granules material (Manual Filling)

- Production controller is responsible to fill the container with sodium bi carbonate Granules material up to the required capacity and make sure the storage container should be with approved label.

- Fill sodium bi carbonate Granules material into cartridge, fill cartridge up to the required weight or follow graduation scale and then transfer to bag from bag top opening side and make sure filled bag should be accurate in weight.

- Quality controller is responsible to make sure that the filled sodium bi carbonate Granules material accurate in weight as required and record the results in specified form.

Repeat process for filling sodium bi carbonate Granules material ed in bag until the required quantity completed.

- After completed all bag filling process, production controller is responsible to transfer filled bag for next process for sealing bag top side by manual sealing machine.

  - Bag Top Side Sealing

- Production controller is responsible to arrange the filled bag and machine for sealing to side of bag.

- Machine operator is responsible to make sure the sealing machine approved parameters are set.

- Sealing machine parameters for sealing bag top side in a given table.

| Bag Top Side Sealing Parameters | Sealing Temperature ± 5.0 oC | Sealing Width ± 1.0mm | Time Sec |
| --- | --- | --- | --- |
|  | - - 300 | 5– 12 | 8– 15 |

    - Quality controller is responsible to make sure that bag proper sealed and no gap in sealing, no holes and burn marks the Granules material should not across or pass on filter side when by shaking the product and valve has no sink or streak or canal line on the contracting neck in order to control the leakage on dialysis machine and record the results in specified form.

    - Hold sealed bag for 45 minutes or keep the products cold before perform leak test.

    - After completed specified period, production controller is responsible to transfer all sealed bag for next process leakage test.

  - Leak Testing Process for Sodium Bi carbonate Bag with Neck holder.

  - Operators are responsible to set the required parameters (0.6 bar ± 0.2 bar pressure) for leak test as per approved parameters.

  - Connect required tubing or leak test adaptor, insert the steel adaptor into the neck holder and immerse the bag into water from holder side and apply required pressure (0.6 bar ± 0.2 for 8 – 15 second and observe any bubble appear from tubing connection, holder sealing side or edges of the sealed area.

  - Observe the air bubbles and if bubbles appeared from the sealing area considered Rejected and if no bubbles appear then it’s Acceptable.

  - Quality controller is responsible to get sample from final approved products as per AQL sampling plan and retest cartridges to make sure cartridges are leakage free and record the results in specified form.

Acceptance Criteria

- No leakage from sealed surface and edges, Acceptable

- Conductivity (historical company criterion): 13.5-14 mS/cm appears in source WI. HOLD for technical reconciliation. Apply only an approved SKU-specific tested solution specification, dilution, water quality, temperature/method and QC result field; this is NOT a universal ISO 23500-4 acceptance limit. OPEN TECHNICAL INPUT REQUIRED. [GAP-012]

- Homogeneous sealed no leakage, Acceptable

- No leakage from holder sealed area, Acceptable

- Production controller is responsible to transfer good products to next process for labelling.

- After leak test apply Silicon 1000CST inside the Neck Holder Part of the Bag and Place the Cap of the Neck Holder Tightly to secure the product from contamination.

  - Labelling Process (Sodium bi Carbonate Bag with Neck holder)

- Production controller is responsible to issue required sticker for sodium bi carbonate bag with neck holder and print Product reference, lot number, production date, expiry date, UDI QR Code and any other tender information on labels.

- Quality controller is responsible to make sure label sticker information should be as per approved labels and as per packaging order and all other information’s are correct.

- Clean bag by dry cloth and make sure bag should be free from dust particles before placing labels.

- Remove label sticker from paper and fix on the middle of the bag below the neck holder and make sure that the sticker should be proper adhesive on bag body and sticker should be bubble free.

- Production controllers are responsible to arrange in specified space after labeling and transfer to tertiary packaging section for final packaging.

  - Tertiary Packaging (Sodium bi Carbonate Bag with Neck Holder)

- Production controller is responsible to issue production order for sodium bi carbonate bag with neck holder for tertiary packaging and make tertiary packaging sticker.

- Quality controller is responsible to make sure that the sticker information’s are correct as per approved labels and production order.

- Fix supporting material in tertiary packaging and place bag one by one in tertiary packaging and make sure that bag should be in horizontal position in tertiary packaging.

- Close carton with tap and place approved sticker by glue.

- Production controller is responsible to arrange final finished products in as designated place.

## 12.  Cross-Configuration Acceptance, Labelling and Evidence Governance (PROPOSED)

• For Fresenius 4008 neck-holder bag retain original cutting, welding, fill, cooling, leakage and packaging steps; obtain the referenced F-823-i14-7 or approved successor before declaring routine record coverage for this configuration. GAP-018.

• Approved label/IFU must meet the applicable ISO 6.1/6.2/6.4 and conditional 6.6 information requirements. Keep controlled artwork/master IFU, validated shelf-life and after-dissolution use conditions, mixing-water requirements and microbial warnings. Add an applicable ISO clause matrix; do not invent unapproved values. GAP-009/010/011.

• Document evidence of qualified packaging and closure integrity, stability, product/material contact compatibility and machine functioning; specify actual verification cadence through QA-approved risk plans. ISO 4.1.9/4.1.11. GAP-005/007/008.

• Maintain one traceable product/SKU evidence register (ISO requirement, WI control, in-process field or separate record, latest approved scope, results and status). Blank templates and this draft never constitute compliance. GAP-001 through GAP-019.

• Final lot review must confirm approved incoming material disposition, in-process form completeness, manufacturing/qualification scope, test exceptions, deviations/NCRs and approved final release form; retain legible controlled record ID and released/held status. GAP-019.

- Final Inspection and Release:

- Quality Controller is responsible to check the products before transfer to Finished Good Store.

- The Quality Controller must inspect the product as per AQL general inspection method for visual

checking and special inspection method for physical checking and record the results in specified final inspection and release form.

- The quality assurance manager or quality control manager have to review and approve the final inspection and release form.

- The Production personnel must prepare a transfer slip and get approval by Q.A / Q.C manager.

- After Approval the Production controllers are responsible to transfer batch from production to store   along with transfer slip.

- Reference:

- Procedure for Inspection and Testing 				:	SOP-8-2-6

- Inspection & Test report for Sodium Bicarbonate			:    	F-823-T1

- Inspection & Test Report for General Product 			:	F-823-T90

- Inprocess control for Injection Molding & Extrusion 		:	F-823-i15

- Inprocess control for Sodium Bi carbonate cartridge assembly 	:	F-823-i14

- Inprocess control for Sodim Bi Carbonate Granules material 	:	F-823-i14-1

- In-process control for Sodium Bicarbonate Gambro Type		:    	F-832-i14-2 [DOCUMENT CONTROL HOLD: WI F-832 identifier conflicts with supplied F-823 form header. Verify approval before renumbering. GAP-018]

- In-process control for Sodium Bicarbonate Fresenius 5008		:    	F-832-i14-3 [DOCUMENT CONTROL HOLD: WI F-832 identifier conflicts with supplied F-823 form header. Verify approval before renumbering. GAP-018]

- In-process control for Sodium Bicarbonate Fresenius 5008 bag	:    	F-832-i14-5 [DOCUMENT CONTROL HOLD: WI F-832 identifier conflicts with supplied F-823 form header. Verify approval before renumbering. GAP-018]

- In-process control for Sodium Bicarbonate Fresenius 5008 bag	:    	F-832-i14-5 [DOCUMENT CONTROL HOLD: WI F-832 identifier conflicts with supplied F-823 form header. Verify approval before renumbering. GAP-018]

- In-process Control for Sodium Bi-Carbonate Bag			:	F-832-i14-6 [DOCUMENT CONTROL HOLD: WI F-832 identifier conflicts with supplied F-823 form header. Verify approval before renumbering. GAP-018]

- In-process control for Sodium Bicarbonate Fresenius 4008 bag neck	:	F-823-i14-7

- Final inspection and Release for Sodium Bi Carbonate 		:	F-823-F10

- Transfer Slip (Finished Goods / Bulk /Raw Material / product parts 	:	F-7-5-7-3

- References

ISO23500-4 Concentrates for haemodialysis and related therapies

- Modification / Documentation Change History

| Date | Changes |
| --- | --- |
| 10-04-2011 | Establish |
| 20-05-2012 | Modify Procedure 5.0 |
| 10-11-2014 | Modify item 5.3 |
| 03-04-2015 | Modify assembly procedure 6.1,6.2 |
| 22-06-2016 | Item 6.1 modify |
| 26.11.2018 | Item 4.3, 5.1.1, 5.4, 6.2.1, 6.3, 6.3.1, &9 new added |
| 15.01.2020 | 5.2, 6.2, 6.2.1, 6.3, updated |
| 29.06.2021 | 6.2.10 Parameters updated |
| 08.12.2022 | Updated Section 6. (Assembly Procedures & all the Machine Parameters updated as per Machine Validations) |
| 10.01.2023 | Section 6.2 (High Power Sealing Machine) parameters updated |
| 01.05.2023 | Receiving inspection for bulk components and components manufacturing process updated |
| 01.09.2023 | Complete procedure modified and process complete procedure new added step by step procedure of each machine, separate types of sodium bi carbonate Mfg. process, added new bag with holder type, bag with neck type holder added complete Mfg. procedure. |
| 9.09.2025 | Update Section 3.5,10,10.2,10.3,10.4.10.5,10.6,11.7. |

# 18. Proposed Revision Control and Open Technical Inputs

GAP-001 [P1]: Chemical-grade release criteria are not demonstrated. Comparing sodium bicarbonate results with a supplier COA alone does not establish compliance with the applicable pharmacopoeial monograph and its applicable gene

GAP-002 [P1]: No controlled reconstitution and solute-concentration demonstration is connected to the dry powder product. Existing conductivity acceptance is not a substitute for a documented composition evaluation.

GAP-003 [P1]: The WI does not define microbiological assurance for packaged bicarbonate powder and the traceable validation/verification route for dialysis-fluid microbiological performance.

GAP-004 [P1]: No demonstrated endotoxin-control and validated process evidence is linked to the bicarbonate powder configurations.

GAP-005 [P1]: Machine-specific function of the powder container and prevention of undissolved powder entering the dialysis-fluid stream are not established by the current visual fitting checks alone.

GAP-006 [P1]: Actual fill-weight decisions and applicability of listed ±10 g company tolerances are not anchored to a verified usage category or approved validated setpoints.

GAP-007 [P1]: Visual seal/leak appearance and manual immersion test do not alone substantiate chemical/physical compatibility of the powder packaging and closure through claimed shelf life.

GAP-008 [P2]: Product-contact manufacturing equipment material compatibility is not traceable; visual housekeeping and setpoint checks alone are insufficient to demonstrate non-interaction.

GAP-009 [P1]: In-process labels are checked largely for print/UDI but no complete controlled checklist verifies applicable general concentrate labelling information or package-insert linkage.

GAP-010 [P1]: Dry powder-specific labelling and reconstitution/handling information cannot be verified from the WI or the attached in-process forms.

GAP-011 [P1]: Concentrate-generator-specific labelling performance and machine identification are not traceably reviewed for Gambro and applicable Fresenius variants.

GAP-012 [P1]: A generic conductivity range of 13.5-14 mS/cm is used in several leak acceptance sections without identifying the solution, dilution, temperature, equipment or applicability; ISO does not prescribe this single globa

GAP-013 [P1]: Gambro machine sensor pressure entry uses inconsistent unit: 0.760 mbar (±0.2) versus 0.8 bar (±0.2), a potentially material three-order unit mismatch. Do not select the correct setting by inference.

GAP-014 [P1]: Gambro welding setpoints differ between controlled sources and Fresenius bag welding table contains an incomplete lower temperature, preventing a reliable, approved parameter comparison.

GAP-015 [P2]: Manual leak tests have qualitative remarks but incomplete raw readings, exposure time and sample trace for Fresenius 5008; pressure-duration fields are necessary to verify the WI procedure.

GAP-016 [P2]: Dimensional/product-part requirements are expressed as generic “per specification” in F-823-i15 and as unclear arithmetic in Gambro, preventing transparent measurement-based disposition.

GAP-017 [P2]: In-process sample frequency, inspection level and actual result notation are not consistently traceable to a controlled sampling plan and shift event.

GAP-018 [P2]: Document identifiers and controlled revisions conflict between WI contents, filenames and attached in-process form headers, and referenced additional records were not provided.

GAP-019 [P2]: Sealing/chemical conformity evidence is not systematically linked to batch disposition, change status and unresolved qualifications; blank in-process forms cannot close ISO findings.

Original WI production steps and tables have been retained as documentary baseline unless explicitly annotated or superseded above. In every conflict, only the current approved company master specification/validated machine recipe may be executed after QC/QA decision. Approval, implementation, effectiveness and actual test evidence remain pending. No existing approval signatures are transferred to this draft.