SPINAL NEEDLE

ISO 9626:2016

WI & Inspection Report Gap Analysis

Evidence-based QC / Medical Device Compliance Assessment

| Product | Spinal Needle |
| --- | --- |
| Primary Standard | ISO 9626:2016 - Stainless steel needle tubing for the manufacture of medical devices |
| Inspection Record | F-823-T58 Rev. 5, document date 20.07.2024, Page 1/1 |
| Product WI | Not identified in supplied Wi.pdf or Library search |
| Linked Specification | DD-01-PS referenced by report; not supplied |
| Assessment Date | 07 October 2026 |
| Assessment Status | Document / evidence gap analysis - not a product conformity certificate |

Confidential QC working analysis

1. Executive Summary

The current evidence set does not support a defensible claim of conformity to ISO 9626:2016 for the Spinal Needle tubing component. The main issue is not confirmed product nonconformity; it is insufficient controlled evidence and document linkage. No product-specific Spinal Needle WI was identified, the current inspection report leaves the “According to WI” field blank, and the repeatedly referenced specification DD-01-PS was not supplied for verification.

The current inspection report F-823-T58 Rev. 5 is materially more complete as a record template than the older historical report found in prior project evidence: it is Page 1/1 and dated after ISO 9626:2016. Therefore, the prior “missing page 2” and “report predates the standard” findings do not apply to the current report. However, the report still only partially addresses several tubing-specific ISO controls and does not provide traceable evidence for material grade, acidity/alkalinity, stiffness, or resistance to breakage.

Overall classification: Significant Evidence and Document Control Gaps. No compliance percentage or PASS status is assigned because the governing Spinal Needle WI, DD-01-PS, and supporting qualification / periodic / supplier records were not available for verification.

| Evidence item | Finding | Impact |
| --- | --- | --- |
| Spinal Needle WI | No product-specific Spinal Needle WI was identified in the supplied 18-page Wi.pdf or Library search. | ISO-to-WI implementation and revision control cannot be verified. |
| Current inspection report | F-823-T58 Rev.5, document date 20.07.2024, Page 1/1. “According to WI” is blank. | Current record is identifiable, but its governing instruction is not traceable. |
| DD-01-PS | Referenced for colour code, needle length, needle OD, stylet OD, introducer gauge/OD and flow rate; not supplied. | Numeric acceptance criteria and revision status cannot be verified. |
| Supporting evidence | No alloy certificate/CoA verification, tubing specification, pH, stiffness, breakage or Annex D method record supplied. | Several ISO requirements remain evidence gaps rather than confirmed implementation failures. |

2. Documents Reviewed

| Document | Identification / Revision | Pages reviewed | Assessment note |
| --- | --- | --- | --- |
| ISO 9626:2016 | Second edition, 2016-08-01 | PDF pp.7-19 (ISO pp.1-13) | Primary technical standard for rigid stainless-steel needle tubing. |
| Wi.pdf | 18-page scanned WI compilation | All 18 pages | Pages 1-4: Semi Auto Biopsy Needle; 5-7: Ureteric Catheter; 8-18: Endobronchial Tube. No Spinal Needle WI identified. |
| Inspection & Test Report for Spinal Needle | F-823-T58 Rev.5; document date 20.07.2024 | Page 1/1 | Rows 1-26 reviewed. “According to WI” field is blank. |
| Referenced specification | DD-01-PS | Not supplied | Referenced by rows 2, 9, 10, 11, 20 and 23; content/revision not verifiable. |
| Historical project evidence | Prior analysis identified older report F-PS1-118 Rev.3 dated 10.07.2013 | Historical only | Used only to confirm that a previous report identifier existed. It is not treated as the current controlled report. |

3. Scope

Assess the ISO 9626:2016 requirements applicable to the rigid stainless-steel needle tubing component used in the Spinal Needle product.

Compare those requirements with the available product WI evidence and the current inspection record F-823-T58 Rev.5.

Separate implementation gaps from evidence gaps and document-control gaps.

Do not assume every ISO test must be performed for each production lot. Controls may be located in raw-material specifications, supplier CoA/CoC verification, qualification, validation, periodic verification or type testing, provided the route and evidence are controlled and traceable.

Applicability boundary: ISO 9626:2016 addresses rigid stainless-steel needle tubing, not the complete finished Spinal Needle design. Product-level checks such as bevel type, point sharpness, gluing/pullout, logo, sterility status and flow rate are therefore not treated as substitutes for tubing-specific ISO 9626 requirements.

4. Methodology

Identify the correct product WI and current inspection record.

Extract applicable ISO 9626 requirements from Clauses 4 and 5 and normative Annexes A-D.

Map each requirement to WI evidence and inspection-report rows without treating generic phrases as equivalent to numeric or method-specific criteria.

Classify findings as Covered, Partial Gap, Gap, Evidence Gap or Document Control Gap.

Assign P1/P2/P3 based on material, dimensional, mechanical, performance, evidence and traceability impact.

| Covered | Partial Gap | Gap | Evidence Gap | Document Control Gap |
| --- | --- | --- | --- | --- |

5. Document Identification

| Product | Spinal Needle |
| --- | --- |
| WI | Not identified / not available for verification |
| WI Revision / Date | N/A - no Spinal Needle WI located |
| ISO | ISO 9626:2016 |
| Inspection Report | F-823-T58 - Inspection & Test Report for Spinal Needle |
| Report Revision / Date | Rev. 5 / document date 20.07.2024 |
| Report Pages | Page 1 of 1 |
| Other linked forms/specifications | DD-01-PS (referenced, not supplied); “According to WI” field is blank. |

Document Control Observation - P1

The current report is a controlled copy, but the “According to WI” field is blank and no Spinal Needle WI was identified in the supplied WI compilation or Library search. This prevents a controlled link between the inspection record and the governing product instruction. The current report number is F-823-T58; the older F-PS1-118 found in historical evidence is not treated as a mismatch because no current WI or supersession chain was supplied.

6. ISO 9626:2016 Requirements Review

| ID | ISO clause | Requirement | ISO page | Method / reference | Acceptance criterion | Nature |
| --- | --- | --- | --- | --- | --- | --- |
| R1 | 4 - Materials | Tubing material shall be a stainless-steel grade listed in ISO 15510 and selected for the intended use. | PDF p.8 (ISO p.2) | Controlled material specification / material certificate or equivalent traceable control. | Approved grade is listed in ISO 15510 and traceably identified. | Mandatory |
| R2 | 5.1 - General | Tubing selection for the specific application and intended use shall use a risk-based approach. | PDF p.8 (ISO p.2) | Risk assessment / design or material-selection record. | Documented risk-based selection appropriate to intended use. | Mandatory, system-level |
| R3 | 5.2 - Surface finish and visual appearance | Outside surface shall be smooth and free from defects; tube shall appear straight and of regular roundness. Product-specific finish criteria may be needed. | PDF p.8 (ISO p.2) | Normal/corrected vision plus controlled surface-finish specification where applicable. | Smooth, defect-free outside surface; straight and regular roundness. | Mandatory |
| R4 | 5.3 - Cleanliness | Tubing surfaces shall be free from metal soil and processing agents; product-specific cleanliness criteria may be required. | PDF p.8 (ISO p.2) | Visual examination plus controlled cleanliness specification where needed. | No metal soil or processing agents. | Mandatory |
| R5 | 5.4 - Limits for acidity and alkalinity | The pH of an extract prepared per Annex A shall be within one pH unit of the control fluid. | PDF p.8 (ISO p.2); Annex A PDF p.14 (ISO p.8) | Annex A: 3 g tubing in 250 ml grade-3 water; 37-40 °C for 60 ± 2 min; compare extract and control with laboratory pH meter. | Absolute pH difference between extract and control ≤ 1 pH unit. | Mandatory |
| R6 | 5.5 - Size designation | Tubing designation includes nominal OD in mm, corresponding gauge size and wall-thickness designation. | PDF p.8 (ISO p.2) | Controlled tubing specification / identification. | Complete designation: metric OD + gauge + wall type. | Mandatory |
| R7 | 5.6 - Dimensions | Tubing dimensions shall meet Table 1 for the applicable designated size and wall type; Table 1 provides ODmin, ODmax and IDmin. | PDF pp.8-11 (ISO pp.2-5) | Dimensional measurement against Table 1 or a controlled specification that incorporates the applicable limits. | Applicable ODmin/ODmax and IDmin limits met for the selected wall type. | Mandatory |
| R8 | 5.7 - Sample size | Where sampling is used, sample size shall be risk-based, included in the tube specification and based on intended use. | PDF p.11 (ISO p.5) | Documented risk-based sampling rationale in tubing specification / inspection plan. | Sampling basis and sample size documented and risk justified. | Conditional - where sampling applies |
| R9 | 5.8 - Stiffness | When Annex B parameters exist, tubing deflection shall not exceed the applicable Table 2 limit. Where parameters are not defined, the manufacturer shall set risk-based stiffness requirements. | PDF pp.11-12 (ISO pp.5-6); Annex B PDF pp.15-16 (ISO pp.9-10) | Annex B: span/force per Table 2, loading speed 1-10 mm/min, measure deflection to 0.01 mm; report size, wall type, date and alloy. | Deflection ≤ applicable Table 2 maximum, or approved risk-based limit where ISO parameters are undefined. | Mandatory / conditional by size |
| R10 | 5.9 - Resistance to breakage | Tubing tested per Annex C shall show no visible breakage. | PDF p.13 (ISO p.7); Annex C PDF pp.17-18 (ISO pp.11-12) | Annex C: bend at Table C.1 distance; 25 ± 1° RW, 20 ± 1° TW, 15 ± 1° ETW/UTW; 20 reversal cycles at 0.5 Hz; visual examination. | No visible breakage. | Mandatory |
| R11 | 5.10 - Resistance to corrosion | When tested per Annex D, the immersed half of the tubing shall show no evidence of corrosion resulting from the test. | PDF p.13 (ISO p.7); Annex D PDF p.19 (ISO p.13) | Annex D: approximately half immersed in 0.5 mol/L NaCl at 23 ± 2 °C for 7 h ± 5 min; compare immersed and unimmersed portions. | No evidence of corrosion caused by the immersion test. | Mandatory |

Note: ISO 9626 does not state that each control must appear on every routine final-inspection record. A requirement may be controlled through approved raw-material specifications, supplier evidence, qualification, periodic verification or other documented routes, provided the control is traceable and suitable for the intended use.

7. ISO ↔ WI ↔ Inspection Report Unified Gap Matrix

| ID | ISO Requirement | ISO Clause / Page | WI Coverage | WI Page | Inspection Report Coverage | Report Row / Page | Status | Priority |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| R1 | Stainless-steel material grade / ISO 15510 | 4; PDF p.8 (ISO p.2) | Not verifiable - no Spinal Needle WI identified. | N/A | NO. No alloy/grade, material certificate or CoA verification field is present. | p.1/1 - no row | Evidence Gap | P1 |
| R2 | Risk-based tubing selection for intended use | 5.1; PDF p.8 (ISO p.2) | Not verifiable - product WI/risk link unavailable. | N/A | NOT EXPECTED IN THIS REPORT. No linked risk/material-selection record is identified. | p.1/1 - none | Evidence Gap | P2 |
| R3 | Surface finish, defects, straightness, regular roundness | 5.2; PDF p.8 (ISO p.2) | Not verifiable. | N/A | PARTIAL. Row 12 requires “Good, smooth surface, no roughness”; it does not explicitly cover straightness, regular roundness or all serviceability-impairing tube defects. | Row 12, p.1/1 | Partial Gap | P2 |
| R4 | Cleanliness - free from metal soil / processing agents | 5.3; PDF p.8 (ISO p.2) | Not verifiable. | N/A | NO. “No flashing / no short shot” in rows 3-5 relates to molded components and is not equivalent to tubing cleanliness. | Rows 3-5 reviewed; p.1/1 | Evidence Gap | P2 |
| R5 | Acidity / alkalinity of tubing extract | 5.4 + Annex A; PDF pp.8,14 (ISO pp.2,8) | Not verifiable. | N/A | NO. No pH/extract result or linked test record is identified. | p.1/1 - none | Evidence Gap | P1 |
| R6 | Size designation: nominal OD + gauge + wall thickness | 5.5; PDF p.8 (ISO p.2) | Not verifiable. | N/A | PARTIAL. Row 2 references gauge/specification and row 10 references Needle OD through DD-01-PS; wall-thickness designation is not shown. | Rows 2,10; p.1/1 | Partial Gap | P1 |
| R7 | Dimensional limits: ODmin / ODmax / IDmin by wall type | 5.6 + Table 1; PDF pp.8-11 (ISO pp.2-5) | Not verifiable. | N/A | PARTIAL. Needle OD, stylet OD and introducer gauge/OD reference DD-01-PS, but the specification is unavailable and the form does not show IDmin or wall type. | Rows 9-11,20; p.1/1 | Partial Gap | P1 |
| R8 | Risk-based sample size included in tube specification | 5.7; PDF p.11 (ISO p.5) | Not verifiable. | N/A | PARTIAL / EVIDENCE. The form has Sample Level, Sample Size and AQL fields, but no evidence links those values to the ISO 5.7 risk-based tubing specification requirement. | Header/AQL section, p.1/1 | Evidence Gap | P2 |
| R9 | Needle-tubing stiffness / deflection | 5.8 + Table 2 + Annex B; PDF pp.11-12,15-16 | Not verifiable. | N/A | NO. No stiffness/deflection test, Table 2 limit or separate controlled record reference is present. | p.1/1 - none | Evidence Gap | P1 |
| R10 | Resistance of tubing to breakage | 5.9 + Annex C; PDF pp.13,17-18 | Not verifiable. | N/A | NO. Rows 17,18,21 and 24 are gluing/pullout/tensile checks; they do not demonstrate Annex C repeated-bending resistance of the tubing. | Rows 17,18,21,24; p.1/1 | Evidence Gap | P1 |
| R11 | Resistance to corrosion | 5.10 + Annex D; PDF pp.13,19 | Not verifiable. | N/A | PARTIAL. Row 22 states “Needles should be corrosion free”, but it does not identify Annex D conditions, sample identity/wall type/alloy, or a linked controlled test record. | Row 22; p.1/1 | Partial Gap | P1 |
| E1 | Controlled acceptance criteria in DD-01-PS | Linked evidence for rows 2,9,10,11,20,23 | Not available. | N/A | Report repeatedly relies on DD-01-PS, but the specification/revision was not supplied. | Rows 2,9-11,20,23; p.1/1 | Evidence Gap | P1 |
| D1 | WI / inspection-record linkage | Document control | No Spinal Needle WI identified in supplied Wi.pdf or Library search. | N/A | “According to WI” is blank on F-823-T58. | Report header, p.1/1 | Document Control Gap | P1 |

8. Key Non-Equivalence Observations

Tensile / pullout tests are not Annex C breakage testing. Rows 17, 18, 21 and 24 address joint strength or pullout/tensile force. ISO 9626 Clause 5.9 / Annex C tests the tubing itself under repeated bending and requires no visible breakage.

Flow rate is not a substitute for Table 1 dimensions. Row 23 controls flow rate through DD-01-PS, while Clause 5.6 separately requires defined tubing dimensions including OD limits and minimum ID for the applicable wall type.

General smoothness is not full Clause 5.2 / 5.3 coverage. Row 12 partially addresses surface smoothness but does not explicitly establish tube straightness, regular roundness, or freedom from metal soil / processing agents.

“Corrosion free” is not enough without method traceability. Row 22 contains the acceptance intent, but Annex D defines solution concentration, temperature, duration, partial immersion and comparison method. A separate controlled record is acceptable if it is linked and retrievable.

9. Critical Findings

| No. | Priority | Finding | Technical impact |
| --- | --- | --- | --- |
| 01 | P1 | Product-specific WI linkage is not verifiable | No Spinal Needle WI was identified and the current report leaves “According to WI” blank. Governing instruction/revision cannot be traced from the inspection record. |
| 02 | P1 | Material grade evidence is missing | Clause 4 requires stainless steel listed in ISO 15510. The current evidence set does not identify alloy/grade or a linked material certificate / CoA verification route. |
| 03 | P1 | Dimensional acceptance cannot be verified | Needle OD and related dimensions are controlled by DD-01-PS, but that specification is unavailable. IDmin and wall designation are not visible on the form. |
| 04 | P1 | No evidence for acidity / alkalinity control | No Clause 5.4 / Annex A pH verification or linked evidence was supplied. |
| 05 | P1 | No evidence for stiffness control | No Annex B / Table 2 stiffness record, qualification result or controlled alternative requirement was supplied. |
| 06 | P1 | No evidence for tubing breakage resistance | Existing tensile and pullout checks are technically different from Annex C repeated-bending resistance. |
| 07 | P1 | Corrosion requirement is only partially defined | Row 22 states a corrosion-free outcome but does not identify Annex D test conditions or a linked controlled method/record. |
| 08 | P1 | DD-01-PS acceptance criteria are not available | The current report relies on DD-01-PS for several critical dimensional/performance characteristics, so exact limits and revision status cannot be verified. |
| 09 | P2 | Surface finish and cleanliness criteria are incomplete | Smooth surface/no roughness is stated, but straightness, regular roundness, metal soil and processing-agent cleanliness are not explicitly controlled on the visible report. |
| 10 | P2 | Sampling basis is not traceable to ISO 5.7 risk rationale | Sample Level, Sample Size and AQL fields exist, but the risk-based basis and tubing-specification linkage were not supplied. |

10. Document Control Findings

| ID | Observation | Classification | Required control |
| --- | --- | --- | --- |
| DC-01 | No Spinal Needle WI identified; F-823-T58 “According to WI” field is blank. | Document Control Gap - P1 | Confirm the current controlled WI number/revision and populate or link it in the inspection record. If no WI exists, create and approve the governing instruction. |
| DC-02 | DD-01-PS is referenced repeatedly but was not supplied. | Evidence Gap - P1 | Confirm document number, current revision, approval status and exact acceptance limits used by QC; ensure the inspection record identifies the applicable revision or controlled source. |
| DC-03 | Historical project evidence identifies older report F-PS1-118 Rev.3, while the current supplied form is F-823-T58 Rev.5. | Traceability observation - P3 | No mismatch is asserted. Confirm supersession/history only if the older identifier remains referenced in active documents or records. |

11. Recommended Actions

| ID | Requirement | Current State | Gap | Required Action |
| --- | --- | --- | --- | --- |
| A1 | Product WI / linkage | No Spinal Needle WI verified; report WI field blank. | Document-control and implementation route cannot be traced. | Confirm the current Spinal Needle WI. If none exists, create/approve one. Update F-823-T58 to reference the exact controlled WI number and revision or a controlled electronic cross-reference. |
| A2 | Material grade | No alloy/grade or CoA verification evidence. | Clause 4 conformity cannot be demonstrated. | Add or confirm a controlled needle-tubing material specification identifying the stainless-steel grade listed in ISO 15510. Define supplier CoA/material-certificate verification and lot traceability. Routine chemical re-testing is not required unless the approved control strategy requires it. |
| A3 | Dimensions and wall classification | Report references DD-01-PS for OD/length but DD-01-PS is unavailable; IDmin and wall type are not visible. | Table 1 limits cannot be checked. | Confirm the current DD-01-PS revision and ensure it contains the applicable nominal OD/gauge, wall designation (RW/TW/ETW/UTW), ODmin/ODmax and IDmin, or an explicit controlled equivalent. Record the specification revision used for inspection. |
| A4 | Acidity / alkalinity | No test or linked evidence supplied. | Clause 5.4 evidence missing. | Establish a controlled test or external/supplier/periodic verification route aligned to Clause 5.4 and Annex A. Record extract/control pH and confirmation that the difference is ≤ 1 pH unit. Define frequency in the QMS; do not automatically require every production lot. |
| A5 | Stiffness | No stiffness/deflection evidence supplied. | Mechanical-performance evidence missing. | Create or link a controlled method based on Annex B / Table 2 for applicable sizes. For sizes without ISO parameters, establish a risk-based device-specific stiffness requirement. Record size, wall type, applied parameters, measured deflection, date and alloy. |
| A6 | Resistance to breakage | Current tensile/pullout tests are not equivalent. | Clause 5.9 evidence missing. | Create or link Annex C testing as qualification/periodic/routine control according to the approved quality plan. Record Table C.1 distance, applicable wall-type angle, 20 cycles at 0.5 Hz and the no-visible-breakage result. |
| A7 | Corrosion | Row 22 only states “corrosion free”. | Acceptance intent present, but method/evidence incomplete. | Reference Annex D or an approved equivalent controlled method. Record sample identity, wall type, alloy, date and result after the defined NaCl immersion conditions. A separate qualification/periodic record is acceptable if linked. |
| A8 | Surface finish / cleanliness | Current report partially addresses surface condition. | Visual-control completeness is weak. | Add explicit criteria or a controlled specification reference for straightness, regular roundness, absence of serviceability-impairing defects, metal soil and processing agents. |
| A9 | Sampling | AQL/sample fields exist, but ISO 5.7 risk basis is not shown. | Sampling rationale not traceable. | Document the sampling basis in the tubing specification/inspection plan using risk-assessment principles and intended use. If AQL is used, link it to the approved plan rather than treating AQL alone as proof of ISO 5.7 compliance. |

12. Priority Action Plan

| Priority | Action | Deliverable / evidence required |
| --- | --- | --- |
| P1 | Confirm the controlled Spinal Needle WI and link it to F-823-T58. | Approved WI number, revision, effective date and report linkage. |
| P1 | Recover and verify DD-01-PS. | Current approved specification with dimensional limits, gauge/OD mapping, wall type and ID control. |
| P1 | Establish material-grade verification. | Approved raw-material specification + ISO 15510 grade + CoA/material-certificate verification record. |
| P1 | Establish / link Clause 5.4, 5.8, 5.9 and 5.10 verification routes. | Controlled methods and records for acidity/alkalinity, stiffness, breakage and corrosion; frequency justified by quality plan/risk. |
| P2 | Complete surface-finish / cleanliness wording and sampling traceability. | Controlled criteria and risk-based sampling rationale linked to the tubing specification. |

13. Final Conclusion

Based on the documents available for this assessment, ISO 9626:2016 conformity for the Spinal Needle tubing component cannot be fully verified. The current inspection report F-823-T58 Rev.5 provides useful product-level controls and partial coverage of surface condition, size/OD references and corrosion acceptance. However, the absence of a product-specific WI, the unavailable DD-01-PS specification, and the lack of traceable evidence for material grade, pH, stiffness and repeated-bending breakage prevent closure of the P1 requirements.

The recommended corrective approach is not to add every ISO test to every production-lot inspection sheet. Instead, define a controlled compliance route for each requirement and link the appropriate evidence: incoming material specification and CoA verification, approved dimensional specification, qualification/periodic methods, and a WI/report cross-reference. Once those records are available, the matrix should be re-assessed to distinguish true implementation gaps from previously unresolved evidence gaps.

14. Source Traceability

| Source | Exact traceability used in this assessment |
| --- | --- |
| ISO 9626:2016 | Scope: PDF p.7 (ISO p.1). Clauses 4 and 5.1-5.6: PDF p.8 (ISO p.2). Table 1: PDF pp.9-11 (ISO pp.3-5). Clause 5.7-5.8 / Table 2: PDF pp.11-12 (ISO pp.5-6). Clauses 5.9-5.10: PDF p.13 (ISO p.7). Annex A: PDF p.14 (ISO p.8). Annex B: PDF pp.15-16 (ISO pp.9-10). Annex C: PDF pp.17-18 (ISO pp.11-12). Annex D: PDF p.19 (ISO p.13). |
| Wi.pdf | All 18 pages reviewed. Pages 1-4 concern Semi Auto Biopsy Needle; 5-7 Ureteric Catheter; 8-18 Endobronchial Tube. No Spinal Needle WI identified. |
| Current inspection record | F-823-T58 Rev.5, document date 20.07.2024, Page 1/1. Rows 1-26 mapped directly in this report; “According to WI” field blank. |
| DD-01-PS | Referenced by current report rows 2, 9, 10, 11, 20 and 23. Not supplied; no numeric acceptance limits or revision were assumed. |

Appendix A - Current Inspection Report Row Map

Source: F-823-T58 Rev.5, Inspection & Test Report for Spinal Needle, document date 20.07.2024, Page 1/1. The table below transcribes the visible requirement wording used for row-level mapping. Blank inspection result fields are not interpreted as failures because this is an uncompleted form/template image.

| Row | Description | Requirement on current report | ISO 9626 mapping note |
| --- | --- | --- | --- |
| 1 | Packing Condition | Good without any damage with label with clear product informations | No direct ISO 9626 tubing requirement used for this assessment. |
| 2 | Colour code | As per required gauge and specification (DD-01-PS) | Partial support for gauge/specification only; does not show wall type. |
| 3 | Holder status | No damage, No flashing, No short shot | No direct ISO 9626 tubing requirement used for this assessment. |
| 4 | Needle holder status | No damage, No flashing, No short shot | No direct ISO 9626 tubing requirement used for this assessment. |
| 5 | Needle protector status | No damage, No flashing, No short shot | No direct ISO 9626 tubing requirement used for this assessment. |
| 6 | Fitting of Needle with holder | Good fitting, no loose, no pullout | No direct ISO 9626 tubing requirement used for this assessment. |
| 7 | Fitting of stylet with holder | Good fitting, no loose, no pullout | No direct ISO 9626 tubing requirement used for this assessment. |
| 8 | Fitting of protector with needle holder | Good fitting, no loose, no pullout | No direct ISO 9626 tubing requirement used for this assessment. |
| 9 | Needle Length | As per Specification (DD-01-PS) | Product dimensional control; ISO 9626 focuses tubing dimensions. |
| 10 | Needle O.D | As per Specification (DD-01-PS) | Partial support for OD; DD-01-PS unavailable; ID/wall type absent. |
| 11 | Stylet O.D | As per Specification (DD-01-PS) | Stylet OD is not automatically a tubing requirement under ISO 9626 without component applicability evidence. |
| 12 | Needle status | Good, smooth surface, no roughness | Partial Clause 5.2 surface-finish coverage only. |
| 13 | Needle bevel | Quincke bevel / Pencil point bevel | No direct ISO 9626 tubing requirement used for this assessment. |
| 14 | Needle tip | Sharp tip without damage and bend | No direct ISO 9626 tubing requirement used for this assessment. |
| 15 | Stylet status | Good smooth surface, no roughness | No direct ISO 9626 tubing requirement used for this assessment. |
| 16 | Stylet tip | Sharp tip without damage and bend | No direct ISO 9626 tubing requirement used for this assessment. |
| 17 | Needle gluing w/ N.holder | Strong gluing / no pullout | Joint/pullout control; not Annex C breakage. |
| 18 | Stylet gluing w/ holder | Strong gluing / no pullout | Joint/pullout control; not Annex C breakage. |
| 19 | Introducer Needle | With introd. Needle / Without Introd. Needle | No direct ISO 9626 tubing requirement used for this assessment. |
| 20 | Gauge & OD of Introd. Needle | As per specification (DD-01-PS) | Partial dimensional control for introducer needle; DD-01-PS unavailable. |
| 21 | Tensile test introducer Needle with holder | OD ≥ 0.550 and < 0.750 mm: 3 N; OD ≥ 0.750 and < 1.150 mm: 5 N; OD ≥ 1.150 and < 1.850 mm: 10 N; OD ≥ 1.850 mm: 15 N | Joint tensile control; not Annex C repeated-bending test. |
| 22 | Corrosion Resistance | Needles should be corrosion free | Partial Clause 5.10 acceptance intent; Annex D method traceability absent. |
| 23 | Flow Rate | Flow rate should be as per specification (DD-01-PS) | Performance control; not a substitute for ISO 9626 Table 1 dimensions. |
| 24 | Peak tensile force Needle with holder & Stylet with holder | Needle with holder and stylet with holder tensile force should be >15 N | Joint tensile control; not Annex C repeated-bending test. |
| 25 | Logo | with Logo / Without Logo | No direct ISO 9626 tubing requirement used for this assessment. |
| 26 | Sterility status | Sterile / Bulk | No direct ISO 9626 tubing requirement used for this assessment. |