SPINAL NEEDLE

ISO 9626:2016

WI & Inspection Report Gap Analysis

Evidence-based QC / Medical Device Compliance Assessment

| Product | Spinal Needle |
| --- | --- |
| Primary Standard | ISO 9626:2016 - Stainless steel needle tubing for the manufacture of medical devices |
| Inspection Record | F-PS1-118 Rev. 3, dated 10.07.2013 - supplied page 1 of 2 only |
| Product WI | Not identified in supplied Wi.pdf or retrieved from Library search |
| Assessment Date | 06 October 2026 |
| Assessment Status | Document / evidence gap analysis - not a product conformity certificate |

Confidential QC working analysis

# 1. Executive Summary

Conformity to ISO 9626:2016 cannot be demonstrated from the supplied evidence set. The main limitation is not proof of product nonconformity; it is the absence of a verifiable Spinal Needle work instruction, the absence of the referenced controlled specification DD-01-PS, and the fact that only page 1 of 2 of the inspection report was supplied. The available inspection report partially addresses surface condition, outer diameter by reference, and corrosion, but it does not provide traceable evidence for several tubing-specific requirements such as material grade, acidity/alkalinity, stiffness, resistance to breakage, and the Annex D corrosion method.

Assessment conclusion Overall classification: Significant Evidence and Document Control Gaps. No PASS / compliance percentage is assigned because the required product WI, DD-01-PS specification, page 2 of the inspection report, and supporting test records were not available for verification.

| Evidence item | Finding | Impact |
| --- | --- | --- |
| Spinal Needle WI | No product-specific Spinal Needle WI was identified in the supplied 18-page Wi.pdf. The file contains WIs for other products only. | ISO-to-WI implementation cannot be verified. |
| Inspection report | F-PS1-118 Rev.3 is identified, but only page 1 of 2 is available. | Tests possibly recorded on page 2 cannot be assumed absent or present. |
| DD-01-PS | Referenced for colour code, needle length, needle OD, stylet OD, introducer gauge/OD and flow rate, but not supplied. | Numeric acceptance criteria and revision control cannot be verified. |
| Supporting evidence | No material certificate/CoA, tubing specification, qualification record, stiffness/breakage/pH/corrosion test record supplied. | Several ISO requirements remain evidence gaps rather than confirmed implementation failures. |

# 2. Documents Reviewed

| Document | Identification / Revision | Pages reviewed | Assessment note |
| --- | --- | --- | --- |
| ISO 9626:2016 | Second edition, 2016-08-01 | Scope and normative requirements; PDF pp.7-19 (ISO pp.1-13) | Primary technical standard for rigid stainless-steel needle tubing. |
| Wi.pdf | 18-page scanned WI compilation | All 18 pages | Contains Semi Auto Biopsy Needle, Ureteric Catheter and Endobronchial Tube WIs; no Spinal Needle WI identified. |
| Inspection & Test Report for Spinal Needle | F-PS1-118 Rev.3; Date 10.07.2013 | Supplied page 1 of 2 only | Rows 1-26 reviewed. Page 2 is not available. |
| Referenced specification | DD-01-PS | Not supplied | Referenced repeatedly by F-PS1-118; content/revision cannot be verified. |

# 3. Scope

Assess the ISO 9626:2016 requirements that apply to the rigid stainless-steel needle tubing component used in the Spinal Needle product.

Compare those requirements against the product-specific WI evidence and the supplied inspection record.

Separate implementation gaps from evidence gaps and document-control gaps.

Do not assume that each ISO test must be performed for every production lot. Where appropriate, controls may be located in raw-material specifications, supplier CoA/CoC verification, qualification, validation, periodic verification or type testing - but such controls require traceable documented evidence.

Applicability boundary ISO 9626:2016 applies to rigid stainless-steel needle tubing. The report identifies a needle, stylet, OD and corrosion checks, but does not state the stainless-steel alloy/grade. This assessment therefore treats ISO 9626 as applicable to the needle-tubing component, while flagging material confirmation as required evidence.

# 4. Methodology

1. Identify the correct product WI and inspection record.

2. Extract applicable ISO 9626 requirements from Clauses 4 and 5 and normative Annexes A-D.

3. Map each requirement to WI evidence and inspection-report rows without treating generic phrases as equivalent to numeric or method-specific ISO criteria.

4. Classify as Covered, Partial Gap, Gap, Evidence Gap or Document Control Gap.

5. Assign P1/P2/P3 based on material/dimensional/mechanical/performance criticality, evidence completeness and document-control impact.

| Covered | Partial Gap | Gap | Evidence Gap | Document Control Gap |
| --- | --- | --- | --- | --- |

# 5. Document Identification

| Field | Identified value |
| --- | --- |
| Product | Spinal Needle |
| WI | Not identified / not available for verification |
| WI Revision | N/A - no Spinal Needle WI located |
| WI Date | N/A |
| ISO | ISO 9626:2016 |
| Inspection Report | F-PS1-118 - Inspection & Test Report for Spinal Needle |
| Report Revision | Rev. 3 |
| Report Date | 10.07.2013 |
| Report Pages Available | Page 1 of 2 only |
| Other linked forms/specifications | DD-01-PS (referenced, not supplied); "According to MI" field is blank on the report. |

Document Control Observation - P1 The inspection report contains an "According to MI" field but it is blank, and no Spinal Needle WI was identified in the supplied WI compilation. This prevents a controlled link between the inspection record and the governing product instruction.

# 6. ISO 9626:2016 Requirements Review

| ID | ISO clause | Requirement | ISO page | Method / reference | Acceptance criterion | Nature |
| --- | --- | --- | --- | --- | --- | --- |
| R1 | 4 - Materials | Tubing shall be made from stainless steels listed in ISO 15510; the selected material shall suit the intended use. | PDF p.8 (ISO p.2) | Approved material specification / material certificate or equivalent control. | Selected stainless-steel grade is within ISO 15510 and is traceably identified. | Mandatory |
| R2 | 5.1 - General | Tubing selection for the specific application and intended use shall use a risk-based approach. | PDF p.8 (ISO p.2) | Risk assessment / design or material-selection record. | Documented risk-based selection appropriate to intended use. | Mandatory, system-level |
| R3 | 5.2 - Surface finish and visual appearance | Outside surface smooth and free from defects; manufacturer-specific finish criteria where needed; tube appears straight and of regular roundness. | PDF p.8 (ISO p.2) | Normal or corrected vision; controlled surface-finish specification where applicable. | No serviceability-impairing defects; smooth outside surface; straight and regular roundness. | Mandatory |
| R4 | 5.3 - Cleanliness | Tubing surfaces shall be free from metal soil and processing agents; product-specific cleanliness specifications may be required. | PDF p.8 (ISO p.2) | Visual examination plus controlled cleanliness specification where needed. | No metal soil or processing agents. | Mandatory |
| R5 | 5.4 - Limits for acidity and alkalinity | pH of the Annex A extract shall be within one pH unit of the control fluid. | PDF p.8 (ISO p.2); Annex A PDF p.14 (ISO p.8) | Annex A: 3 g tubing / 250 ml grade-3 water; 37 to 40 C for 60 +/- 2 min; compare extract with control using pH meter. | Absolute pH difference between extract and control <= 1 pH unit. | Mandatory |
| R6 | 5.5 - Size designation | Tubing size is designated by nominal OD in mm, corresponding gauge and wall thickness designation. | PDF p.8 (ISO p.2) | Controlled tubing specification / identification. | Complete designation including nominal OD, gauge and wall type. | Mandatory |
| R7 | 5.6 - Dimensions | Tubing dimensions shall comply with Table 1. Table 1 defines ODmin, ODmax and IDmin by designated metric size and wall type. | PDF pp.8-11 (ISO pp.2-5) | Dimensional measurement against Table 1 or a controlled specification incorporating those limits. | Applicable ODmin/ODmax and IDmin limits met for the specified wall type. | Mandatory |
| R8 | 5.7 - Sample size | Where sampling is applicable, sample size shall be risk-based, included in the tube specification and based on intended use. | PDF p.11 (ISO p.5) | Documented risk-based sampling rationale in tubing specification / inspection plan. | Sampling basis and sample size are documented and risk justified. | Conditional - where sampling applies |
| R9 | 5.8 - Stiffness | When tested per Annex B, deflection shall not exceed Table 2. If standard parameters are not defined for a tube, manufacturer-specific stiffness requirements shall be risk-based. | PDF pp.11-12 (ISO pp.5-6); Annex B PDF pp.15-16 (ISO pp.9-10) | Annex B: span and force per Table 2, loading speed 1-10 mm/min, record deflection to 0.01 mm. | Deflection <= applicable Table 2 maximum, or approved risk-based limit where ISO parameters are not defined. | Mandatory / conditional by size |
| R10 | 5.9 - Resistance to breakage | Tubing tested per Annex C shall show no visible breakage. | PDF p.13 (ISO p.7); Annex C PDF pp.17-18 (ISO pp.11-12) | Annex C: bend at Table C.1 distance; angle depends on wall type; 20 reversal cycles at 0.5 Hz; visual examination. | No visible breakage. | Mandatory |
| R11 | 5.10 - Resistance to corrosion | When tested per Annex D, the immersed half shall show no evidence of corrosion resulting from the test. | PDF p.13 (ISO p.7); Annex D PDF p.19 (ISO p.13) | Annex D: approx. half immersed in 0.5 mol/L NaCl at 23 +/- 2 C for 7 h +/- 5 min; compare immersed vs unimmersed portion. | No evidence of corrosion caused by the immersion test. | Mandatory |

Note: ISO 9626 does not require every control to appear on the routine final inspection form. A requirement may be controlled through approved raw-material specifications, supplier evidence, qualification or periodic testing, provided the control route and evidence are documented and traceable.

# 7. ISO - WI - Inspection Report Unified Gap Matrix

| ID | ISO requirement | ISO clause / page | WI coverage | WI page | Inspection report coverage | Report row / page | Status | Priority |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| R1 | Stainless-steel material grade / ISO 15510 | 4; PDF p.8 (ISO p.2) | Not verifiable - no Spinal Needle WI identified. | N/A | NO. No alloy/grade, material certificate or CoA verification field is visible on supplied page 1. | p.1/2 - no row | Evidence Gap | P1 |
| R2 | Risk-based tubing selection for intended use | 5.1; PDF p.8 (ISO p.2) | Not verifiable - product WI/risk link unavailable. | N/A | NOT EXPECTED IN THIS REPORT. No linked risk/material-selection record is shown. | p.1/2 - none | Evidence Gap | P2 |
| R3 | Surface finish, defects, straightness, regular roundness | 5.2; PDF p.8 (ISO p.2) | Not verifiable. | N/A | PARTIAL. Row 12 requires "Good, smooth surface, no roughness"; it does not explicitly cover straightness, regular roundness or serviceability-impairing tube defects. | Row 12, p.1/2 | Partial Gap | P2 |
| R4 | Cleanliness - free from metal soil / processing agents | 5.3; PDF p.8 (ISO p.2) | Not verifiable. | N/A | NO on supplied page 1. "No flashing/no short shot" rows relate to molded component condition and are not equivalent to tubing cleanliness. | Rows 3-5 reviewed; p.1/2 | Evidence Gap | P2 |
| R5 | Acidity / alkalinity of tubing extract | 5.4 + Annex A; PDF pp.8,14 (ISO pp.2,8) | Not verifiable. | N/A | NO on supplied page 1. No pH/extract result or linked test record is identified. | p.1/2 - none | Evidence Gap | P1 |
| R6 | Size designation: nominal OD + gauge + wall thickness | 5.5; PDF p.8 (ISO p.2) | Not verifiable. | N/A | PARTIAL. Colour code is "as per required gauge and specification" and Needle OD is referenced to DD-01-PS, but wall thickness designation is not shown. | Rows 2, 10; p.1/2 | Partial Gap | P1 |
| R7 | Dimensional limits: ODmin / ODmax / IDmin by wall type | 5.6 + Table 1; PDF pp.8-11 (ISO pp.2-5) | Not verifiable. | N/A | PARTIAL. Needle OD is "As per Specification (DD-01-PS)". The linked specification is unavailable, and the form does not visibly record ID or wall type. | Rows 9-11, 20; p.1/2 | Partial Gap | P1 |
| R8 | Risk-based sample size included in tube specification | 5.7; PDF p.11 (ISO p.5) | Not verifiable. | N/A | PARTIAL. Form contains Sample Level, Sample Size and AQL fields, but no evidence that the sampling basis is risk-assessed and included in the tubing specification. | Header / AQL section, p.1/2 | Evidence Gap | P2 |
| R9 | Needle-tubing stiffness / deflection | 5.8 + Table 2 + Annex B; PDF pp.11-12,15-16 | Not verifiable. | N/A | NO on supplied page 1. No stiffness/deflection test, Table 2 limit or separate record reference is visible. | p.1/2 - none | Evidence Gap | P1 |
| R10 | Resistance of tubing to breakage | 5.9 + Annex C; PDF pp.13,17-18 | Not verifiable. | N/A | NO on supplied page 1. Rows 17, 18, 21 and 24 are gluing/pullout/tensile checks; they do not demonstrate Annex C repeated-bending resistance to breakage. | Rows 17,18,21,24; p.1/2 | Evidence Gap | P1 |
| R11 | Resistance to corrosion | 5.10 + Annex D; PDF pp.13,19 | Not verifiable. | N/A | PARTIAL. Row 22 states "Needles should be corrosion free", but the form does not specify Annex D conditions, test method, sample identity/wall type, alloy or a linked test record. | Row 22; p.1/2 | Partial Gap | P1 |
| E1 | Controlled acceptance criteria in DD-01-PS | Linked evidence for rows 2,9,10,11,20,23 | Not available. | N/A | Report repeatedly relies on DD-01-PS, but the specification/revision was not supplied. | Rows 2,9-11,20,23; p.1/2 | Evidence Gap | P1 |
| D1 | WI / inspection record linkage | Document control | No Spinal Needle WI identified in supplied Wi.pdf. | N/A | "According to MI" is blank on F-PS1-118. | Report header, p.1/2 | Document Control Gap | P1 |
| D2 | Current standard vs report revision history | Document control / change impact | No product WI available to show current standard adoption. | N/A | F-PS1-118 Rev.3 is dated 10.07.2013, while the assessed ISO edition is 2016. No change-impact or post-2016 form revision evidence was supplied. | Header, p.1/2 | Document Control Gap | P1 |
| E2 | Inspection record completeness | Evidence completeness | N/A | N/A | Only page 1 of 2 was supplied; page 2 content cannot be verified. | Header: "Page 1 of 2" | Evidence Gap | P2 |

# 8. Key Non-Equivalence Observations

Tensile / pullout tests are not Annex C breakage testing. Rows 17, 18, 21 and 24 address joint strength or pullout/tensile force. ISO 9626 Clause 5.9 / Annex C tests the tubing itself under repeated bending and requires no visible breakage.

Flow rate is not a substitute for Table 1 dimensions. Row 23 controls flow rate through DD-01-PS, but ISO 9626 Clause 5.6 separately requires defined tubing dimensions, including OD limits and minimum ID for the applicable wall type.

General smoothness is not full Clause 5.2 / 5.3 coverage. Row 12 partially addresses smooth surface condition, but does not explicitly establish tube straightness, regular roundness, or absence of metal soil / processing agents.

"Corrosion free" is not enough without method traceability. Row 22 has a high-level acceptance statement, but Annex D defines solution concentration, temperature, time, partial immersion and comparison method. A separate controlled method/record may be used, but it must be linked and retrievable.

# 9. Critical Findings

| No. | Priority | Finding | Technical impact |
| --- | --- | --- | --- |
| 01 | P1 | Product-specific WI / MI linkage is not verifiable | No Spinal Needle WI was identified in the supplied WI file, and the inspection report has a blank "According to MI" field. The governing instruction and its revision cannot be traced from the record. |
| 02 | P1 | Material grade evidence is missing | ISO 9626 Clause 4 requires stainless steel from ISO 15510. The supplied record does not identify the alloy/grade or a linked material certificate / CoA verification route. |
| 03 | P1 | Dimensional acceptance cannot be verified | Needle OD and other characteristics are controlled by reference to DD-01-PS, but that controlled specification was not supplied. The visible form does not show ID or wall-type designation required to select Table 1 limits. |
| 04 | P1 | No evidence for acidity / alkalinity control | No Annex A / Clause 5.4 pH test or linked evidence was supplied. |
| 05 | P1 | No evidence for stiffness control | No Annex B / Table 2 stiffness result, qualification record or controlled alternative requirement was supplied. |
| 06 | P1 | No evidence for tubing breakage resistance | Existing tensile and pullout checks are different tests and cannot demonstrate Annex C repeated-bending resistance. |
| 07 | P1 | Corrosion requirement is only partially defined | Row 22 states that needles should be corrosion free, but no Annex D method conditions or linked controlled test record are identified. |
| 08 | P1 | Form revision predates the assessed ISO edition | F-PS1-118 Rev.3 is dated 2013. If ISO 9626:2016 is the current governing standard, a documented change-impact assessment or updated controlled record is needed to show the 2016 requirements were incorporated. |
| 09 | P2 | Surface finish and cleanliness criteria are incomplete on the visible report | Smooth surface / no roughness is recorded, but straightness, regular roundness, metal soil and processing-agent cleanliness are not explicitly addressed on supplied page 1. |
| 10 | P2 | Sampling basis is not traceable to a risk-based tubing specification | AQL / sample fields exist, but the ISO 5.7 risk-based basis and tubing-specification linkage were not supplied. |

# 10. Document Control Findings

| ID | Observation | Classification | Required control |
| --- | --- | --- | --- |
| DC-01 | No Spinal Needle WI identified in supplied Wi.pdf; F-PS1-118 "According to MI" field is blank. | Document Control Gap - P1 | Confirm the current controlled WI/MI number and revision. Populate/link it in the inspection record or update the record design. |
| DC-02 | F-PS1-118 Rev.3 is dated 10.07.2013, before ISO 9626:2016. | Document Control Gap - P1 | Perform and retain a documented standards-change impact review. Revise the WI/report/specification if the 2016 requirements were not incorporated. |
| DC-03 | DD-01-PS is referenced repeatedly but was not supplied. | Evidence Gap - P1 | Confirm document number, current revision, approval status and the exact acceptance limits used by QC. |
| DC-04 | Only page 1 of 2 of F-PS1-118 was supplied. | Evidence Gap - P2 | Retrieve page 2 and confirm whether any ISO-related tests or linked records are documented there. |

# 11. Recommended Actions

| ID | Requirement | Current state | Gap | Required action |
| --- | --- | --- | --- | --- |
| A1 | Product WI / MI linkage | No Spinal Needle WI was verified; report MI field is blank. | Document control and ISO implementation cannot be traced. | Confirm the current Spinal Needle WI/MI. If none exists, create/approve one. Update F-PS1-118 to reference the exact controlled document number and revision or use a controlled cross-reference field. |
| A2 | Material grade | No alloy/grade or CoA verification evidence. | Clause 4 conformity cannot be demonstrated. | Add a controlled needle-tubing material specification identifying the stainless-steel grade listed in ISO 15510. Define supplier CoA/material certificate verification, lot traceability and acceptance record. Routine chemical re-testing is not required unless your approved control strategy requires it. |
| A3 | Dimensions and wall classification | Report uses "As per DD-01-PS" for OD/length but DD-01-PS is unavailable; ID and wall type are not visible. | Table 1 limits cannot be checked. | Confirm DD-01-PS current revision and ensure it contains the applicable nominal OD/gauge, wall designation (RW/TW/ETW/UTW), ODmin/ODmax and IDmin or an explicit controlled equivalent. Record the specification revision used for inspection. |
| A4 | Acidity / alkalinity | No test or linked evidence supplied. | Clause 5.4 evidence missing. | Establish a controlled test or external/supplier/periodic verification route aligned to Clause 5.4 and Annex A. Record extract/control pH and confirmation that the difference is <= 1 pH unit. Define justified frequency in the QMS; do not automatically require every production lot. |
| A5 | Stiffness | No stiffness/deflection evidence supplied. | Mechanical performance evidence missing. | Create or link a controlled method based on Annex B / Table 2 for applicable sizes. For sizes without ISO parameters, establish a risk-based device-specific stiffness requirement. Record size, wall type, applied parameters, measured deflection, date and alloy. |
| A6 | Resistance to breakage | Current tensile/pullout tests are not equivalent. | Clause 5.9 evidence missing. | Create or link Annex C testing as qualification/periodic/routine control according to the approved quality plan. Record the Table C.1 distance, wall-type angle, 20 cycles at 0.5 Hz and no visible breakage result. |
| A7 | Corrosion | Row 22 only says "corrosion free". | Acceptance intent is present, but method/evidence is incomplete. | Revise the control to reference Annex D or an approved equivalent controlled method. Record sample identity, wall type, alloy, date, and result after approx. half immersion in 0.5 mol/L NaCl at 23 +/- 2 C for 7 h +/- 5 min. A separate periodic/qualification record is acceptable if linked. |
| A8 | Surface finish / cleanliness | Current report only partially addresses ISO wording. | Visual inspection completeness is weak. | Add explicit criteria or a controlled specification reference for straightness, regular roundness, absence of serviceability-impairing defects, metal soil and processing agents. |
| A9 | Sampling | AQL fields exist, but ISO 5.7 risk basis is not shown. | Sampling rationale not traceable. | Document the sampling basis in the tubing specification/inspection plan using risk assessment principles and intended use. If AQL is used, link it to the approved plan rather than treating AQL alone as proof of ISO 5.7 compliance. |
| A10 | Report completeness / change control | Page 2 missing; report revision dated 2013. | Evidence set and current-state confirmation incomplete. | Retrieve page 2 and perform document change-impact reconciliation against ISO 9626:2016. Update revision history and linked documents as required. |

# 12. Priority Action Plan

| Priority | Action | Deliverable / evidence required |
| --- | --- | --- |
| P1 | Confirm the controlled Spinal Needle WI/MI and link it to F-PS1-118. | Approved WI/MI number, revision, effective date and report linkage. |
| P1 | Recover and verify DD-01-PS. | Current approved specification with dimensional limits, gauge/OD mapping and wall type / ID control. |
| P1 | Establish material-grade verification. | Approved raw-material specification + ISO 15510 grade + CoA/material certificate verification record. |
| P1 | Establish / link Clause 5.4, 5.8, 5.9 and 5.10 verification routes. | Controlled methods and records for acidity/alkalinity, stiffness, breakage and corrosion; frequency justified by quality plan/risk. |
| P1 | Perform standards-change impact review for the 2013 report vs ISO 9626:2016. | Approved change assessment and revised documents if gaps are confirmed. |
| P2 | Strengthen visible surface/cleanliness and sampling traceability. | Updated acceptance criteria / specification link and risk-based sampling rationale. |
| P2 | Retrieve F-PS1-118 page 2. | Complete inspection report evidence set. |

# 13. Final Conclusion

Conclusion: The supplied evidence does not support a conclusion of full compliance with ISO 9626:2016 for the Spinal Needle tubing component. The strongest issues are evidence and document-control gaps: the product WI/MI is not available, DD-01-PS is not available, page 2 of the inspection report is missing, and the report revision predates the 2016 ISO edition. The visible report partially covers surface condition, dimensional references and corrosion intent, but does not provide verifiable controls for material grade, acidity/alkalinity, stiffness or resistance to breakage. These findings do not by themselves prove that the manufactured product is nonconforming; they show that conformity cannot be demonstrated from the current controlled-document/evidence set.

Release decision caution This gap analysis is a documentation and evidence assessment. It should not be used alone to release or reject product. Product disposition should follow the approved QMS, current controlled specifications, validated/qualified test evidence and authorized QC/QA decision-making.

# 14. Source Traceability

| Source | Exact traceability used in this assessment |
| --- | --- |
| ISO 9626:2016 | Scope: PDF p.7 (ISO p.1). Clauses 4 and 5.1-5.6: PDF p.8 (ISO p.2). Table 1: PDF pp.9-11 (ISO pp.3-5). Clauses 5.7-5.8: PDF p.11 (ISO p.5). Table 2: PDF p.12 (ISO p.6). Clauses 5.9-5.10: PDF p.13 (ISO p.7). Annex A: PDF p.14 (ISO p.8). Annex B: PDF pp.15-16 (ISO pp.9-10). Annex C: PDF pp.17-18 (ISO pp.11-12). Annex D: PDF p.19 (ISO p.13). |
| Wi.pdf | All 18 supplied pages reviewed. The file contains other-product WIs; no Spinal Needle WI was identified. |
| F-PS1-118 Rev.3 | Supplied photograph of page 1 of 2. Header and rows 1-26 reviewed. Key mapped rows: 2, 9-12, 17-18, 20-24. |
| DD-01-PS | Referenced by F-PS1-118 but not supplied; therefore no acceptance limits or revision were assumed. |
| Prior / Library search | A search for a Spinal Needle WI / F-PS1-118 did not retrieve a product-specific WI. No unrelated prior WI was substituted. |

Assessment rule applied: absence of evidence was not converted into a claim that the requirement is definitely not implemented elsewhere. Where another controlled source could reasonably exist, the finding is classified as an Evidence Gap unless the available document itself is demonstrably incomplete or inconsistent.

# Appendix A - Supplied Inspection Record Evidence

F-PS1-118 Rev.3 - Inspection & Test Report for Spinal Needle - supplied page 1 of 2.

Figure A1. Supplied inspection report page used for row-level mapping. Page 2 was not supplied.