CLOSURE REGISTER

# ISO 8637-2 Applicability
& Gap Closure Register

Sterile Blood Lines — evidence, documentary status and operational verification

| DOCUMENT PROFILE | REVIEW STATUS |
| --- | --- |
| Company | Saudi Mais Co. for Medical Products |
| Product | Sterile Blood Lines — Arterial and Venous |
| Original WI | WI-8-2-2-21, Rev.5, dated 01.09.2023 (9 pages) |
| Original in-process form | F-823-i16, Rev.8, dated 01.03.2024 (Arterial Line, photographed page 1/1) |
| ISO supplied | ISO 8637-1:2017 — not applicable as primary blood-circuit standard |
| ISO required for full assessment | ISO 8637-2:2024 — controlled copy not provided |
| Issued as | PROPOSED / TECHNICAL REVIEW DRAFT — no approval or compliance claim |
| Audit date | 18 October 2026 |

# 1. Closure rules and scope guard

ISO 8637-1:2017 Clause 1, PDF p.7 / printed p.1 excludes extracorporeal blood circuits. The applicable controlled ISO 8637-2:2024 was not supplied. Therefore all full-standard requirement extraction, clauses/pages/test-condition verification, product-specific documentary coverage percentages and ISO operational conformity conclusions are BLOCKED pending source review. The gap rows below capture legitimate company WI/form and missing-evidence actions, NOT a fabricated complete normative ISO checklist.

# 2. Source-based requirement coverage inventory

| REQ ID / SOURCE | WHAT IS VERIFIED | EXACT CLAUSE / PAGE | DOCUMENTARY STATUS | OPERATIONAL STATUS |
| --- | --- | --- | --- | --- |
| STD-SCOPE-001 / ISO 8637-1:2017 | Scope excludes extracorporeal blood circuits; ISO 8637-1 is not primary source for blood lines. | §1, PDF p.7 / printed p.1 | Verified source-scope decision | N/A to direct ISO 8637-1 device requirements |
| STD-APP-001 / ISO 8637-2:2024 | Official ISO title/scope covers disposable extracorporeal blood and fluid circuits. | Official published ISO metadata; text of Clause 1 pages unavailable | Source required for exact document pagination & clause extraction | Unable to verify |
| STD-REQ-OPEN / ISO 8637-2:2024 | ALL specific applicable clauses, tests, annexes, tables and product-configuration exceptions. | OPEN — CONTROLLED ISO STANDARD NOT PROVIDED | Not evaluated; NO percentage of complete coverage | Unable to verify |
| WI-ASSY-001 / WI Rev.5 | Company assembly and 37 arterial/venous component criteria. | §7 pp.5–7 | Reconstruction drafted; technical limitations open | Execution evidence absent |
| WI-LEAK-001 / WI Rev.5 | Individual leak test for each blood line; company tester pressure/time settings. | §8 pp.7–8 | Reconstruction drafted; tester validation open | Execution evidence absent |
| WI-LOT-001 / WI Rev.5 | Sampling and references for incoming, packaging, sterilization and final release. | §§5–6, 9–12 pp.1–9 | Reconstruction drafted; approved route records unseen | Execution evidence absent |

# 3. All audit findings / individual closure control

| GAP | CHANGE / CONTROL | SOURCE REF | EVIDENCE NEEDED | ACTUAL EVIDENCE VERIFIED | DOCUMENTARY COVERAGE | CURRENT OPERATIONAL STATUS | RESPONSIBLE FUNCTION |
| --- | --- | --- | --- | --- | --- | --- | --- |
| GAP-001 / P1 | Incorrect ISO source supplied / revision applicability | WI §13, p.9; ISO 8637-1:2017 §1, PDF p.7, printed p.1 | Controlled ISO 8637-2 copy, approved edition applicability memo and clause matrix; QA confirmation; closure after complete clause-by-clause evaluation. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Technical Input | QA / Document Control |
| GAP-002 / P1 | Unverified proof of 100% leak-test execution | WI §8, pp.7–8; form row 19, p.1 | Approved leak-test work method; equipment/program qualification; 100% cycle logs or equivalent reconciled record for representative lots and verified rejection segregation. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | Production / QC / Engineering |
| GAP-003 / P1 | Adhesive joint / pull-out control lacks verified method | WI §7.4 pp.5 and rows 5–16 pp.5–6; form rows 5–16 | Approved bond process specification and retention method, justified acceptance limits, method qualification/validation as applicable, batch sample measurements. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Technical Input | Engineering / QC / QA |
| GAP-004 / P1 | Product conformity evidence for sterile blood-contact device not supplied | WI §§10–11, p.9; ISO source scope mismatch | Approved sterilization validation and cycle records, test results, release records, controlled technical-file verification and applicability review. None supplied for review. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | QA / QC / Sterilization |
| GAP-005 / P2 | Sampling frequency not unambiguously recorded | WI §§7.6–7.7 p.5; form p.1 hourly grid | Approved interval definition and completed time-stamped checks compared with production history. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Approval | QC / Production |
| GAP-006 / P2 | Tubing dimensions lack verifiable drawing criteria in supplied pack | WI §5.3.1 p.3 and rows 7,9,10,12,15 pp.6; form rows 7,9,10,12,15 | Approved controlled drawing/BOM, dimensional specifications, gauge calibration and representative completed worksheets. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Technical Input | Engineering / QC |
| GAP-007 / P2 | AQL sampling rule lacks referenced plan and actual sampling data | WI §§5.2, 5.3.1, 7.6, 11 pp.2,3,5,9; form p.1 | Approved sampling procedure and populated evidence for representative batches; validation of plan applicability. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Technical Input | QC / QA |
| GAP-008 / P2 | Leak-test time tolerances and parameter checks incompletely represented | WI §8 p.8; form row 19 p.1 | Approved tester program and setting limits, qualification/validation, calibrated pressure reading, completed setup check records. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | Production / QC / Engineering |
| GAP-009 / P2 | Batch and check identity incomplete for reproducible event trace | WI §§7.2, 7.7 pp.5; form header p.1 | Approved traceability procedure, executed orders and in-process records with linkage. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Approval | Production / QC |
| GAP-010 / P2 | Inspection and measurement equipment identity not documented | WI §§5.2, 5.3.1, 8 pp.2–3,8; form p.1 instrument column | Calibrated equipment register, test method and linked samples; no evidence supplied. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | QC / Metrology |
| GAP-011 / P2 | Venous line execution record is not supplied for comparison | WI venous rows 21–37, pp.6–7; supplied form arterial-only | Approved and completed venous in-process template/records, or approved combined-form control. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | QC / Document Control |
| GAP-012 / P2 | Failure, recheck and reject disposition linkage not explicit | WI §8 p.8 and §11 p.9; form legend and remarks p.1 | Approved reject/NCR procedure, completed deviation examples, traceable quarantined/reworked quantities and authorized decisions. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | QC / Production / QA |
| GAP-013 / P2 | Critical component performance checks lack method/source linkage | WI §6 pp.4–5; §7 rows 4–18 pp.5–6; form p.1 | Applicable component specifications, verified methods, qualification testing as needed and completed QC worksheets. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Technical Input | Engineering / QC |
| GAP-014 / P2 | Material acceptance and supplier COA controls not independently verifiable | WI §5.1 pp.1–2 and §6 pp.3–5 | Approved incoming specs, sampled inspections, COAs, analytical results, release status and lot linkage. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | Incoming QC / Supplier Quality |
| GAP-015 / P2 | Packaging / sterilization / final-release chain lacks actual linked records in provided review | WI §§9–12 pp.8–9 | Approved packaging and sterilization WIs, actual completed records, batch genealogy and authorized release decisions. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Evidence | Production / Packaging QC / QA |
| GAP-016 / P3 | Section numbering and reference-version clarity | WI §7.6 / 7.6.1 p.5; §13 p.9; form header | Approved document change record, revised WI/form master list and controlled release. | NOT PROVIDED — verify actual controlled records | Proposed WI and in-process form: DRAFT; not approved | Open — Approval | Document Control / QA |

# 4. ISO clause-by-clause completion gate

ISO 8637-2:2024 clause matrix cannot be populated with exact clause/page/test limits without a verified, authorized text. This register intentionally does not invent or assume an ISO clause. Required future matrix columns are: requirement ID, applicability, exact clause, ISO printed page and PDF page, factual ISO paraphrase, test/acceptance/method, WI mapping, assembly or alternative evidence route, objective evidence reference, sign-off and verified closure status. Review any normative annexes and national/regulatory edition decisions before issuing product-specific WI amendments.

# 5. Closure and approval summary

| STATUS METRIC | VALUE / INTERPRETATION |
| --- | --- |
| P1 / P2 / P3 findings | 4 / 11 / 1 — document/control/evidence findings, not confirmed ISO violations |
| Confirmed ISO 8637-2 clauses fully reviewed | 0 — standard text not in supplied review package (does NOT equal 0% conformity) |
| Verified Closed findings | 0 — no executed objective evidence in supplied package |
| Proposed documentary work | Four drafts; includes reconstructed WI and redesigned IN-PROCESS control record |
| Approved / implemented | Not verified — no claimed current approval of drafts |
| Verified operational/product conformity | Unable to conclude from provided evidence |

# 6. Technical and management inputs outstanding

• Obtain approved controlled ISO 8637-2:2024 and confirm applicability, prior edition transition and any later regulatory requirements.

• Obtain TF-21-PDBM and relevant drawings/BOM/component material tolerances and controlled specifications.

• Approve leak tester method / detection threshold / validation, cycle record retention and unit-count reconciliation.

• Approve quantitative/functional test methods and limits for joints, clamping, patency, dimensions and any other identified characteristics.

• Obtain and reconcile original 2-hour sampling plan, AQL parameters, QC record, instrument calibration and venous form.

• Review linked incoming, packaging, sterilization, lab and finished product release records; approve document revisions only after review.

# 7. Source references

1. Saudi Mais: WI-8-2-2-21 Rev.5, pp.1–9 (supplied .doc). 2. Saudi Mais: photographed F-823-i16 Rev.8, Arterial Line, p.1 (supplied image). 3. ISO 8637-1:2017, Scope Clause 1, PDF p.7, ISO printed p.1 (supplied PDF). 4. ISO published record for ISO 8637-2:2024, https://www.iso.org/standard/80229.html ; FDIS development record https://www.iso.org/standard/90347.html. Other referenced company documents not supplied; no claim of review.