PROPOSED QUALITY RECORD

# In-process Control
Blood Line Assembly

Arterial + Venous — draft controlled form / not a final inspection record

| DOCUMENT PROFILE | REVIEW STATUS |
| --- | --- |
| Company | Saudi Mais Co. for Medical Products |
| Product | Sterile Blood Lines — Arterial and Venous |
| Original WI | WI-8-2-2-21, Rev.5, dated 01.09.2023 (9 pages) |
| Original in-process form | F-823-i16, Rev.8, dated 01.03.2024 (Arterial Line, photographed page 1/1) |
| ISO supplied | ISO 8637-1:2017 — not applicable as primary blood-circuit standard |
| ISO required for full assessment | ISO 8637-2:2024 — controlled copy not provided |
| Issued as | PROPOSED / TECHNICAL REVIEW DRAFT — no approval or compliance claim |
| Audit date | 18 October 2026 |

# Form control notes

Source: photographed F-823-i16 Rev.8 dated 01.03.2024 (Arterial Line) and original WI Rev.5 venous checks. This is a PROPOSED unapproved replacement/addition. Preserve existing form ID unless Document Control formally authorizes a new or amended form. The venous form identifier remains OPEN TECHNICAL INPUT REQUIRED.

Instructions: use ONE copy of the applicable configuration worksheet per actual QC check event at the original WI maximum 2-hour interval, unless approved control plan requires otherwise. OK = Pass; X = Fail; N/A = justified not applicable; PENDING = technical value/method or source missing. Retain immediate reject disposition evidence. Do not record 100% leak test as a 2-hour sample result.

# A. Manufacturing / quality record header

| MANDATORY TRACE FIELD | ENTER THE VERIFIED VALUE |
| --- | --- |
| Date / shift / actual check time | ________________________________________ |
| Product name / code / variant | ________________________________________ |
| Arterial or venous (select one) | ________________________________________ |
| Assembly lot / order / quantity produced | ________________________________________ |
| TF-21-PDBM drawing & BOM revision | ________________________________________ |
| Line / station / assembly machine ID | ________________________________________ |
| Operator / production supervisor | ________________________________________ |
| QC sampling inspector | ________________________________________ |
| Approved sampling plan / rev / population | ________________________________________ |
| Sample size n / inspected sample IDs | ________________________________________ |
| Event type: startup / routine / changeover / recheck | ________________________________________ |
| Leak tester / program ID (if evaluated) | ________________________________________ |
| Inspection / measuring equipment ID & status | ________________________________________ |

# B. Event evaluation and general rules

| CONTROL POINT | EXPECTED RECORD / EVALUATION |
| --- | --- |
| Check interval | Previous valid check time: ____ This check time: ____ Interval: ____ hours Production active between: ____ and ____ |
| Sampling | Population: ____ Plan ID/rev: ____ Level/AQL: ____ n: ____ Ac: ____ Re: ____ Defective count: ____ Lot disposition: P / F / Pending |
| Method | Visual results may be OK/X; quantifiable performance requirements require a linked method and actual numeric evidence. |
| Drawing & equipment | Drawing source/rev: ____ measuring equipment ID: ____ calibration/verification status: ____ |
| N/A / skipped controls | N/A reason & authorization: ____. No operation conducted: identify production downtime instead of assuming a missing mandatory inspection. |

# C. Arterial line — component-level in-process verification

Configuration = ARTERIAL. Each row preserves the characteristic / source acceptance wording of WI Rev.5. Each event requires actual finding and decision. Measured tests refer to separate, approved quantitative worksheets. Original rows: 1–20.

| WI ROW | CHECK / COMPONENT | CURRENT WI ACCEPTANCE BASELINE (NOT NEW ISO LIMIT) | WI METHOD | ACTUAL OBSERVATION / VALUE / EVIDENCE | RESULT |
| --- | --- | --- | --- | --- | --- |
| 1 | Appearance | Clear and good appearance, no black spot, no flashing, no holes and free from dust particle. | Visual | ________ ________ | P / F / N/A / Pend. |
| 2 | End Plug | Smooth surface, no flashing, good fitting with FLL connector, no pullout. | Visual | ________ ________ | P / F / N/A / Pend. |
| 3 | Female connector | Connector should be transparent, no crack, no damaged, good fitting with end plug. | Visual | ________ ________ | P / F / N/A / Pend. |
| 4 | Clamps (Red) Status | Clamp should be smooth surface, no flashing, no crack, no damaged, no leakage. | Visual | ________ ________ | P / F / N/A / Pend. |
| 5 | T-injection Site | Strong gluing with tubes no pullout, no crack, no damaged, no leakage, red color, with rubber material good fitting inside port. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 6 | Transducer upper and lower part | smooth surface, no crack no damaged, no leakage, proper sealed upper and lower part, strong gluing with tube, no pullout. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 7 | Transducer Tubing | Soft and kink resistance tubes, accurate in dimension, strong gluing with transducer filter and pump segment connector, no pullout. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 8 | Pump Segment Connectors | Smooth surface, accurate in dimension, no crack, no damaged, no leakage, good fitting with pump segment tube, no pullout. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 9 | Pump Segment Tube | Soft and kink resistance flexible tubes, accurate in dimension, strong gluing with connectors, no damaged, no holes, no leakage, | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 10 | Heparin Tube Status | Should be clear and transparent, no black spots, accurate in dimension, good fitting with FLL connector, no leakage. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 11 | Chamber status | Clear and transparent PVC 2way chamber, proper fitting with chamber connector, strong gluing of tubing with chamber, no pullout, no leakage and free from black spot and dust particle. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 12 | Chamber Tube | Tube should be transparent, kink resistance, accurate in dimension, free from black spots, blockage and leakage | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 13 | Dialyzer Connector (Red) | Strong gluing with tube, no pullout, no damaged, no leakage, no flashing and black spots. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 14 | Dialyzer connector cap | Should be good fitting with dialyzer connector, no damaged and crack. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 15 | Chamber tube with Transducer Filter | Clear and transparent kink resistance tube, good fitting with transducer, good fitting, strong gluing, no pullout, no crack and damaged. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 16 | FLL connector and cap status | FLL connector should be strong gluing with tube, no pullout, no crack, no damaged, cap good fitting with FLL connector, no leakage | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 17 | Spike with protector Status | No damaged, no crack, no blockage, should be sharp tip, good fitting with protector. | Visual | ________ ________ | P / F / N/A / Pend. |
| 18 | Male luer lock connectors | Transparent connectors, good fitting, no cracks, no damaged, no leakage. | Visual | ________ ________ | P / F / N/A / Pend. |
| 19 | Leakage test | Applied pressure (38-42 Kpa) / (380 – 420mbar) / (5.5 – 6.2 Psi) Charge time 15sec, Testing time 10sec and should be free from leakage. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 20 | Winding Status | Should not bend, no blockage | Visual | ________ ________ | P / F / N/A / Pend. |

Applicable drawings/test limits and methods with numerical result requirement: worksheet ID __________. Deviations and rejection reference __________. N/A reasons __________.

# C. Venous line — component-level in-process verification

Configuration = VENOUS. Each row preserves the characteristic / source acceptance wording of WI Rev.5. Each event requires actual finding and decision. Measured tests refer to separate, approved quantitative worksheets. Original rows: 21–37.

| WI ROW | CHECK / COMPONENT | CURRENT WI ACCEPTANCE BASELINE (NOT NEW ISO LIMIT) | WI METHOD | ACTUAL OBSERVATION / VALUE / EVIDENCE | RESULT |
| --- | --- | --- | --- | --- | --- |
| 21 | Appearance | Clear and good appearance, no black spot, no flashing, no holes and free from dust particle | Visual | ________ ________ | P / F / N/A / Pend. |
| 22 | End Plug | Smooth surface, no flashing, good fitting with FLL connector, no pullout | Visual | ________ ________ | P / F / N/A / Pend. |
| 23 | Female connector | Connector should be transparent, no crack, no damaged, good fitting with end plug | Visual | ________ ________ | P / F / N/A / Pend. |
| 24 | Clamps (Blue) Status | Clamp should be smooth surface, no flashing, no crack, no damaged, no leakage. | Visual | ________ ________ | P / F / N/A / Pend. |
| 25 | T-injection Site | Strong gluing with tubes no pullout, no crack, no damaged, no leakage, blue Colour, with rubber material good fitting inside port. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 26 | Chamber tube with dialyzer Connector & cap | Tube good and strong fitting with dialyzer connector and T injection site, strong gluing, no pullout, no leakage between tube and connector, accurate in dimension, cap good fitting with connector. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 27 | Transducer upper and lower part (monitor line) | smooth surface, no crack no damaged, no leakage, proper sealed upper and lower part, strong gluing with tube, no pullout | Visual | ________ ________ | P / F / N/A / Pend. |
| 28 | Transducer Tubing with chamber (monitor line) | soft and kink resistance tubes, accurate in dimension, strong gluing with transducer filter and chamber, no pullout. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 29 | Chamber tube with FLL connector status (monitor line) | Tube good and strong gluing with chamber and transducer filter, proper sealed, transducer filter no pull out from tubing, strong gluing, no leakage, good fitting of end plug and FLL connector | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 30 | Chamber status | Clear and transparent PVC 3way chamber, proper fitting with chamber connector, strong gluing of tubing with chamber, no pullout, no leakage, and free from black spot and dust particle. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 31 | Chamber filter Status | Good fitting inside chamber, no blockage, and should be free from damaged and crack | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 32 | Level Adjustment line Status | Clear and transparent tube, accurate in dimension, good fitting with chamber by strong gluing, no pull out, no leakage and blockage. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 33 | Spike with protector Status | No damaged, no crack, no blockage, should be sharp tip, good fitting with protector | Visual | ________ ________ | P / F / N/A / Pend. |
| 34 | Male luer lock connectors | Transparent connectors, good fitting, no cracks, no damaged, no leakage | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 35 | Recirculating Hook double female ports Status | Good fitting with Male luer lock, no damaged, no crack, no leakage | Physical | ________ ________ | P / F / N/A / Pend. |
| 36 | Leakage test | Applied pressure (38-42 Kpa) / (380 – 420mbar) / (5.5 – 6.2 Psi) Charge time 15sec, Testing time 10sec and should be free from leakage. | Visual / Physical | ________ ________ | P / F / N/A / Pend. |
| 37 | Winding Status | Should not bend, no blockage | Visual | ________ ________ | P / F / N/A / Pend. |

Applicable drawings/test limits and methods with numerical result requirement: worksheet ID __________. Deviations and rejection reference __________. N/A reasons __________.

# D. Dimensional and performance measurements — attached/linked evidence

| COMPONENT / SAMPLE ID | METHOD / GAUGE / DRAWING REV | ACCEPTANCE LIMIT / UNIT | ACTUAL MEASUREMENT / UNIT | RESULT / EVIDENCE ID |
| --- | --- | --- | --- | --- |
| ________________ | ________________ | OPEN TECHNICAL INPUT REQUIRED | ________________ | ________________ |
| ________________ | ________________ | OPEN TECHNICAL INPUT REQUIRED | ________________ | ________________ |
| ________________ | ________________ | OPEN TECHNICAL INPUT REQUIRED | ________________ | ________________ |
| ________________ | ________________ | OPEN TECHNICAL INPUT REQUIRED | ________________ | ________________ |
| ________________ | ________________ | OPEN TECHNICAL INPUT REQUIRED | ________________ | ________________ |

# E. Independent 100% leak-test reconciliation — cross-reference only

Original WI §8 requires individually testing EVERY arterial and venous unit. The following summary does not replace individual tester cycle evidence. The section is a linked periodic review or batch reconciliation, NOT an ISO claim and NOT final release.

| FIELD | VERIFIED VALUE |
| --- | --- |
| Assembly lot / production order | ____________________________ |
| Date / shift / tester ID / approved program | ____________________________ |
| Approved pressure setting (company WI: 38–42 kPa) | ____________________________ |
| Charge time (company WI: 15 s ±2 s) | ____________________________ |
| Test duration (company WI: 10 s ±2 s) | ____________________________ |
| Actual verified settings / program / QC check ref | ____________________________ |
| 100% machine cycle log / record ID | ____________________________ |
| Number of units assembled / submitted to test | ____________________________ |
| PASS quantity | ____________________________ |
| FAIL quantity | ____________________________ |
| Aborted/invalid/retested quantity; authorization | ____________________________ |
| Reconciled quantity / discrepancy and resolution | ____________________________ |
| Leak-test rejects / HOLD or NCR ID | ____________________________ |
| QC reviewer / date-time / decision | ____________________________ |

# F. Discrepancies, actions and QC review

| FIELD | DETAIL / VALUE |
| --- | --- |
| Failed characteristic / time / sample / quantity | ___________________________________________________________________ |
| Containment and affected period / identification | ___________________________________________________________________ |
| NCR / reject reference and segregation status | ___________________________________________________________________ |
| Recheck evidence / approval (if permitted) | ___________________________________________________________________ |
| Checked by QC: name / sign / time | ___________________________________________________________________ |
| Reviewed by QC supervisor / QCM: name / sign / time | ___________________________________________________________________ |
| Next sampling check due / verification | ___________________________________________________________________ |

# G. Controlled workflow boundaries

Assembly in-process record does not establish final sterile product conformity. Separately link (when generated): packaging F-823-i4 _____ ; sterilization WI-7-5-1-1 cycle ID _____ ; final inspection/release F-823-F6 _____. Do not prematurely sign sterile release on this form.

Document-control approval: Pending. Proposed template revision: unassigned. Controlled issue / effective date: Pending. Reference gaps: GAP-002 through GAP-015.