TECHNICAL QUALITY REVIEW  /  MEDICAL DEVICES

# Suture Removal Set
ISO Closure & Evidence Register

Provisional requirement / gap mapping — no verified ISO closure

————————————————————————————————————————————————————————————————

Product  Suture Removal Set

Source baseline  WI-8-2-2-89-1 Rev.1; F-823-i178 Rev.2

Primary ISO  OPEN — not supplied; clause/page mapping cannot be certified

Scope  In-process only

DRAFT • NOT APPROVED • NO ISO CLOSURE CLAIM

Prepared for QC Manager | 27 October 2026 | Evidence-based reconstruction from supplied records.

# 1. Closure definitions

Documentary Covered: complete proposed wording exists in a draft (not equivalent to approval). Open — Technical Input: missing approved limits, standards, source documents or sampling details. Open — Evidence: approved operational records not provided. Verified Closed: approved current control AND actual objective evidence reviewed. NO rows in this register qualify as Verified Closed.

# 2. Requirement / gap closure register

| REQ | Gap ID | Control topic | ISO Says / clause | Documentary coverage | Actual verified evidence | Operational status | Outstanding verification |
| --- | --- | --- | --- | --- | --- | --- | --- |
| RQ-001 | GAP-001 | In-process frequency and checkpoint timing | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Approval / Evidence | Approved sampling/frequency decision; blank revised form; completed actual-time logs for representative run; deviations if any |
| RQ-002 | GAP-002 | Complete set content and quantity at point of packing | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input / Evidence | Approved bill of materials/product configuration; representative completed count record linked to lot and sample; nonconformity records for failures |
| RQ-003 | GAP-003 | Dimensional limits cannot be corroborated to controlled specification | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input | Controlled TF-89-PS and relevant drawing/variant; approved measured limits and equipment details; completed dimensional observations |
| RQ-004 | GAP-004 | Stitch cutter / blade identity and pouch requirements | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input / Evidence | Approved component BOM/TF, supplier or component specification, expiry policy, sample records and pouch checks; nonconforming item record when applicable |
| RQ-005 | GAP-005 | Forceps defect definition differs between WI and form | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input / Evidence | Approved visual-defect acceptance guide, completed in-process observations, NCR for failures |
| RQ-006 | GAP-006 | Drape and gauze visual/material criteria not fully synchronized | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input / Evidence | Approved material specification TF-89-PDBM, product specification and variant list; completed visual checks and linked rejected records |
| RQ-007 | GAP-007 | Actual measurements and unambiguous results | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Evidence | Verified completed measurement records including traceable instrument IDs; approved acceptance values; QC review |
| RQ-008 | GAP-008 | Sampling quantity and AQL application are not defined for in-process records | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input | Approved in-process sampling plan and examples of completed field values; acceptance-decision records |
| RQ-009 | GAP-009 | Handling of an unacceptable in-process finding | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Approval / Evidence | Approved hold/NCR workflow; actual NCR, tag and disposition evidence; restarted checkpoint results for sampled cases |
| RQ-010 | GAP-010 | Packaging/seal/label control route not evidenced | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Evidence | Controlled F-823-i4 and linked completed records, approved labeling/packaging limits, validation records if applicable |
| RQ-011 | GAP-011 | QC review / approval responsibilities and traceability | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Approval | Approved role/approval matrix; revised form and executed signed or electronic approval records consistent with policy |
| RQ-012 | GAP-012 | Form revision identity and referenced controlled documents | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Approval | Master list, revision approval history, effective controlled documents and aligned sheet/tab |
| RQ-013 | GAP-013 | ISO applicability and full requirement verification remain open | Official ISO clause/page: NOT VERIFIED | Proposed draft only | None verified | Open — Technical Input / Evidence | Official controlled applicable ISO copies and applicability/risk determination, TF-89-PS, TF-89-PDBM, F-823-i4, F-823-T90, F-823-F75 and sample executed QC records |

# 3. Coverage of other reviewed WI topics

| Current WI area | Source reference | In-process disposition |
| --- | --- | --- |
| Purpose / scope | WI p.1 §§1–2 | Source reviewed; revised to narrow in-process scope; not an ISO coverage claim |
| Roles | WI p.1 §4 | Reviewed; approval path remains open (GAP-011) |
| Receiving inspection | WI p.1 §5.1 | Excluded from this assignment; original approved process remains separately controlled |
| Component composition / visual | WI p.1 §§5.1.1–5.2.2, p.2 §§5.2.3–5.2.5 | Reconstructed as in-process checkpoints subject to TF/BOM verification |
| Specification | WI p.2 §6 | Referenced TF-89-PS unverified |
| Packaging / labelling | WI p.2 §7 | Linked to F-823-i4 after verification; no duplicate testing |
| Final inspection and release | WI p.2 §8 | Excluded and not superseded |
| Forms and change control | WI p.2 §9, p.3 §10 | Reviewed; ref alignment open (GAP-012) |

# 4. Primary ISO assessment not yet executable

The full normative ISO coverage checklist requires the official product-specific ISO designation, edition and full text. Consequently, exact clauses, ISO printed/PDF page references, normative acceptance values, applicability and N/A justification, and 100%-documentary coverage SHALL NOT be invented. The official ISO catalog verifies only general scope of potential standards, not the source-specific clauses or product compliance.

Candidate official index pages (not a substituted controlled ISO copy): https://www.iso.org/standard/59752.html ; https://www.iso.org/standard/70799.html ; https://www.iso.org/standard/70800.html ; https://www.iso.org/standard/20417 .

# 5. Evidence intake checklist for QC Manager

☐  Approved applicable product ISO full text + exact edition (and amendments).

☐  Controlled TF-89-PS and TF-89-PDBM, BOM, cutter component specification and drawings.

☐  Approved in-process sampling plan and details of AQL use when applicable.

☐  Controlled F-823-i4 and applicable label/packaging controls.

☐  Current NC/Hold SOP and role approval matrix.

☐  Representative completed in-process records, deviations and quality review evidence.

Verified Closed count: 0. ISO clause coverage percentage: NOT ASSESSABLE from available files. All proposed records are controlled DRAFTS pending review and approval.