# Urine Bag Final Compliance Matrix

Final audit of proposed WI and evidence routes

Prepared 07 October 2026. Evidence-based reconstruction against the supplied edition only. Proposed controls and responsibilities require QMS approval; no lot disposition or conformity certification is issued by this document.

# Audit conclusion

All 16 requirement groups, including three conditional attachment routes, have a draft WI control, acceptance, method and proposed record route. This is document reconstruction, not verified product conformity. No operational requirement is classified Fully Compliant: approved inputs, implemented records and completed test evidence remain absent. Coverage labels describe the planned route. Seven open-input packages and 20 existing findings remain open until controlled approval and evidence verification.

# Final ISO WI and evidence matrix

| ISO source and requirement | Revised WI and acceptance | Report or evidence control | Coverage and remaining gap |
| --- | --- | --- | --- |
| R01 Rated volume and test volume 4.1 / 6.1 PDF 12–14 (ISO 2–4) | §9.1 Rated volume ≤ measured test volume, in ml. | Test volume worksheet and linked lot evidence Proposed report R worksheet / Section E or approved A/D/F route. F/E with A or J only when approved plan requires | Open Technical Input Required O01 O03/O04/O05 apply. No test result supplied. |
| R02 Leakage without load 4.2 / 6.2 PDF 12,14–16 (ISO 2,4–6) | §9.2 No leakage or staining after both orientations. | Unloaded leakage worksheet and evidence reference Proposed report R worksheet / Section E or approved A/D/F route. F/E; D or lot test only under approved plan | Draft control complete; evidence pending O03/O04/O05 apply. No test result supplied. |
| R03 Leakage under load 4.3 / 6.3 PDF 12,16–17 (ISO 2,6–7) | §9.3 No leakage or staining under 200 ±10 N for 1 min to 1 min 10 s. | Loaded leakage worksheet Proposed report R worksheet / Section E or approved A/D/F route. F/E; A/J only if required by approved plan | Draft control complete; evidence pending O03/O04/O05 apply. No test result supplied. |
| R04 Leakage under impact 4.4 / 6.4 PDF 12,17 (ISO 2,7) | §9.4 No leakage on any of three fresh bags; existing no-burst check also retained. | Impact worksheet with individual three-bag results Proposed report R worksheet / Section E or approved A/D/F route. F/E; periodic or lot verification only under plan | Draft control complete; evidence pending O03/O04/O05 apply. No test result supplied. |
| R05 Nonreturn valve reflux 4.5 / 6.5 PDF 12,18 (ISO 2,8) | §9.5 Reflux flowrate ≤10 ml/min. | NRV worksheet with volume time and calculated rate Proposed report R worksheet / Section E or approved A/D/F route. F/E with D or A/J as justified | Draft control complete; evidence pending O03/O04/O05 apply. No test result supplied. |
| R06 Cutout attachment 4.6 / 6.6 PDF 12,18–21 (ISO 2,8–11) | §9.6 No applicable attachment point or part fails to support bag. | Attachment worksheet for selected design with justified N/A for others Proposed report R worksheet / Section E or approved A/D/F route. F/E; D or lot control only under plan | Open Technical Input Required O02 O03/O04/O05 apply. No test result supplied. |
| R07 Button and buttonhole attachment 4.6 / 6.7 PDF 12,18–21 (ISO 2,8–11) | §9.7 No applicable attachment point or part fails to support bag. | Attachment worksheet for selected design with justified N/A for others Proposed report R worksheet / Section E or approved A/D/F route. F/E; D or lot control only under plan | Open Technical Input Required O02 O03/O04/O05 apply. No test result supplied. |
| R08 Integral suspensory system 4.6 / 6.8 PDF 12,18–21 (ISO 2,8–11) | §9.8 No applicable attachment point or part fails to support bag. | Attachment worksheet for selected design with justified N/A for others Proposed report R worksheet / Section E or approved A/D/F route. F/E; D or lot control only under plan | Open Technical Input Required O02 O03/O04/O05 apply. No test result supplied. |
| R09 Inlet tubing attachment strength 4.7 / 6.9 PDF 12,22 (ISO 2,12) | §9.9 Inlet tubing remains attached under 40 ±1 N for 1 min to 1 min 10 s. | Tubing attachment worksheet Proposed report R worksheet / Section E or approved A/D/F route. F/E; A/J if justified | Draft control complete; evidence pending O03/O04/O05 apply. No test result supplied. |
| R10 Flow initiation and average filling rate 4.8 / 6.10 PDF 12,22–23 (ISO 2,12–13) | §9.10 Flow begins within 1 min; average filling rate ≥10 ml/s. | Flow worksheet with start time fill time and ml/s Proposed report R worksheet / Section E or approved A/D/F route. F/E; D or A/J only under approved plan | Draft control complete; evidence pending O03/O04/O05 apply. No test result supplied. |
| R11 Declared dimensions 4.9 / Annex A PDF 12,24–25 (ISO 2,14–15) | §9.11 Measure declared dimensions using Annex A; approved drawing supplies product limits. ISO specifies no design sizes or tolerances. | Annex A dimension worksheet and approved drawing Proposed report R worksheet / Section E or approved A/D/F route. E/F identification; B/A/J per drawing/plan | Open Technical Input Required O01/O02 O03/O04/O05 apply. No test result supplied. |
| R12 Test temperature 5.1 PDF 12 (ISO 2) | §9.12 Atmosphere and reagent 23 ±2 °C. | Test environmental log Proposed report R worksheet / Section E or approved A/D/F route. All ISO verification routes | Open Technical Input Required O04 O03/O04/O05 apply. No test result supplied. |
| R13 End user sample state 5.2 PDF 13 (ISO 3) | §9.13 Test product samples in condition supplied to end user. | Sample-state and qualification evidence register Proposed report R worksheet / Section E or approved A/D/F route. F/E/J with linked processing qualification G | Open Technical Input Required O02/O07 O03/O04/O05 apply. No test result supplied. |
| R14 Test report completeness 7 and 6.1.5–6.10.5 / A.4 PDF 14,16–18,20–24 (ISO 4,6–8,10–14) | §9.14 Identify ISO part, bag completely, testing date/place, number tested, and each applicable method result. | Proposed report sections A–G and worksheets Proposed report R worksheet / Section E or approved A/D/F route. All routes | Open Technical Input Required O05 O03/O04/O05 apply. No test result supplied. |
| R15 Coloured water reagent 6.2.2.2 / 6.3.2.1 PDF 14,16 (ISO 4,6) | §9.15 Tap water with erythrosin E127 at 0.3 g/l for 6.2 and 6.3. | Reagent preparation record linked to F-1019-2 after confirmation Proposed report R worksheet / Section E or approved A/D/F route. All 6.2/6.3 testing | Open Technical Input Required O05 O03/O04/O05 apply. No test result supplied. |
| R16 Apparatus geometry and capability 6.1.3.2 / 6.2.3.2 / 6.3.3 / 6.5.3 / Figures 1–2 PDF 13–18 (ISO 3–8) | §9.16 Use required filling geometry and accuracies; Figures 1–2 retained as controlled source, not an assumed generic rig. | Rig qualification and calibration/capability record Proposed report R worksheet / Section E or approved A/D/F route. E for setup with use verification A/J/F | Open Technical Input Required O04 O03/O04/O05 apply. No test result supplied. |

# Control strategy matrix

Proposed allocation only. A routine lot; B incoming material; C supplier certificate; D periodic; E qualification; F type; G validation; H external laboratory; I risk-based verification; J final inspection. Supplier records may support completed bag ISO performance only when their method, configuration and supplied-state scope are verified. Certificate presence is not assumed.

| Requirement | Best proposed control location | Frequency and evidence |
| --- | --- | --- |
| R01 4.1 / 6.1 | F/E with A or J only when approved plan requires | Frequency to be defined by approved risk assessment / quality plan. Test volume worksheet and linked lot evidence |
| R02 4.2 / 6.2 | F/E; D or lot test only under approved plan | Frequency to be defined by approved risk assessment / quality plan. Unloaded leakage worksheet and evidence reference |
| R03 4.3 / 6.3 | F/E; A/J only if required by approved plan | Frequency to be defined by approved risk assessment / quality plan. Loaded leakage worksheet |
| R04 4.4 / 6.4 | F/E; periodic or lot verification only under plan | Three fresh bags per impact set. Frequency to be defined by approved risk assessment / quality plan. Impact worksheet with individual three-bag results |
| R05 4.5 / 6.5 | F/E with D or A/J as justified | Frequency to be defined by approved risk assessment / quality plan. NRV worksheet with volume time and calculated rate |
| R06 4.6 / 6.6 | F/E; D or lot control only under plan | Frequency to be defined by approved risk assessment / quality plan. Attachment worksheet for selected design with justified N/A for others |
| R07 4.6 / 6.7 | F/E; D or lot control only under plan | Frequency to be defined by approved risk assessment / quality plan. Attachment worksheet for selected design with justified N/A for others |
| R08 4.6 / 6.8 | F/E; D or lot control only under plan | Frequency to be defined by approved risk assessment / quality plan. Attachment worksheet for selected design with justified N/A for others |
| R09 4.7 / 6.9 | F/E; A/J if justified | Frequency to be defined by approved risk assessment / quality plan. Tubing attachment worksheet |
| R10 4.8 / 6.10 | F/E; D or A/J only under approved plan | Frequency to be defined by approved risk assessment / quality plan. Flow worksheet with start time fill time and ml/s |
| R11 4.9 / Annex A | E/F identification; B/A/J per drawing/plan | Frequency to be defined by approved risk assessment / quality plan. Annex A dimension worksheet and approved drawing |
| R12 5.1 | All ISO verification routes | Frequency to be defined by approved risk assessment / quality plan. Test environmental log |
| R13 5.2 | F/E/J with linked processing qualification G | Frequency to be defined by approved risk assessment / quality plan. Sample-state and qualification evidence register |
| R14 7 and 6.1.5–6.10.5 / A.4 | All routes | Frequency to be defined by approved risk assessment / quality plan. Proposed report sections A–G and worksheets |
| R15 6.2.2.2 / 6.3.2.1 | All 6.2/6.3 testing | Frequency to be defined by approved risk assessment / quality plan. Reagent preparation record linked to F-1019-2 after confirmation |
| R16 6.1.3.2 / 6.2.3.2 / 6.3.3 / 6.5.3 / Figures 1–2 | E for setup with use verification A/J/F | Frequency to be defined by approved risk assessment / quality plan. Rig qualification and calibration/capability record |

# Inspection record allocation

| Section | Required content and condition |
| --- | --- |
| A routine lot | Company receiving/appearance/fit/printing/sealing/dimensions/packaging checks; ISO tests only if assigned by approved plan. |
| B incoming | PO identity, packaging, BOM/material criteria from approved specifications; no ISO material grade invented. |
| C supplier | Certificate identity, scope, specification and results/declaration verified by QC/Supplier Quality. |
| D periodic | Defined tests and due dates from approved risk/quality plan; not arbitrarily monthly or yearly. |
| E qualification and type | Preferred design-performance route for ISO 6.1–6.10 and Annex A; actual scope approved in plan. |
| F external | Use when internal capability or approved plan calls for external testing; link method and result scope. |
| G other evidence | Rig qualification, reagent preparation, final-state/sterilization, product drawing, linked forms and release evidence. |

# Cross document consistency check

| Check | Audit result |
| --- | --- |
| Product scope | Revised WI and template use non-body-worn urine drainage bag. Attachment subtype remains O02; no paediatric/urimeter extension. |
| Volume | Both use rated ≤ measured test volume; 0.75 l4; O01 supplies manufacturer rating. |
| Loads | Both use total 200 ±10 N compression, 100 ±5 N NRV, 40 ±1 N inlet tension. 20 kg shorthand removed as governing ISO limit. |
| NRV | Both use ≤10 ml/min; volume/time recorded; not >10 ml and not millimeters/min. |
| Flow | Both use onset within 1 min; mean ≥10 ml/s; 150 ±5 mm head. No-flow observation stopping window remains 1 min to 1 min 10 s, not a relaxed acceptance deadline. |
| Durations and impact | Unloaded 17 ±1 h then 4 ±1 h; support 24 ±1 h; NRV 5 min to 5 min 10 s; three fresh impact bags, 500 ±10 mm drop. |
| Dimensions | Annex A nearest 0.5/10/1 mm definitions consistent; ISO supplies no product dimensional limits. OD remains company requirement. |
| Reagent | E127 0.3 g/l for ISO 6.2/6.3; safranin not assumed equivalent. |
| Environment and state | 23 ±2 °C atmosphere/reagent and end-user supplied state in WI and report. |
| Document identifiers | Existing WI/form match F-823-T16; draft revision not fabricated. Linked forms/revisions remain O05. |
| Sampling | No invented AQL/frequency; three-bag impact requirement retained. Prior S-2/AQL1.5 context not substituted for controlled approval. |
| Material and sterilization | No grade/dose/chemical limit invented. Existing packaging/sterilization/release retained with O06/O07 dependencies. |

# Open technical inputs

## O01 Approved product specification and drawing

OPEN ITEM — TECHNICAL INPUT REQUIRED. Rated volume, declared dimensions, OD/ID/length limits and revisions are absent. Dependent requirements: R01/R11. Required controlled document: Controlled product specification and drawing. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

## O02 Product configuration and attachment system

OPEN ITEM — TECHNICAL INPUT REQUIRED. Confirm product codes, non-body-worn adult scope, supplied NRV, attachment design, connector and final processing state. Dependent requirements: R05–R13. Required controlled document: Product specification and approved configuration list. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

## O03 Sampling and control plan

OPEN ITEM — TECHNICAL INPUT REQUIRED. Lot size mappings, inspection level, AQL, sample count, test allocation/frequency and escalation are not available. S-2 and AQL 1.5 mentioned in prior context are not treated as an approved product plan. Dependent requirements: All. Required controlled document: Approved inspection plan and risk assessment. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

## O04 Equipment and rig definition

OPEN ITEM — TECHNICAL INPUT REQUIRED. No equipment IDs, qualification, calibration or controlled Figure 1/2 rig drawings available. Dependent requirements: R01–R12/R16. Required controlled document: Controlled method setup and equipment file. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

## O05 Linked controlled records

OPEN ITEM — TECHNICAL INPUT REQUIRED. Confirm form revisions, ownership, test worksheet allocation and supersession. Dependent requirements: R14/R15. Required controlled document: Document master list and controlled forms. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

## O06 Incoming materials and supplier controls

OPEN ITEM — TECHNICAL INPUT REQUIRED. No material/packaging specification, BOM, supplier CoA/CoC or approved verification plan available; no grade or chemical limit is invented. Dependent requirements: Company WI §5.1/8. Required controlled document: Approved BOM, raw material/packaging specifications and supplier quality plan. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

## O07 Sterilization and release linkage

OPEN ITEM — TECHNICAL INPUT REQUIRED. WI-7-5-1-1, cycle evidence, laboratory results, packaging parameters and release records unavailable. Dependent requirements: R13 and company WI §8–10. Required controlled document: Approved sterilization WI and validation, packaging specification and release records. Owner proposed: QA with QC and Engineering; O06 also Purchasing/Supplier Quality; O07 also Production and Sterilization. Closure requires approved source and linked evidence, not a blank field.

# Closure and release evidence checklist

For each R ID: approve applicability → issue WI and specification → verify acceptance and setup → assign control location/frequency/sample count → obtain actual signed results or scoped supplier/external evidence → QC review → link to lot/configuration → authorized release. An open input or missing required record prevents a Fully Covered operational status. N/A requires approved configuration evidence and cannot conceal an untested requirement.

# Final conclusion

Existing WI covers much of the test sequence but does not fully control the supplied ISO. Existing form does not prove conformity and contains a critical NRV criterion error. Revised draft closes textual limits/method/unit/traceability gaps and introduces usable worksheets. Qualification/periodic/supplier/external evidence can replace routine repetition only through the approved quality plan, not by assumption. Open technical inputs and three document-control findings require approval and records before conformity or release can be asserted.