# Urine Bag Proposed Inspection Report

Editable lot record and linked ISO test worksheets

Prepared 07 October 2026. Evidence-based reconstruction against the supplied edition only. Proposed controls and responsibilities require QMS approval; no lot disposition or conformity certification is issued by this document.

Candidate replacement of F-823-T16 under change control only. Report number F-823-T16 proposed revision ____; current source Rev.3. Approval/effective date ____ . This blank template proves no test or compliance. Use Sections A–G to assign evidence without automatically making every ISO test a lot test.

# Identification and sampling

Product name Urine Drainage Bag    Product code ____    Size / rated volume ____ ml

Lot / batch ____    Supplier ____    PO ____    Quantity ____

Inspection date ____    Testing place ____    Inspector ____

WI number WI-8-2-2-39    Approved WI revision ____

ISO 8669-2:1996 / BS EN ISO 8669-2:1997

Product specification ID / revision ____    Drawing ____

Raw material specification ID / revision ____    Applicable processing state ____

Lot size ____    Inspection level ____    Sample size ____    AQL if applicable ____

Sampling plan ID / revision ____    Code letter ____    Ac ____    Re ____

Control allocation / risk plan ID ____    Attachment design ____    NRV fitted ____

# Equipment and environmental records

| Equipment name | ID | Calibration or capability status | Due date |
| --- | --- | --- | --- |
| ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ |
| ____ | ____ | ____ | ____ |

Atmosphere ____ °C    Reagent ____ °C    Required 23 ±2 °C. Reagent record ____ . Rig drawing/qualification ____ . Sample state and sterilization evidence ____ .

# Section A Routine lot inspection

Apply only characteristics allocated to this lot by approved plan. Existing company inspections below are retained; missing specification limits require Pending, not acceptance by generic wording.

| No and inspection | WI and source | Acceptance and method | Sample size and actual result | Result and remarks |
| --- | --- | --- | --- | --- |
| A01 Packing condition | WI §8/13–15 Company control | No damage; secured; approved packing specification Visual | n ____ actual ____ | Pass / Fail / Pending ____ |
| A02 Bag and welding | WI §8/13–15 Company control | Clear appearance; no visible weld defect; structural strength verified by linked ISO tests Visual plus evidence | n ____ actual ____ | Pass / Fail / Pending ____ |
| A03 Inlet and outlet fitting | WI §8/13–15 Company control | No loose fitting; no damage, flashing or short shot Physical / visual | n ____ actual ____ | Pass / Fail / Pending ____ |
| A04 Capacity and scale | WI §8/13–15 Company control | Rated volume per specification; scale clearly printed; R01 evidence required Visual plus volume evidence | n ____ actual ____ | Pass / Fail / Pending ____ |
| A05 Logo and printing | WI §8/13–15 Company control | Approved Mais artwork; clear lot/reference printing Visual / approved artwork | n ____ actual ____ | Pass / Fail / Pending ____ |
| A06 Fixing lace | WI §8/13–15 Company control | Flexible; no crack/damage; adjustable button where fitted; R06–R08 support evidence separate Visual / physical | n ____ actual ____ | Pass / Fail / Pending ____ |
| A07 Inlet and outlet sealing | WI §8/13–15 Company control | No leakage under approved company method; ISO R02/R03 separately evidenced Approved method required | n ____ actual ____ | Pass / Fail / Pending ____ |
| A08 Tube dimensions and length | WI §8/13–15 Company control | Approved drawing limits ____; ID/OD/length actual values; Annex A where declared Dimensional | n ____ actual ____ | Pass / Fail / Pending ____ |
| A09 Packaging label and sterile release | WI §8/13–15 Company control | Approved §13 label; §14–15 sterilization and lab release evidence Records / visual | n ____ actual ____ | Pass / Fail / Pending ____ |

ISO test performed on this lot? Record applicable R ID ____; worksheet ____; sample count ____; actual ____; Pass/Fail ____; remarks ____ . Otherwise link qualified evidence in Section G. Additional approved lot-specific tests may be entered without changing ISO acceptance limits.

# Section B Incoming and raw material verification

PO/shipment matching, packaging integrity and component identity against approved BOM/specification. Specification and revision ____; supplier lot ____; characteristic/limit ____; method/equipment ____; sample count ____; actual ____; Pass/Fail/Pending ____; incoming record ____ . No raw-material grade or chemical limit is supplied by this ISO.

# Section C Supplier CoA and CoC verification

Certificate ID ____; supplier ____; product/material/lot ____; standard/specification/revision ____; scope and numerical results ____; authenticity/authorized issue check ____; QC/Supplier Quality reviewer ____; date ____; accepted supporting evidence or pending ____ . Generic CoC does not establish all bag performance. If supplier performs ISO tests, verify final configuration, method, sample state and result detail.

# Section D Periodic verification

Approved plan ID/revision ____; requirement R ID ____; interval/next due from approved plan ____; previous report ____; method/actual/result ____; linked worksheet ____ . Frequency to be defined by approved risk assessment / quality plan.

# Section E Qualification and type tests

Preferred candidate route for R01–R11 final-state performance and design characteristics; assignment requires approved plan. Product configuration ____; supplied-state processing ____; qualifying lot/sample IDs ____; test date/place ____; sample counts ____; worksheet results ____; approved report ____; validity/change assessment ____ .

# Section F External laboratory tests

External testing is optional when internal capability is unavailable or plan requires it. Laboratory ____; scope/capability verification ____; sample identity/state ____; ISO method ____; report/date ____; result ____; reviewer/acceptance ____ . No chemical test is invented for this ISO.

# Section G Other controlled evidence

Link apparatus qualification, reagent record, sample-state/sterilization evidence, packaging control, approved drawing, risk/quality plan and final release. Every applicable R ID must have either actual report result or valid referenced evidence. Missing evidence = Pending, not N/A.

| Requirement | Control section A–G | Evidence ID and revision | Result validity and review |
| --- | --- | --- | --- |
| R01 4.1 / 6.1 | ____ | Test volume worksheet and linked lot evidence ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R02 4.2 / 6.2 | ____ | Unloaded leakage worksheet and evidence reference ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R03 4.3 / 6.3 | ____ | Loaded leakage worksheet ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R04 4.4 / 6.4 | ____ | Impact worksheet with individual three-bag results ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R05 4.5 / 6.5 | ____ | NRV worksheet with volume time and calculated rate ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R06 4.6 / 6.6 | ____ | Attachment worksheet for selected design with justified N/A for others ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R07 4.6 / 6.7 | ____ | Attachment worksheet for selected design with justified N/A for others ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R08 4.6 / 6.8 | ____ | Attachment worksheet for selected design with justified N/A for others ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R09 4.7 / 6.9 | ____ | Tubing attachment worksheet ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R10 4.8 / 6.10 | ____ | Flow worksheet with start time fill time and ml/s ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R11 4.9 / Annex A | ____ | Annex A dimension worksheet and approved drawing ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R12 5.1 | ____ | Test environmental log ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R13 5.2 | ____ | Sample-state and qualification evidence register ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R14 7 and 6.1.5–6.10.5 / A.4 | ____ | Proposed report sections A–G and worksheets ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R15 6.2.2.2 / 6.3.2.1 | ____ | Reagent preparation record linked to F-1019-2 after confirmation ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |
| R16 6.1.3.2 / 6.2.3.2 / 6.3.3 / 6.5.3 / Figures 1–2 | ____ | Rig qualification and calibration/capability record ID ____ rev ____ | Pass / Fail / Pending / justified N/A Reviewer ____ |

# ISO inspection and test requirements

| Test and ISO | WI reference and criterion | Method equipment sample size | Actual and result |
| --- | --- | --- | --- |
| R01 Rated volume and test volume 4.1 / 6.1 | §9.1 Rated volume ≤ measured test volume, in ml. | Figure 1 system; graduated cylinder ±2%; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R02 Leakage without load 4.2 / 6.2 | §9.2 No leakage or staining after both orientations. | Figure 2 system; graduated cylinder ±2%; white absorbent material; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R03 Leakage under load 4.3 / 6.3 | §9.3 No leakage or staining under 200 ±10 N for 1 min to 1 min 10 s. | Flat plate; restraints ±5°; verified total force system; cylinder; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R04 Leakage under impact 4.4 / 6.4 | §9.4 No leakage on any of three fresh bags; existing no-burst check also retained. | Cylinder ±2%; Figure 2; height measure; smooth rigid impact surface Approved plan n ____ 3 fresh bags per set | Actual ____ Pass / Fail / Pending Evidence ____ |
| R05 Nonreturn valve reflux 4.5 / 6.5 | §9.5 Reflux flowrate ≤10 ml/min. | 100 ±5 N total-force system; plate/restraints; cylinders ±2%; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R06 Cutout attachment 4.6 / 6.6 | §9.6 No applicable attachment point or part fails to support bag. | Rod 6 ±0.1 mm; full-width clamp; verified masses; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R07 Button and buttonhole attachment 4.6 / 6.7 | §9.7 No applicable attachment point or part fails to support bag. | Supplied-type straps; support; clamp; verified masses; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R08 Integral suspensory system 4.6 / 6.8 | §9.8 No applicable attachment point or part fails to support bag. | Suitable support; full-width clamp; verified masses; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R09 Inlet tubing attachment strength 4.7 / 6.9 | §9.9 Inlet tubing remains attached under 40 ±1 N for 1 min to 1 min 10 s. | Clamp; verified 40 ±1 N weight system; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R10 Flow initiation and average filling rate 4.8 / 6.10 | §9.10 Flow begins within 1 min; average filling rate ≥10 ml/s. | Figure 2 system; cylinder ±2%; timer Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R11 Declared dimensions 4.9 / Annex A | §9.11 Measure declared dimensions using Annex A; approved drawing supplies product limits. ISO specifies no design sizes or tolerances. | Suitable dimensional tools with 0.5 mm capability Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R12 Test temperature 5.1 | §9.12 Atmosphere and reagent 23 ±2 °C. | Verified temperature instrument Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R13 End user sample state 5.2 | §9.13 Test product samples in condition supplied to end user. | Sample identification and processing records Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R14 Test report completeness 7 and 6.1.5–6.10.5 / A.4 | §9.14 Identify ISO part, bag completely, testing date/place, number tested, and each applicable method result. | Controlled report and worksheets Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R15 Coloured water reagent 6.2.2.2 / 6.3.2.1 | §9.15 Tap water with erythrosin E127 at 0.3 g/l for 6.2 and 6.3. | Reagent preparation measuring equipment Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |
| R16 Apparatus geometry and capability 6.1.3.2 / 6.2.3.2 / 6.3.3 / 6.5.3 / Figures 1–2 | §9.16 Use required filling geometry and accuracies; Figures 1–2 retained as controlled source, not an assumed generic rig. | Controlled rig drawing and equipment verification Approved plan n ____ Frequency per approved plan | Actual ____ Pass / Fail / Pending Evidence ____ |

# Method specific worksheets

Copy result lines as needed for each actual sample. Identify bag/sample ____ and test date/place ____ on every worksheet. Full methods are in revised WI §9; record actual values, not only check marks.

## R01 Rated volume and test volume

ISO 4.1 / 6.1 | Revised WI §9.1

Rated ____ ml; test volume ____ ml; l4 ____ mm; mark ____ mm; head ____ mm; start/stop time ____; elapsed ____ s; cylinder accuracy ____; leak-invalid/repeat ____; rated ≤ test result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R02 Leakage without load

ISO 4.2 / 6.2 | Revised WI §9.2

Test volume ____ ml; 75% fill ____ ml; head ____ mm; E127 ____ g/l; drain cycled ten times ____; horizontal start/end ____ elapsed ____ h; leakage/location ____; vertical start/end ____ elapsed ____ h; leakage/location ____; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R03 Leakage under load

ISO 4.3 / 6.3 | Revised WI §9.3

75% fill ____ ml; head ____ mm; plate size ____; plate angle ____ °; total force incl plate ____ N; elapsed ____ s; staining/leakage/location ____; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R04 Leakage under impact

ISO 4.4 / 6.4 | Revised WI §9.4

50% fill ____ ml; head ____ mm; lowest-edge drop height ____ mm. Fresh bag 1 bottom: ID ____ leak/location/burst ____; fresh bag 2 face: ID ____ leak/location/burst ____; fresh bag 3 side: ID ____ leak/location/burst ____; overall result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R05 Nonreturn valve reflux

ISO 4.5 / 6.5 | Revised WI §9.5

Valve presence/location ____; cut tube length ____ mm; 50% fill ____ ml; plate size/angle ____; actual total force ____ N; pre-load flow excluded ____; collection elapsed ____ s / ____ min; reflux collected ____ ml; measurement accuracy ____; rate ____ ml/min; ≤10 ml/min result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R06 Cutout attachment

ISO 4.6 / 6.6 | Revised WI §9.6

Cutout applicability and evidence ____; rod ____ mm; minimum cutouts per instructions ____; mass of test-volume water ____; clamp mass ____; added mass ____; combined ratio ____; duration ____ h; breakthrough to outside/support failure ____; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R07 Button and buttonhole attachment

ISO 4.6 / 6.7 | Revised WI §9.7

Button design applicability/evidence ____; supplied strap type ____; spacing ____; combined load ratio ____; duration ____ h; button/buttonhole support failure ____; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R08 Integral suspensory system

ISO 4.6 / 6.8 | Revised WI §9.8

Integral system applicability/evidence ____; support/instructions ____; combined load ratio ____; duration ____ h; system support failure ____; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R09 Inlet tubing attachment strength

ISO 4.7 / 6.9 | Revised WI §9.9

Sample ____; downward inlet/clamp setup ____; actual tensile force ____ N; elapsed ____ s; shock avoided ____; detached yes/no ____; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R10 Flow initiation and average filling rate

ISO 4.8 / 6.10 | Revised WI §9.10

Test volume ____ ml; reservoir volume 50% ____ ml; head ____ mm; tube support/no kink ____; flow onset ____ s; within 1 min yes/no ____; total fill time ____ s; no-flow stop ____ s; leakage-invalid/repeat ____; mean rate ____ ml/s; ≥10 ml/s result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R11 Declared dimensions

ISO 4.9 / Annex A | Revised WI §9.11

Drawing ID/rev ____; instrument capability ____; empty/flat/tap closed ____; d1 ____ mm; d2 ____ mm and minimum pathway location ____; l1 ____ mm; l2 ____ mm; l3 ____ mm; l4 ____ mm; l5 ____ mm; applicable drawing limits ____; tube OD company check ____ mm; result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R12 Test temperature

ISO 5.1 | Revised WI §9.12

Atmosphere ____ °C; reagent ____ °C; instrument ID ____; time ____; 23 ±2 °C result ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R13 End user sample state

ISO 5.2 | Revised WI §9.13

Supplied configuration ____; sterile/bulk state ____; processing record ____; representative end-user sample evidence ____; applicability review ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R14 Test report completeness

ISO 7 and 6.1.5–6.10.5 / A.4 | Revised WI §9.14

ISO reference ____; complete bag identity ____; date/place ____; number tested ____; all required method results linked ____; review ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R15 Coloured water reagent

ISO 6.2.2.2 / 6.3.2.1 | Revised WI §9.15

Tap water ____; E127 ID ____; mass ____ g; final volume ____ l; concentration ____ g/l; preparation record ____; 0.3 g/l verified ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

## R16 Apparatus geometry and capability

ISO 6.1.3.2 / 6.2.3.2 / 6.3.3 / 6.5.3 / Figures 1–2 | Revised WI §9.16

Figure 1/2 drawing rev ____; head/nozzle/reservoir/baffles verified ____; cylinder capability ____; plate geometry/±5° ____; load verification ____; rig qualification ____; reviewer ____.

Evidence ID ____; equipment IDs ____; number of bags ____; Pass / Fail / Pending / justified N/A ____; remarks ____; inspector ____; reviewer ____ .

# Disposition and approval

Disposition: [ ] Accepted   [ ] Rejected   [ ] Quarantine   [ ] Pending Additional Test. Conditional status only under approved QMS authority/procedure; no such authorization was supplied.

Applicable routine tests passed ____; linked ISO evidence valid ____; deviations/NCR ____; open release blockers ____ .

QC Inspector ____ Date ____    QC Reviewer ____ Date ____

QC Supervisor if authorized ____ Date ____    QC Manager ____ Date ____

Final Release Authority under approved QMS ____ Signature ____ Date ____

Legend: Pass = meets defined criterion; Fail = does not meet; Pending = incomplete verification; N/A = not applicable with approved design justification. Blank or Good/OK alone is not a numeric result.